## Sheet1
uns_tag | extracted_content | SRS Tags (Combined) | Content
Not applicable, no user requirement |  | SRS.Allura.Manf.VerifySystem ¦ SRS.Allura.Phys.TechnicalRoomCabinetsSpace ¦ SRS.Allura.I&V.ManualTest ¦ SRS.Allura.SLS.ECO.EnergyConsumption ¦ SRS.Allura.I&V.TestInterface ¦ SRS.Allura.I&V.Environment ¦ SRS.Allura.Manf.LoadConfiguration | The Manufacturer Engineer shall be able to verify that the system is manufactured according to the specification independent of Commercial licenses. Tolerances shall be specified for critical measurements which can be influenced by the manufacturing process. In combination with IP based video infrastructure, system components in the technical room shall be placed in max 3 (monoplane systems) and max 4 (biplane systems) 19’’ cabinets. Electrical Most electrical requirements stem from targets as derived from requirements stated in section 12.3 Standards, Legal Requirements. For manual testing the system shall have: A viewer for logging and tracing; Capabilities for content and detail level scaling of system logging and tracing; File handling possibilities for logging and tracing and system data; Customizable EPX data extraction. Support for the use of product test tools; Support for fault injection for safety requirements verification. Applicable Standards International markets The list below applies to x-ray imaging equipment including embedded software that cannot be used separately from equipment. The list is applicable to show compliance with the 3rd edition of IEC60601-1 and its collateral and particular standards. Table: International standard Edition Title IEC 60601-1:2005 + Corrigendum C1:2006 + Corrigendum C2:2007 + Interpretation ISH1:2008 + Interpretation ISH2:2009 + Amendment A1:2012 3.1 Medical electrical equipment – Part 1: General requirements for basic safety and essential performance IEC 60601-1-2:2007 3.0 Collateral standard: Electromagnetic compatibility – Requirements and tests IEC 60601-1-2:2014 + Amendment A1:2020 4.1 Collateral standard: Electromagnetic disturbances – Requirements and tests IEC 60601-1-3:2008 + Amendment A1:2013 2.1 Collateral standard: General requirements for radiation protection in diagnostic X-Ray equipment IEC 60601-1-6:2010 + Amendment A1:2013 3.1 Collateral standard: Usability IEC 60601-1-9:2007 + Amendment A1:2013 1.1 Collateral Standard: Requirements for environmentally conscious design IEC 60601-2-43:2010 + Amendment A1:2017 + Amendment A2:2019 2.2 Particular requirements for the basic safety and essential performance of X-ray equipment for interventional procedures IEC 60601-2-54:2009 + Amendment A1:2015 + Amendment A2:2018 1.2 Particular requirements for the basic safety and essential performance of X-ray equipment for radiography and radioscopy IEC 60825-1:2014 3.0 Safety of laser products – Part 1: Equipment classification and requirements IEC 60825-2:2004 + Amendment A1:2006 + Amendment A2:2010 + Interpretation ISH1:2008 3.2 Safety of laser products – Part 2: Safety of optical fiber communication systems IEC 62304:2006 + Amendment A1:2015 1.1 Medical device software – Software life cycle processes IEC 62366:2007 + Amendment A1:2014 1.1 Medical devices – Application of usability engineering to medical devices IEC 62366-1:2015 + Amendment A1:2020 1.1 Medical devices – Part 1: Application of usability engineering to medical devices ISO 14971:2019 3.0 Medical devices – Application of risk management to medical devices The following standards are normatively referenced and are needed to demonstrate compliance with the “top-level” standards listed above. Table: Normative reference Edition Title IEC 60336:2005 + C1:2006 4.0 X-ray tube assemblies for medical diagnosis – Characteristics of focal spots IEC 60522:1999 2.0 Determination of the permanent filtration of X-ray tube assemblies IEC 60526:1978 + C1:2010 2.0 High-voltage cable plug and socket connections for medical X-ray equipment IEC 60529:1989 + A1:1999 + A2:2013 + C1:2013 + C2:2015 + A2/C1:2019 2.2 Degrees of protection provided by enclosures (IP Code) IEC 60580:2019 3.0 Medical electrical equipment – Dose area product meters IEC 60601-1-8:2006 + A1:2012 2.1 Note 1 Part 1-8: Collateral standard: General requirements, tests and guidance for alarm systems in medical electrical equipment and medical electrical systems IEC 60601-2-28:2017 3.0 Part 2-28: Particular requirements for the basic safety and essential performance of X-ray tube assemblies for medical diagnosis IEC 60613:2010 3.0 Electrical, thermal and loading characteristics of rotating anode X-ray tubes for medical diagnosis IEC 60627:2013 3.0 Diagnostic X-ray imaging equipment – Characteristics of general purpose and mammographic anti-scatter grids IEC 60806:1984 1.0 Determination of the maximum symmetrical radiation field from a rotating anode X-ray tube for medical diagnosis IEC 60950-1:2005 + C1:2006 + A1:2009 + Corrigendum of A1 (2012) +A2:2013 2.2 Note 2 Information technology equipment – Safety. Part 1: General requirements IEC 62368-1:2018 3.0 Note 2 Audio/video, information and communication technology equipment – Part 1: Safety requirements IEC 61331-n:2014 (series) 2.0 Protective devices against diagnostic medical X-radiation (Parts 1 to 3) IEC 61910-1:2014 1.0 Medical electrical equipment – Radiation dose documentation – Part 1: Equipment for radiography and radioscopy DICOM (2016c) Digital Imaging and Communications in Medicine (http://medical.nema.org) IEC 62220-1-3:2008 1.0 Characteristics of digital X-ray imaging devices. Part 1-3: Determination of the detective quantum efficiency – detectors used in dynamic imaging ISO 10993-1:2018 5.0 Biological evaluation of medical devices. Part 1: Evaluation and testing within a risk management process ISO 15223-1:2021 4.0 Medical devices – Symbols to be used with medical device labels, labelling and information to be supplied ISO 20417:2021 1.0 Medical devices – Information to be supplied by the manufacturer IEC 80001-1:2010 1.0 Application of risk management for IT-networks incorporating medical devices – Part 1: Roles, responsibilities and activities DIN6868-157:2014 - Image Quality assurance in diagnostic X-ray departments - Part 157: X-ray ordinance acceptance and constancy tests of image display systems in their environment. DIN6862-2:2019 - Identification and characterization of radiological images in medical diagnosis - Part 2: Passing on of radiographs and related records in digital radiography, digital fluoroscopy and computed tomography ISO 17664-2:2021 1.0 Processing of health care products — Information to be provided by the medical device manufacturer for the processing of medical devices — Part 2: Non-critical medical devices Note 1: IEC 60601-1-8 does not apply, as stated in IEC 60601-2-54:2009 and IEC 60601-2-43:2010. However, it must be checked that audible signals generated by x-ray imaging equipment do not interfere with alarm signals generated by other medical devices conforming with IEC 60601-1-8. Note 2: This standard is only applicable to IT equipment European Economic Area The list below contains all European Directives and Regulations applicable to medical devices and to medical electrical equipment emitting ionizing radiation. Compliance with harmonized EN standards gives presumption of conformity with the General Safety and Performance Requirements (MDR, Annex I). Table: Directives Published Title EU 2017/745 2017-04-05 Amended: 2020-04-23 Regulation (EU) 2017/745 of the European Parliament and the Council of 5 April 2017 on medical devices, amending Directive 2001/83/EC, Regulation (EC) No 178/2002 and Regulation (EC) No 1223/2009 and repealing Council Directives 90/385/EEC and 93/42/EEC 2014/53/EU 2014-04-16 Amended: 2018-07-04 Directive of the European Parliament and of the Council on the making available on the market of radio equipment and repealing Directive 1999/5/EC 94/62/EC 1994-12-20 Amended: 2018-05-30 European Parliament and Council directive on packaging and packaging waste 2006/66/EC 2006-09-06 Amended: 2013-11-20 Directive of the European Parliament and of the Council on batteries and accumulators and waste batteries and accumulators EC 1907/2006 2006-12-18 Amended: 2019-10-10 Regulation of the European Parliament and of the Council concerning the Registration, Evaluation, Authorisation and Restriction of Chemicals (REACH) 2013/59/EURATOM Note 1 2013-12-05 Amended: 2014-01-17 Council directive laying down basic safety standards for protection against the dangers arising from exposure to ionizing radiation 2011/65/EU (RoHS) 2011-06-08- Amended: 2021-11-23 Directive of the European Parliament and of the Council on the restriction of the use of certain hazardous substances in electrical and electronic equipment (commonly referred to as RoHS); SL with Annex IV 2012/19/EU 2012-07-24 Amended 2018-06-14 Directive of the European Parliament and of the Council on waste of electrical and electronic equipment (WEEE, recast) EU 2021/2226 2021-12-14 Commission Implementing Regulation (EU) 2021/2226 of 14 December 2021 laying down rules for the application of Regulation (EU) 2017/745 of the European Parliament and of the Council as regards electronic instructions for use of medical devices. DE: Sachverständigen-Prüfrichtlinie (SV-RL) Note 2 2020-07-01 Richtlinie für die technische Prüfung von Röntgeneinrichtungen und genehmigungsbedürftigen Störstrahlern durch Sachverständige nach dem Strahlenschutzgesetz und der Strahlenschutzverordnung English: Guideline for the technical inspection of X-ray equipment and interference radiation requiring approval by experts in accordance with the Radiation Protection Act and the Radiation Protection Ordinance. Note 1 The EURATOM directive does not directly apply to medical electrical equipment but more to the use of (equipment emitting) ionizing radiation. On request from customers, a statement must be provided how Philips Healthcare can assist in meeting the requirements of this directive. Note 2 The compliance to this requirement is managed by the local organization (supplier). A local independent inspection body will verify this compliance. Switzerland Table: Regulation Published Title CH: 814.542.1 2017-04-26 Verordnung des EDI über den Strahlenschutz bei medizinischen Röntgensystemen (Röntgenverordnung, RöV) English: Ordinance of the Swiss Federal Department of Radiation Protection in Medical X-Ray Systems (X-ray Ordinance) United Kingdom Table: Statutory Instrument Published Title SI 2002 No. 618 2002-05-21 Amended: 2020-12-08 Statutory Instruments 2002 No. 618, Consumer Protection, The Medical Devices Regulations 2002 (as last amended by SI 2020 No.1478) EN standards used for compliance with other directives Table: EN standard Edition Notes Title EN 62321-n (2013-2014) 1.0 1 Determination of certain substances in electrotechnical products – Parts 1, 2, 3-1, 3-2, 4, 5 IEC 63000:2016 1.0 - Technical documentation for the assessment of electrical and electronic products with respect to the restriction of hazardous substances EN 50419:2006 2.0 - Comply with Article 11(2) of WEEE Directive 2002/96/EC and is also applicable to comply with Article 15(2) of the recast Directive 2012/19/EU EN 301 489-1:2017 2.1.1 - ElectroMagnetic Compatibility (EMC) standard for radio equipment and services; Part 1: Common technical requirements EN 301 489-3:2017 2.1.1 - Electromagnetic compatibility and Radio spectrum Matters (ERM); ElectroMagnetic Compatibility (EMC) standard for radio equipment and services; Part 3: Specific conditions for Short-Range Devices (SRD) operating on frequencies between 9 kHz and 40 GHz EN 301 489-17:2017 3.2.0 - ElectroMagnetic Compatibility (EMC) standard for radio equipment and services; Part 17: Specific conditions for Broadband Data Transmission Systems; Harmonised Standard covering the essential requirements of article 3.1(b) of Directive 2014/53/EU EN 300 220-1:2017 3.1.1 - Short Range Devices (SRD) operating in the frequency range 25 MHz to 1 000 MHz;Part 1: Technical characteristics and methods of measurement EN 300 220-2:2017 3.1.1 - Short Range Devices (SRD) operating in the frequency range 25 MHz to 1 000 MHz; Part 2: Harmonised Standard covering the essential requirements of article 3.2 of Directive 2014/53/EU for non-specific radio equipment EN 300 440:2017 2.1.1 - Radio equipment to be used in the 1 GHz to 40 GHz frequency range; Harmonised Standard covering the essential requirements of article 3.2 of Directive 2014/53/EU EN 62479:2010 1.0 - Assessment of the compliance of low-power electronic and electrical equipment with the basic restrictions related to human exposure to electromagnetic fields (10 MHz to 300 GHz) EN 303 204:2016 2.1.2 - Network Based Short Range Devices (SRD); Radio equipment to be used in the 870 MHz to 876 MHz frequency range with power levels ranging up to 500 mW; Harmonised Standard covering the essential requirements of article 3.2 of the Directive 2014/53/EU EN 301 893:2017 2.1.1 - 5 GHz RLAN; Harmonised Standard covering the essential requirements of article 3.2 of Directive 2014/53/EU EN 300 328:2019 2.2.2 - Wideband transmission systems; Data transmission equipment operating in the 2,4 GHz band; Harmonised Standard for access to radio spectrum Note 1 These standards are only applicable in case that analytical tests of certain substances need to be made. Canada The table below contains the Regulations from Health Canada and Department of Justice that apply to medical electrical equipment and products emitting ionizing radiation. Interventional x-ray imaging systems must comply with these regulations in Canada. Table: Canada Regulations Date Title Statutory Orders and Regulations SOR/98-282 Last amended: 2022-03-02 Medical Devices Regulations Consolidated Regulations of Canada C.R.C. Chapter 1370 Last amended: 2021-03-31 Radiation Emitting Devices Regulations (Schedule II, Part XII, Diagnostic X-ray equipment) Canadian Deviations to International standards Table: Canadian deviation Edition Title CAN/CSA-C22.2 No. 60601-1:14 3.1 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance (IEC 60601-1:2005, third edition, 2005-12, including amendment 1:2012, modified) CAN/CSA-C22.1:2021 25 Canadian Electrical Code, Part I, Safety Standard for Electrical Installations. CAN/CSA C22.2 0.4-17 4.0 Bonding of Electrical Equipment USA The first table below contains the USA Federal Regulations that apply to medical electrical equipment and products emitting ionizing radiation and/or laser light. Interventional x-ray imaging systems must comply with these regulations in USA. The second table contains national deviations from international standards. Table: USA Federal Regulations Date Title Code of Federal Regulations, Title 21, Subchapter J, Part 1010 April 01, 2019 Performance standards for electronic products: general Code of Federal Regulations, Title 21, Subchapter J, Part 1020.30 April 01, 2019 Performance standards for ionizing radiation emitting products – Diagnostic X-ray systems and their major components Code of Federal Regulations, Title 21, Subchapter J, Part 1020.31 April 01, 2019 Performance standards for ionizing radiation emitting products – Radiographic equipment Code of Federal Regulations, Title 21, Subchapter J, Part 1020.32 April 01, 2019 Performance standards for ionizing radiation emitting products – Fluoroscopic equipment Code of Federal Regulations, Title 21, Subchapter J, Part 1040.10 April 01, 2019 Performance standards for light-emitting products – Laser products Code of Federal Regulations, Title 47 May 2017 Telecommunication (relevant for FCC certification of monitors) Table: USA deviations to international standards Edition Subject National Electrical Code (NFPA 70) 2020 Note 1 Protection against electrical fire and shock hazards ANSI/AAMI ES60601-1:2005/(R) 2012 and A1:2012, C1:2009/(R)2012 and A2:2010/(R)2012 3.1 Medical electrical equipment – Part 1: General requirements for basic safety and essential performance (modified) NEMA XR-27:2012 + Amendment A1:2013 1.1 X-ray equipment for interventional procedures User quality control mode NEMA XR-31:2016 1.0 Standard attributes on x-ray equipment for interventional procedures UL 2900-1:2017 1.0 Software Cybersecurity for Network-Connectable Products, Part 1: General Requirements Note 1 NEC applies to the installation of x-ray imaging systems in hospitals/clinics/etc and may contain additional requirements for the equipment. Some states may use previous versions of NEC (2005 or 2008). Latest version is considered to be state-of-the-art. Table: Philips’ proprietary information. Unauthorized use is prohibited. The energy consumption of the system shall be equal or less than for the predicate device (Azurion R2.1), when measured on a comparable configuration and under comparable usage. SAFETY, LEGAL & SECURITY STANDARDS Safety The system offers a method (to enable test automation) to: condition/configure/customize the system for specific tests. This includes software installation (“full install”). simulate human interaction and system environmental interaction to enable test automation. access/collect system logging and tracing. The system shall work in its AiT environment: No manual system modifications and content changes are needed for AiT Based on external information, the system shall offer automatic system configuration with only initial user interaction. USER INTERFACE REQUIREMENTS General Requirements
UNS.SmartSuite.2DQAAccuracyAndPrecision | Clinical images, Systematic error [mm], Random error [mm] Inter-observer variabilities obstructions, n.a., < 0.15 Diameter deviation from an equivalent marketed device, < 0.13, n.a. Quantify vessels and obstructions with 2DQCA with accuracy and precision in line with established 2DQCA guidelines. New, N.A. | SRS.Allura.Func.VesselAnalysis.Accuracy ¦ SRS.Allura.Func.VentricularAnalysisAccuracy | Table: Type of measurement Accuracy (systematic error) Precision (random error) QCA Vessel diameter < 0.2mm (for Ø <= 1mm) < 0.1mm (for Ø > 1mm) < 0.2mm QVA Vessel diameter < 0.2mm (for Ø <= 20mm) < 1% (for Ø > 20mm) < 0.2mm Vessel segment length < 1.0mm < 2.0mm Note: Vessel diameter accuracy is specified for measurements performed on a vessel placed in the iso-center, using automatic calibration. Vessel segment-length accuracy is specified for distances up to 50mm between user-defined markers on an un-foreshortened view of a vessel placed in the iso-center, using automatic calibration. Note: Accuracy and precision in table above are specified for a single measurement, which are to be distinguished from accuracy specifications for pooled measurements. Accuracy is decreased when using catheter calibration: Table: Catheter calibration Filled catheter >= 6 French Additional error 7% Note: Errors from using unfilled catheters or catheters below 6 French can be 20% or more. Ventricular Analysis (LVA & RVA) Using Auto Calibration, the volume methods (excluding the regression formulae) shall be accurate within 5% when applied to X-ray images of ellipsoid phantoms that represent average diastolic and systolic volumes in adults and children, respectively. Errors in Calibration Factors propagate into volume measurements multiplied by approximately a factor of 2 (biplane) to 3 (monoplane) and the IfU shall advise the users not to use catheter calibration for LVA and RVA. Ejection Fraction on the other hand is not sensitive for those inaccuracies Printing & Reporting This section focuses on case documenting and related hard-copy printing of images and reports. Print Compose
UNS.SmartSuite.2DQAAutomatedMeasurement | Have automatic contour detection in 2DQCA and 2DQVA of a user-selected vessel and/or stenosis in an angiogram. The user needs to be able to manually edit the provided vessel contours and obstruction borders and trigger a recalculation of the results; Automatically outline and estimate dimensions of user-indicated coronary segments (sizes 1-6 mm) and obstructions from non-subtracted angiograms (2DQCA); Automatically outline and estimate dimensions of user-indicated aortic and peripheral artery (sizes 5-50 mm) segments and obstructions from subtracted and non-subtracted angiograms (2DQVA). New, N.A. | SRS.Allura.Func.VesselAnalysis ¦ SRS.Allura.Func.QA ¦ SRS.Allura.Func.VentricularAnalysis | The system shall provide vessel analysis for coronary (QCA) and vascular arteries (QVA). The Quantitative Analysis shall allow the user to indicate a vessel segment in an un-subtracted or (for QVA only) in a subtracted X-ray image. The application shall then automatically detect the contrast-medium filled lumen in that segment and outline its borders as graphical overlays. The user shall be able to manually correct the contour. The application shall also display the profile of the vessel-diameter for the vessel segment and compute characteristics of the stenosis. QCA and QVA both require a separate license. The system shall provide Coronary Analysis, Vascular Analysis, Left Ventricular Analysis and Right Ventricular Analysis. Coronary Analysis (QCA) and Vascular Analysis (QVA) are monoplane applications i.e. analysis is performed on a single image of one channel only. Left and Right Ventricular Analysis (LVA, RVA) are available as mono- and biplane application i.e. in biplane configurations ventricular analysis can be performed on images of one or two channels. The user selects an image or run for analysis, using available image navigation functions. The user selects the desired analysis function, and performs the related applicable workflow. QA-calibration can be applied to express quantitative outcomes on an absolute scale (e.g. in mm). Results are shown on the review monitor but can also be copied to one of the reference viewports A QA report is generated and stored as a Secondary Capture image in the system database. This report can be displayed, exported or printed. Calibration The system shall provide ventricular analysis for the left (LVA) and right ventricle (RVA) After identification by the user of the aortic valve plane and the apex, the LVA application shall automatically trace and display the outline of a contrast-medium filled left ventricle. Biplane LVA shall similarly establish automatic contours for both the frontal and the lateral plane. RVA and LVA application shall allow the user to manually draw contours around the ventricular lumina in end-diastole and in end-systole, and in biplane series for the frontal and the lateral plane, respectively. RVA and LVA then computes ventricular characteristics like volumes, ejection fraction and performs wall motion analysis. LVA and RVA both require a separate license.
UNS.SmartSuite.2DQACalibration | Automatic and manual calibrate to convert distances measured in the X-ray image to real-world distances. New, N.A. | SRS.Allura.Func.QACalibration | All 2DQA applications and basic measurements (see SRS.Allura.Func.Measurement) shall support auto calibration and manual calibration methods to relate distances on display, expressed in pixels, to real-object distances. Biplane ventricular analysis shall support calibrations for both the frontal and for the lateral channel. Automatic calibration is based on geometry settings assuming that the object of interest is placed in the sytem’s iso-center. When Auto Calibration is active, the user shall be warned that the auto calibration is valid only if the ROI is imaged in the iso-center. Manual calibration is based on an object of known size which may be a non-tapered catheter (with or without contrast agent), a user indicated straight interval or a radio-opaque sphere Manual calibration methods shall allow the user to select and use a different series for calibration than for analysis. When 2DQVA is used in subtracted mode, the application shall automatically switch to and only display un-subtracted images in the calibration work-step Vessel Analysis (QCA & QVA)
UNS.SmartSuite.2DQAEfficientToUse | Receive automated assistance when execute time-consuming tasks like drawing a contour around a lumen or object; Obtain automatically generated reports from user-reviewed results. New, N.A. | SRS.Allura.Func.VesselAnalysis ¦ SRS.Allura.Func.QA ¦ SRS.Allura.Func.VentricularAnalysis | The system shall provide vessel analysis for coronary (QCA) and vascular arteries (QVA). The Quantitative Analysis shall allow the user to indicate a vessel segment in an un-subtracted or (for QVA only) in a subtracted X-ray image. The application shall then automatically detect the contrast-medium filled lumen in that segment and outline its borders as graphical overlays. The user shall be able to manually correct the contour. The application shall also display the profile of the vessel-diameter for the vessel segment and compute characteristics of the stenosis. QCA and QVA both require a separate license. The system shall provide Coronary Analysis, Vascular Analysis, Left Ventricular Analysis and Right Ventricular Analysis. Coronary Analysis (QCA) and Vascular Analysis (QVA) are monoplane applications i.e. analysis is performed on a single image of one channel only. Left and Right Ventricular Analysis (LVA, RVA) are available as mono- and biplane application i.e. in biplane configurations ventricular analysis can be performed on images of one or two channels. The user selects an image or run for analysis, using available image navigation functions. The user selects the desired analysis function, and performs the related applicable workflow. QA-calibration can be applied to express quantitative outcomes on an absolute scale (e.g. in mm). Results are shown on the review monitor but can also be copied to one of the reference viewports A QA report is generated and stored as a Secondary Capture image in the system database. This report can be displayed, exported or printed. Calibration The system shall provide ventricular analysis for the left (LVA) and right ventricle (RVA) After identification by the user of the aortic valve plane and the apex, the LVA application shall automatically trace and display the outline of a contrast-medium filled left ventricle. Biplane LVA shall similarly establish automatic contours for both the frontal and the lateral plane. RVA and LVA application shall allow the user to manually draw contours around the ventricular lumina in end-diastole and in end-systole, and in biplane series for the frontal and the lateral plane, respectively. RVA and LVA then computes ventricular characteristics like volumes, ejection fraction and performs wall motion analysis. LVA and RVA both require a separate license.
UNS.SmartSuite.3DAnatomyScan | Capture the 3D anatomy with a single scan. | SRS.Allura.Func.DRAFor3D ¦ SRS.Allura.Func.ClassicDRA ¦ SRS.Allura.Func.DualPhaseXperCT ¦ SRS.Allura.Func.XperCT | DRA for 3D is a special case of DRA performed at predefined positions, to enable 3D reconstruction. Settings, including start- and end-positions, shall be defined in EPX. The system shall support DRA for 3D at least at the following positions: for FlexArm (XY-fa) configuration at head, doctor and nurse side for other FD20 and FD15 configurations at head and one side position (doctor or nurse) for FD12 only at head position SID is maximized to allow for 3D scans of obese patients. The user shall define the end- and start-position of the DRA scan, implicitly executing a test-scan to detect obstructions. The system shall be able to perform the DRA scan at the side or head of the table (C-arm roll and propellor respectively). DRA acquisition is only possible for the frontal channel. The user shall be able to acquire a second run to enable subtracted reviewing. Settings, including start- and end-positions, shall be defined in EPX, but if allowed in EPX, the user shall be able to choose another end- and/or start-position. DRA requires a license. This is a special case of XperCT where two consecutive DRA scans are being made by executing the scans in different phases of contrast propagation, which is known as Dual phase XperCT. The two scans are performed in opposite direction to reduce the overall procedure time to minimize the amount of contrast medium. For neuro applications, to allow the use of intravenous contrast injections in Dual phase XPerCT scans, the system shall provide a subtracted fluoroscopy flavor with enhanced contrast. It shall enable to detect the arrival of contrast agent in order to start the scan. This application shall be executable from the Control Room. Flexible Dynamic Peripheral Angiography (FDPA) This is a special case of DRA or DAR for 3D enabling reconstruction of soft-tissue. This function is only supported with an FD20 detector. The system shall offer dedicated EPX procedures and related settings for Neuro (DRA and DAR) and Abdomen (DRA only) applications. The system shall perform acquisition with up to 60 images/second to obtain acceptable scan time. The system shall be able to apply slit collimation to reduce patient dose.
UNS.SmartSuite.3Dacquisition | Perform rotational acquisitions from Head and Nurse or Doctor side of the table in the same procedure; Perform 3D rotational acquisitions from Head, Nurse and Doctor side of the table in the same procedure with the FlexArm system and re-use these images for navigation purposes in 7 working positions (with respect to the standard table position): 1) Head side, 2) Nurse side, 3) Doctor side, 4) +45 degrees, 5) -45 degrees, 6) +135 degrees, and 7) -135 degrees. New, Perform 3D rotational acquisitions from Head, and Nurse and Doctor side of the table in the same procedure with the Biplane floor system (20”) and re-use these images for navigation purposes in 6 working positions (with respect to the standard table position): 1) Head side, 2) Nurse side, 3) Doctor side, 4) +45 degrees, 5) -45 degrees and 6) +135 degrees (nurse side); Perform 3D rotational acquisitions of the head from the head side of the table along the predefined trajectory. | SRS.Allura.Func.DRAFor3D ¦ SRS.Allura.Func.ClassicDRA ¦ SRS.Allura.Func.DualPhaseXperCT ¦ SRS.Allura.Func.XperCT | DRA for 3D is a special case of DRA performed at predefined positions, to enable 3D reconstruction. Settings, including start- and end-positions, shall be defined in EPX. The system shall support DRA for 3D at least at the following positions: for FlexArm (XY-fa) configuration at head, doctor and nurse side for other FD20 and FD15 configurations at head and one side position (doctor or nurse) for FD12 only at head position SID is maximized to allow for 3D scans of obese patients. The user shall define the end- and start-position of the DRA scan, implicitly executing a test-scan to detect obstructions. The system shall be able to perform the DRA scan at the side or head of the table (C-arm roll and propellor respectively). DRA acquisition is only possible for the frontal channel. The user shall be able to acquire a second run to enable subtracted reviewing. Settings, including start- and end-positions, shall be defined in EPX, but if allowed in EPX, the user shall be able to choose another end- and/or start-position. DRA requires a license. This is a special case of XperCT where two consecutive DRA scans are being made by executing the scans in different phases of contrast propagation, which is known as Dual phase XperCT. The two scans are performed in opposite direction to reduce the overall procedure time to minimize the amount of contrast medium. For neuro applications, to allow the use of intravenous contrast injections in Dual phase XPerCT scans, the system shall provide a subtracted fluoroscopy flavor with enhanced contrast. It shall enable to detect the arrival of contrast agent in order to start the scan. This application shall be executable from the Control Room. Flexible Dynamic Peripheral Angiography (FDPA) This is a special case of DRA or DAR for 3D enabling reconstruction of soft-tissue. This function is only supported with an FD20 detector. The system shall offer dedicated EPX procedures and related settings for Neuro (DRA and DAR) and Abdomen (DRA only) applications. The system shall perform acquisition with up to 60 images/second to obtain acceptable scan time. The system shall be able to apply slit collimation to reduce patient dose.
UNS.SmartSuite.AcquisitionImageQualityAndDose | Select, adjust and save image acquisition settings, for both fluoroscopy and exposures (incl. single shot): On the fly From various work spots (at table side, in the control room, etc.) For various items, such as: Patient Dose, Frame Rate Display real-time X-ray images; Display relevant (and regulatory required) acquisition parameters (e.g. kV, mA, DAP) in viewed images; Save relevant acquisition parameters in stored images/ image data. New, N.A. | SRS.Allura.Qual.IQGeoStability ¦ SRS.Allura.Qual.IQScatterReduction ¦ SRS.Allura.Conf.ImageStorageOptions ¦ SRS.Allura.Func.PresettingFunctions ¦ SRS.Allura.Func.DisplayFluoroTime ¦ SRS.Allura.Qual.IQWorkingPoint ¦ SRS.Allura.Func.High-DoseFluoroProtection ¦ SRS.Allura.Func.DisplayAndReportPatientDose ¦ SRS.Allura.Func.ExposureTechniques ¦ SRS.Allura.Qual.ClinicalResponseTimes.Acquisition ¦ SRS.Allura.Func.DoseControlTechniques ¦ SRS.Allura.Qual.X-rayPerformance ¦ SRS.Allura.Conf.UI-RelatedItems ¦ SRS.Allura.Qual.IQDisplaySignalTransfer ¦ SRS.Allura.Func.DoseControlMeasuringFields ¦ SRS.Allura.Func.RunLogging ¦ SRS.Allura.Qual.IQContrastResolution2D ¦ SRS.Allura.Func.FluoroFlavours ¦ SRS.Allura.Func.SelectDetectorFieldsize ¦ SRS.Allura.Qual.IQLimitingResolutionSystemLevel ¦ SRS.Allura.Qual.IQHalfValueLayer ¦ SRS.Allura.Func.AntiScatterGrid ¦ SRS.Allura.Qual.IQBeamAlignment ¦ SRS.Allura.Func.X-rayFiltering ¦ SRS.Allura.Func.ActivateFluoro ¦ SRS.Allura.Qual.IQX-rayOutput ¦ SRS.Allura.Func.ChangeFluoroProcessing ¦ SRS.Allura.Qual.IQBrightnessUniformity ¦ SRS.Allura.Func.TimedVariableFramerates ¦ SRS.Allura.Func.ActivateExposure ¦ SRS.Allura.Func.ManualOverrideEPXParameterValues ¦ SRS.Allura.Func.SingleShotFootpedal ¦ SRS.Allura.Func.IntegratedExposure ¦ SRS.Allura.Qual.IQmAsStabilized ¦ SRS.Allura.Qual.IQKVStabilized ¦ SRS.Allura.Qual.IQArtifactsSubjective ¦ SRS.Allura.Func.AutomaticElectronicShutters ¦ SRS.Allura.Func.DisplayInfo ¦ SRS.Allura.Func.ExposureFramespeed | Requirements for geometry stability in certain procedures. Table: Neuro subtract Trace subtract FDPA 3DRA Maximum image shift +/- 100 µm +/- 250 µm +/- 750 µm +/- 0.2° Requirements for the anti-scatter grid to reduce X-ray scatter (applies to all acquisition techniques). Table: Grid lines Primary Transmission Contrast Improvement SID application range Not visible in the final image ≥ 70 % 3.7 85-120 cm Table: sub tag 5122 1K2 (2K)2 .xs Capacity Keys 5 -1 Single + 100.000 25.000 6250 -2 Double6 o 200.000 50.000 12500 -3 Quadruple6 o 400.000 100.000 25000 .gc Frame-rate Keys 1 -1 Basic + 12 img/s 12 img/s 6 img/s -2 Cardiac6 o 30 img/s 3 30 img/s 3,4 6 img/s -8 Pediatric6 o 60 img/s 2 30 img/s 3,4 6 img/s Notes: These figures do not apply to fluoro or to scenarios and represent maximum frame-rates expressed in images/second (img/s) within limits set by detector and configuration. In biplane this is frame rate per channel, and in addition to 60 img/s also 50 img/s is supported. Cardiac and pediatric also supports 25 img/s instead of 30 img/s. With processing algorithm ClarityIQ (see SRS.Allura.Func.X-rayImageProcessing) only <= 25 img/s @ 1K2 is supported. The image sizes give an indication how many of these sized images can be stored, the system is not limited nor bound to these sizes as they vary depending on detector size and configuration. This option requires a license INTEGRATION & INTEROPERABILITY REQUIREMENTS This chapter focuses on integration with other systems/equipment, in and outside the exam room. UPS Integration Presetting is determined by the information in the study/procedure step that is current for acquisition. The user can modify the information. The procedure card to be used can be modified. The performing physician can be selected from a pre-defined physician list. The patient type can be entered/modified manually. The EPX application/procedure group/procedure can be selected from a list (which contains only items that can be performed within the system's current configuration). EPX application, procedure group and procedure can be selected in control room, as well as in exam room. EPX application, procedure group and procedure can be selected during x-ray. The selected settings will then be applied for the next exposure or fluoroscopy run. The total fluoro time for the current study is displayed. Requirements for working point, related to Automatic Gain Control (AGC), to ensure a reliable overall contrast transfer. Table: Acquisition mode AGC mode Accuracy must do [%] Fluoro/Cine/TSL/STI/DRA Active ± 2 Fluoro/Cine Fixed/Locked IQ-DoseAccuracy ± 2 TSL Locked ± 2 Single shot with ATC Active ± 5 STI/DRA Fixed IQ-DoseAccuracy ± 5 Customizable, depending on local legislation and/or preferences, the following measures are available: Limitation on entrance doserate. An audible warning signal, produced during the high-dose fluoro. Automatic switch-back towards normal dose settings after a certain time of continuously high-dose fluoro, and/or at the end of high-dose fluoro. This applies to Dose Area Product (DAP) and skin-entrance-dose (air-kerma). Dose info is displayed during and after acquisition. Generated and displayed values comply with legal demands for required accuracy. The user is warned if accumulated air-kerma within the current body zone is above a customizable threshold value (default value 2 Gray). The cumulative (zone-)dose figures are comprised in the study. Selection and settings of exposure dose-control techniques are determined in EPX and cannot be overridden by the user. Table: Categories/ items Max. notes: (all values in seconds) Switch viewing to fluoroscopy 1.8 from pedal press till the first image shown Switch viewing to exposure 1.8 from pedal press till the first image shown, For non-DSA cine Switch viewing to exposure 2.5 from pedal press till the first image shown, For acquisition types with test-shot lock-in and frame-speeds greater or equal to 4 fr/s Switch fluoro to viewing 1.8 from pedal release till display of viewing image on screen Switch exposure to viewing 1.8 from pedal release till display of viewing image on screen Switch between exposure/fluoro 1.8 For acquisition types without test-shot lock-in and frame-speeds greater or equal to 4 fr/s Switch between exposure/fluoro 2.5 For acquisition types with test-shot lock-in and frame-speeds greater or equal to 4 fr/s Digital Subtraction Angiography (DSA): delay after test-image 1.5 Change live fluoro settings 1.5 Fluo flavour, field-size (incl FD-mode). Change live exposure settings 0.5 Frame-rates for multiphase acquisitions Delay between Xray-OFF and Indicators-OFF 1.0 Fluoro pipeline perception latency 0.200 Essential for eye-hand coordination. X-ray dose-control techniques determine the applied X-ray dose, and have a major impact on resulting image quality. 'Cine control' uses previously acquired images and measurements to optimize X-ray settings. 'Test-Shot Lock-in' determines fixed settings for a whole run, based on testshot(s). 'Automatic Thickness Control (ATC)' uses a previous run to optimize start conditions for a next run. Control strategy as well as default settings are determined by EPX (optimized for e.g. application). Table: Categories/ items must notes Exposure with TC 100KW generator and MRC Continuous power 2.5 KW exposure and/or fluoro Peak power 100 KW exposure Tube current range 50-1000 mA exposure KV range 40-125 KV exposure Tube current range 10-200 mA pulsed Fluoro KV range 40-125 KV pulsed fluoro Clinical Response Times The performance figures are applicable under the following conditions: Normal clinical user mode. DICOM X-Ray Radiation Dose Structured Report active. Without (remote) field service activities. No degradation and exception situation like tube overheating or insufficient storage space. “historical system usage”: number of patients = 80 (-0+10%) number of runs per patient = 17 (-0+10%) number of images per run per channel = 100 (-0+10%) Lower amounts are allowed due to system configuration limitations (image disk space). Note on the must values that are mentioned in this and the following section: All the measured values shall be below 1.25 x must value. The average of the measured values shall be below the must value. The average is calculated over at least 3 values. In case of a user action, the response value is defined as the elapsed time between the user action and the related noticeable response. Table: sub tag item monoplane biplane .xs X-ray switches -1 footswitch 3f + -2 footswitch 6f + -3 wireless footswitch 3f o -4 wireless footswitch 6f o -5 additional footswitch CR 3f o -6 additional footswitch CR 6f o -7 additional footswitch ER 3f o -8 additional footswitch ER 6f o -9 X-ray handswitch + + Table: sub tag item Monoplane Monoplane Monoplane XY (fm/fa) XY (fm/fa) XY (fm/fa) biplane biplane biplane sub tag item AD7 XT AD7 XNT Maq AD7 XT AD7 XNT Maq Tru AD7 XT AD7 XNT Maq ui UI-modules -1 ControlModule NT +1) -2 ControlModule Tilt +1) -3 ControlModule Flex +1) +1) -4 ControlModule ORT o2) o2) +1,3) -5 ControlModule ORT Flex o2) o2) +1,3) -6 ControlModule Geo NT7) +1) -7 ControlModule Geo Tilt7) +1) -8 ControlModule Geo ORT7) o2) o2) + -9 ControlModule Geo NT IBR +1) -10 ControlModule Geo Tilt IBR +1) -11 ControlModule Imaging o4) o4) o4) o4) o4) o4) + + + -12 speed controller o5) o5) o5) o5) o5) o5) -13 pan handle o o o o o o -14 TouchScreenModule +6) +6) +6) +6) +6) +6) +6) +6) +6) -15 review module + + + + + + + + + -16 viewpad (cardio) # # # # # # # # # -17 viewpad (mixed) # # # # # # # # # -18 ControlModule Geo ORT IBR o2) o2) + Notes: 1) Additional ControlModule optional for CR 2) Optional as additional ControlModule on pedestal for ER 3) On pedestal 4) ControlModule imaging optionally available for ER (not in combination with 2)) 5) Required for FDPA, not an option for TruSystem 6) Additional TSM optional for CR and (on pedestal) for ER 7) Not for NPI, applicable for modules already present in the installed base Table: sub item Table type Table type Table type notes tag AD 7XT AD 7XNT Maquet TruSystem .mi miscellaneous -1 TSM swingarm + + + -2 Pedestal o o + trolley to attach UI modules; for AD7X(N)T including Pedestal Wall Conn. Box; for ORT including Surgery Wall Conn. Box -3 Two-way Intercom o o o -4 Keyboard + Mouse in CR + + + in CR; additional second keyboard + mouse in case of additional FlexSpot -5 Mouse in ER o o o Requires a license -7 MultiVision Switch o o o see SRS.Allura.Intgr.Multivision -8 KVM-Switch o o o see SRS.Allura.Intgr.KVM-Switching, requires DVI based video infrastructure -10 MCS Auxiliary box + + + for viewpad-receiver and X-on light on 3rd party booms Requirements for contrast transfer and stability as perceived on display (applies to all applications). Table: Pre condition: Calibrated monitor Monitor type and predefined contrast transfer and predefined working point Luminance L [Cd/m²] Luminance tolerance must [%] Luminance difference between displays in the ceiling suspension Luminance difference between viewing ports on large screen displays [%] L ± 10 D = 25 The field determines which part of the radiated detector-area is taken into account for dose-control. The field is determined in EPX; customizable from a set of available shapes. The user cannot override the measuring field. Information of stored fluoro and exposure runs is collected in the run log, as part of the study. This comprises: Figures on accumulated skin-dose and area-dose. Applied generator settings (e.g. KV/mA/mS) and X-ray spectral filter; per run. Applied geometry (e.g. rotation, angulation) settings; per run. Name of EPX procedure, time of acquisition, and number of images; per run. Fluoro Specific Functions Requirements related to contrast resolution in 2D applications (based on Leeds TO10 and 1.0 mm Cu). Table: Cardio Cardio Vascular Vascular Fluoroscopy Exposure Fluoroscopy Exposure Nr. visible details ≥ 70 ≥ 83 ≥ 74 ≥ 100 Dose 50 ± 10 nGy 150 ± 25 nGy 80 ± 15 nGy 4800 ± 800 nGy Potential 60 ± 2 kV 62 ± 2 kV 65 ± 2 kV 70 ± 2 kV The user can select between 3 EPX defined flavours, determining e.g. fluoro-frame-rate and X-ray dose. The flavour-determined frame-rate can be 30, 25, 20, 15, 12.5, 10, 7.5, 6.25, 5, 3.75, 3.125, 2.5, 1.875, 1.25, 1.0, 0.625, 0.5 img/sec. In biplane, frontal frame-rate equals lateral frame-rate (available range same as for monoplane). A fourth flavour can be selected by the system in situations where low-load is necessary (e.g. for graceful degradation). For legal reasons, a low-dose fluoro flavour is available with a ratio normal/low of at least a factor 2. Configurable, a high-dose flavour with a maximum of 20 R/min is offered. A flavour can also be selected during fluoro. The detector field-size is the detector area that is being acquired and displayed. The user shall be able to select a field-size; for fluoro also during acquisition. The available field-sizes and related image quality characteristics depend on the type of detector. The default shall be restored when a new procedure step is selected for acquisition, and EPX-dependent at EPX procedure selection. Requirements for sharpness in terms of Limiting Resolution System Level. Table: Min. [Cycles/mm]; (large focal spot) Min. [Cycles/mm]; (small focal spot) ≥ 1.7 ≥ 2.8 Half Value Layer (HVL) requirements to guarantee the patient protection against excessive skin dose. Table: Conditions: 100kV ± 1 kV Min [mm Al equivalent] Max [mm Al equivalent] HVL thickness 3.8 4.8 Legal requirement 3.6 n.a. Is placed in front of the detector to reduce the influence of scatter radiation on image quality. The grid can be mounted and removed by hand (without the need for tools). Requirements for alignment of tube and collimator at the detector. Table: Application conditions Must Standard position (rotation 0, angulation 0) + 5 mm Extreme position (rotation 90, angulation -45, min. SID) + 10 mm The system is equipped with a set of 4 X-ray spectral filters that are used to reduce the patient skin dose and to improve image quality. Selection of a filter for exposure is EPX procedure determined; no manual override. The filter for fluoro is coupled to the fluoro-flavour (defined in EPX application); no manual override. Image Generation This section describes the functions related to acquisition of X-ray images. General Concepts As long as the user activates fluoro, X-ray is generated and images are acquired and displayed. Exposure will temporarily interrupt fluoro; after the exposure switch is released, fluoro is automatically resumed (if still activated). After 10 minutes of uninterrupted fluoro, fluoro is automatically stopped. X-ray output requirements of the efficiency conversion for the X-ray generation. Table: Xray tube Min at delivery [µGy/mAs] Min during lifetime [µGy/mAs] MRC 200 1.75 1.57 Enables the user to change processing during LIH or Auto-cycle (e.g. contrast/brightness/edge-enhancement). The changes remain active until the same EPX application, another EPX application, patient type, performing physician or procedure step for acquisition is selected, or roadmap is activated. Exposure Related Functions Requirements for brightness non-uniformity (relative to the center). Table: Minimum Maximum Uniformity of brightness 60 % 120 % This enables the user to acquire a vascular/lock-in run in 3 phases with different frame-rates and durations. Pre-programmed sequences are available (customizable in EPX). The user can modify phase-duration and frame-rate per phase; phase-switch possible during the run This causes X-ray radiation, images are acquired, displayed and stored. The user is warned when the free space is insufficient to make the next exposure run. The run is stopped (in case of max. 1k2) after 1000 images, or when no more image storage free space is available. According to personal preferences the user shall be able to perform overrides on a subset of EPX values. The subset shall be restricted in order to avoid risks of a degraded performance or image quality. Unless explicitly specified otherwise changes shall not be possible during acquisition. Overriding settings at application level shall be persistent until the same EPX application, another EPX application, patient type, performing physician or procedure step for acquisition is selected. Overriding settings at procedure level shall be persistent until the same EPX application, another patient type, performing physician or procedure step for acquisition is selected. Customizing Functions This applies to customization of presettings according local or personal preferences. This enables the user to acquire a single image in vascular/lock-in mode directly via a separate footpedal. This feature provides improved image quality with reduced detector noise in vascular/lock-in runs. Settings are EPX determined and cannot be changed by the user. Dose-control requirements for test-shot lock-in, related to energy balance and X-ray penetration. Table: Conditions mAs stabilized Field-size kV stabilized [kV] Accuracy [%] 2 mm Cu + 20 mm Al and EPX with pre-defined stabilization All Pre-defined ± 30 Dose-control requirements for fluoro and cine-exposure, related to energy balance and X-ray penetration. Table: Conditions kV stabilized Field-size kV stabilized [kV] Accuracy [%] 2 mm Cu + 20 mm Al and EPX with pre-defined stabilization All Pre-defined ± 4 Requirements for artifacts as a collection of irregularities in the final image: The visibility of artifacts shall be within specified limits as recorded by reference images; this applies to: crosstalk, amplifier crosstalk, bright burn, first image gain, detector defects, anti-scatter grid imperfections (e.g. spots, lines), viewing jitter, focus stability, color impression, non-uniformity, viewing angle dependency. Unknown artifacts are judged by image quality and application specialists. The area outside the detected image is automatically blanked to hide detected scatter and noise. Displays detailed information about the current acquisition parameters. The display of information can be customized for live- and review viewport. Determines the rate at which images are acquired. In cardio/cine mode: 60, 50, 30, 25, 20, 15, 10, 7.5, 5, and 3.75 im/sec; EPX defined (within limits set by detector and configuration). In vascular Test-Shot Lock-in (TSL) mode: 12, 8, 6, 4, 3, 2, 1 and 0.5 im/sec; EPX-default, user-overridable (within limits set by detector and configuration). In biplane, frontal frame-rate equals lateral frame-rate (available range same as for monoplane).
UNS.SmartSuite.ActivateFluoroExposure | Activate fluoroscopy and exposure; Use a hand-switch for exposure; Use a wired footswitch for fluoroscopy/exposure; Use a wireless footswitch for fluoroscopy/exposure; Operate with a single action the acquisition system together with the injector; Save a fluoroscopy run retrospectively (store); Save a fluoroscopy run prospectively (grab). In case of biplane: Activate fluoroscopy from frontal and lateral channel separately as well as simultaneously Activate exposure from frontal and lateral channel separately as well as simultaneously Continue the examination with one channel if the other channel has an error. New, N.A. | SRS.Allura.Func.ContrastInjection ¦ SRS.Allura.Func.SelectActiveChannel ¦ SRS.Allura.Conf.UI-RelatedItems ¦ SRS.Allura.Func.ActivateExposure ¦ SRS.Allura.Func.SingleShotFootpedal ¦ SRS.Allura.Func.AcquisitionMethods ¦ SRS.Allura.Func.SetInjectorCoupling ¦ SRS.Allura.Func.ActivateFluoro ¦ SRS.Allura.Func.StoreFluoro | For fluoro, contrast medium is to be administered manually, and no synchronization or coupling to a motorized injector is provided. In a biplane system this enables the user to select the acquisition channel(s): frontal, lateral or biplane. Table: sub tag item monoplane biplane .xs X-ray switches -1 footswitch 3f + -2 footswitch 6f + -3 wireless footswitch 3f o -4 wireless footswitch 6f o -5 additional footswitch CR 3f o -6 additional footswitch CR 6f o -7 additional footswitch ER 3f o -8 additional footswitch ER 6f o -9 X-ray handswitch + + Table: sub tag item Monoplane Monoplane Monoplane XY (fm/fa) XY (fm/fa) XY (fm/fa) biplane biplane biplane sub tag item AD7 XT AD7 XNT Maq AD7 XT AD7 XNT Maq Tru AD7 XT AD7 XNT Maq ui UI-modules -1 ControlModule NT +1) -2 ControlModule Tilt +1) -3 ControlModule Flex +1) +1) -4 ControlModule ORT o2) o2) +1,3) -5 ControlModule ORT Flex o2) o2) +1,3) -6 ControlModule Geo NT7) +1) -7 ControlModule Geo Tilt7) +1) -8 ControlModule Geo ORT7) o2) o2) + -9 ControlModule Geo NT IBR +1) -10 ControlModule Geo Tilt IBR +1) -11 ControlModule Imaging o4) o4) o4) o4) o4) o4) + + + -12 speed controller o5) o5) o5) o5) o5) o5) -13 pan handle o o o o o o -14 TouchScreenModule +6) +6) +6) +6) +6) +6) +6) +6) +6) -15 review module + + + + + + + + + -16 viewpad (cardio) # # # # # # # # # -17 viewpad (mixed) # # # # # # # # # -18 ControlModule Geo ORT IBR o2) o2) + Notes: 1) Additional ControlModule optional for CR 2) Optional as additional ControlModule on pedestal for ER 3) On pedestal 4) ControlModule imaging optionally available for ER (not in combination with 2)) 5) Required for FDPA, not an option for TruSystem 6) Additional TSM optional for CR and (on pedestal) for ER 7) Not for NPI, applicable for modules already present in the installed base Table: sub item Table type Table type Table type notes tag AD 7XT AD 7XNT Maquet TruSystem .mi miscellaneous -1 TSM swingarm + + + -2 Pedestal o o + trolley to attach UI modules; for AD7X(N)T including Pedestal Wall Conn. Box; for ORT including Surgery Wall Conn. Box -3 Two-way Intercom o o o -4 Keyboard + Mouse in CR + + + in CR; additional second keyboard + mouse in case of additional FlexSpot -5 Mouse in ER o o o Requires a license -7 MultiVision Switch o o o see SRS.Allura.Intgr.Multivision -8 KVM-Switch o o o see SRS.Allura.Intgr.KVM-Switching, requires DVI based video infrastructure -10 MCS Auxiliary box + + + for viewpad-receiver and X-on light on 3rd party booms This causes X-ray radiation, images are acquired, displayed and stored. The user is warned when the free space is insufficient to make the next exposure run. The run is stopped (in case of max. 1k2) after 1000 images, or when no more image storage free space is available. This enables the user to acquire a single image in vascular/lock-in mode directly via a separate footpedal. Image generation (also called acquisition) can be done in a number of ways: Fluoro: intended to provide a real-time X-ray image. Exposure: intended for real-time diagnostic imaging and documentation purposes. Compared to fluoro a higher signal to noise ratio (S/N) is necessary for diagnostic imaging and the images are always stored, enabling reviewing. Enables to couple motorized injection to acquisition. This requires a license. The user can select between coupled and uncoupled. At a new procedure step selection for acquisition, and customizable after every run, the injector is set to uncoupled. Timing of injection is based on a fixed "injector response time" and EPX defined image after which injection should start and “Contrast Arrival Time” (user overridable). In case the start of acquisition is delayed, a “Count-down bar” is shown representing remaining time. Type of injector coupling is customizable: One-knob: X-ray hand- or footswitch control both injection and X-ray. Two-knob: X-ray hand- or footswitch controls X-ray, but injection is enabled by injector handswitch. Acquisition Of Physiology Signals As long as the user activates fluoro, X-ray is generated and images are acquired and displayed. Exposure will temporarily interrupt fluoro; after the exposure switch is released, fluoro is automatically resumed (if still activated). After 10 minutes of uninterrupted fluoro, fluoro is automatically stopped. Fluoro images are cached with cyclic overwrite for an EPX defined time-span. During fluoroscopy, the user can grab (i.e. store) single fluoro images or an image sequence. Multiple grabs during a single fluoro run are stored for each grab in its own run. After fluoroscopy, auto-cycle displays the cached sequence. After fluoroscopy, the user can store this complete cached run as a normal run. The cache is cleared at the start of a next fluoro run.
UNS.SmartSuite.AddPatientData | Import demographic patient/procedural information from the hospital information system (e.g. RIS/CIS); Manually enter patient data; Manually add missing information; Modify added information; Automatically push/pull information to/from other systems. Automatically trigger the corresponding procedural information (Procedure Card) when the system is connected to the hospital information system(e.g. RIS/CIS); Enter patient data in the Exam Room. New, N.A. | SRS.Allura.Intgr.NonDICOMWLMSystems ¦ SRS.Allura.Intgr.WLMSystems ¦ SRS.Allura.Func.ModifyStudy ¦ SRS.Allura.Intgr.Flexspot.AdditionalWorkspot ¦ SRS.Allura.Func.Merge ¦ SRS.Allura.Func.AddStudy ¦ SRS.Allura.Intgr.Flexvision ¦ SRS.Allura.Func.Workflow.Flexibility ¦ SRS.Allura.SLS.IHE ¦ SRS.Allura.Intgr.Flexvision.KeyboardMouse ¦ SRS.Allura.Intgr.Interfaces.DataExchange | RIS systems which are not using DICOM Worklist management can be connected via an external Philips product, which translates most RIS protocols to the DICOM protocol. Network Film Printers The system can be connected to the hospitals Worklist management system (e.g. RIS/CIS). This requires a license. Compatibility according DICOM conformance statement. Support for DICOM-SCU Worklist and MPPS (Modality Performed Procedure Step) protocol. This function enables to modify existing study data. The following data cannot be modified locally: Data originating from a WLM system (modify on WLM). Auto-stored data (e.g. run info); exception: the run description. A second workspot (keyboard, mouse, single monitor) next to FlexSpot shall be offered, which can be a workspot in the control room or exam room. One video source can be selected to be shown. Video can originate from the system itself and from auxiliary systems (configurable set, e.g. Interventional Workspot, IVUS). The additional FlexSpot requires a license. The user shall be able to move a run(s) from one procedure step to a procedure step of another patient. Also, the user shall be able to import a study from media to another patient in the local database. Selecting Study For Acquisition Before the start of acquisition the related study for acquisition is defined. In case a WLM system is connected, a study can consist of multiple scheduled procedure steps. In that case, the user selects a procedure step for acquisition. This function enables to add a new entry and related information to the study list. If operational, the WLM system provides the data for scheduling (e.g. patient name/ID, scheduled procedure start date/time). Alternatively, it remains possible to add a study entry locally. FlexVision enables flexible viewing for displaying multiple video sources (live images but also video from other systems) on one large screen display in the exam room with which the user can determine the layout, the position and size of multiple video sources on one large screen display in the exam room. Layouts are predefined. A layout divides the large screen in viewports. The user shall be able to select a preset (a preset contains the layout and for each viewport the specific video source shown in the viewport) to be applied on the large screen. The user shall be able to create, delete, change, save and rename a preset. The user shall be able to select which source is displayed in which viewport. The user shall be able to hide/unhide the video source display in a viewport. The user shall be able to swap the video sources of two viewports in a preset. The user shall be able to enlarge a specific viewport and restore it to its original size. The user shall be able to display video sources in native resolution. Display of 2k2 images shall be supported. Video can originate from the system itself and from auxiliary systems (configurable set, e.g. Interventional Workspot, IVUS). In combination with VoIP based infra, FlexVision requires a license. The display of 2K2 images requires an additional license. The user shall be able to have a workflow according local preferences. Different studies/procedure steps can overlap in time, i.e. The user shall be able to suspend the current acquisition study/procedure step, select another study/procedure step for acquisition and later select the suspended study/procedure step for acquisition again. The user shall be able to select a study for viewing besides an already selected study for acquisition. Sequence of workflow steps is not fixed, i.e. the user shall be able to switch between administration, image generation, viewing, post-processing, printing, reporting, archiving and distribution for the same study. Customizable automated workflow, i.e. Automatic DICOM export. Automatic dose report creation. Automatic DICOM dose structured report. Automatic finishing of the current acquisition study when another study is selected for acquisition. Automatic setting a study to protected against deletion upon finishing of the study. Automatic selecting of procedure card based on RIS code. Concurrency The system implements all of the transactions specified in the Integrating the Healthcare Enterprise (IHE) Technical Framework to support the IHE Integration Profiles, Actors and Options listed below: Table: Integration Profiles Implemented Actors Implemented Options Implemented Cardiac Catheterization Workflow (CATH) Acquisition Modality Broad Worklist Query Scheduled Workflow (SWF) Acquisition Modality Patient Based Worklist Query, Broad Worklist Query Patient Information Reconciliation (PIR) Acquisition Modality none Consistent Presentation of Images (CPI) Acquisition Modality none Radiation Exposure Monitoring (REM) Acquisition Modality none Consistent Time (CT) Time Client none Audit Trail and Node Authentication (ATNA) Secure Node none With FlexVision, the system shall allow for remote control of auxiliary systems via on-screen keyboard and touchpad (on TSM) and / or wireless mouse. This requires a license. The system shall provide communication interfaces to exchange data with compatible auxiliary systems Patient and examination related data Acquisition related data Dose data ECG data Geometry related data
UNS.SmartSuite.AdditionalFlexSpot | Occupy a second workspot (next to FlexSpot) in the CR or ER; Circulating nurse should be able to perform all tasks during a case (e.g. patient scheduling) without having to leave the ER. New, First and second FlexSpot with one or two 27’monitors(s) per workspot; Work in the same way has in the first FlexSpot; View and manipulate different sources on the primary and second workspot. | SRS.Allura.Intgr.Flexspot.AdditionalWorkspot ¦ SRS.Allura.Intgr.Flexspot.SecondFlexspot ¦ SRS.Allura.Conf.Control-RoomDisplays | A second workspot (keyboard, mouse, single monitor) next to FlexSpot shall be offered, which can be a workspot in the control room or exam room. One video source can be selected to be shown. Video can originate from the system itself and from auxiliary systems (configurable set, e.g. Interventional Workspot, IVUS). The additional FlexSpot requires a license. A second workspot (keyboard, mouse, one or two monitors) next to FlexSpot shall be offered, which can be a workspot in the control room or exam room. It shall provide the user with the same features as FlexSpot. Second FlexSpot requires a license. A second monitor for the second workspot shall be offered. This second monitor requires a license. Table: sub tag item DVI based infra VoIP based infra notes . crd control room displays -1 acquisition and review monitor # # Both 24" color monitors with native image resolution of 1920x1080 and aspect ratio 16:9. -2 FlexSpot primary monitor # # Instead of acquisition and review monitor, a 27" color monitor with native image resolution of 2560x1440 and aspect ratio 16:9. -3 FlexSpot secondary monitor o o Requires FlexSpot primary monitor, 27" color monitor with native image resolution of 2560x1440 and aspect ratio 16:9. -4 additional FlexSpot o Requires FlexSpot primary monitor, 27" color monitor with native image resolution of 2560x1440 and aspect ratio 16:9; shall be placed either in the control or exam room. -5 slave monitor1) o o 24” (desk version) or 27” (boom version) or 32" (boom version) color monitor with native image resolution of 1920x1080 and aspect ratio 16:9, e.g. extra live monitor. -6 MultiView o Requires acquisition and review 24" color FHD monitors, -7 Additional MultiView o Requires acquisition and review monitor, extra 24" color FHD monitor, mouse and keyboard; -8 Second FlexSpot primary monitor o Requires FlexSpot primary monitor, 27" QHD color monitor, mouse and keyboard; shall be placed either in the control or exam room. -9 Second FlexSpot secondary monitor o Requires Second FlexSpot primary monitor, 27" color QHD monitor; shall be placed at the same location as Second FlexSpot primary monitor -10 3rd MultiView o Requires extra 27" color QHD monitor, mouse and keyboard. Only in combination with FlexSpot, second FlexSpot, and 4th MultiView -11 4th MultiView o Requires extra 27" color QHD monitor, mouse and keyboard. Only in combination with 3rd MulitView Note: 1) 32" monitors are not part of NPI.
UNS.SmartSuite.Administration | Do administration tasks and store evidence* after any service activity to indicate that the system is operational according to manufacturer specification. *, e.g. import to the Performance Assurance tool and installed base registration. | SRS.Allura.Serv.Backup/Restore ¦ SRS.Allura.Serv.SystemID ¦ SRS.Allura.Serv.SiteLog | It shall be possible to backup system data and if needed restore, e.g. in case of a disk failure. Includes all customization and EPX settings (except manually added image processing tastes) and all system-specific adjustment- and configuration data. Backup data can only be restored on the same system on which it was created Backup data can be restored only if it was created with the same software version or with a software version for which an upgrade path is defined (see Upgradeability Requirements) The system shall be uniquely identified. A service report shall contain the system ID and hospital name. Non-Functional Requirements Upgrading Related The FSE shall be able to use an on-system logbook to register the systems history and comments by the FSE.
UNS.SmartSuite.AmbientVideos | Display ComfortThemes videos on FlexVision controlled from TSM. New, N.A. | SRS.Allura.UI.ComfortThemes | With FlexVision, the user shall be able to play movies or image slideshows to relax the patient. This requires a license.
UNS.SmartSuite.Annotations | Define/set annotations per image and run; Select predefined standard annotations; Reposition and modify annotations; Control the list of annotations; Add free text to images/image selection during or shortly after the examination. Edit run name. Define/set annotations on X-ray images, screenshots and reconstruction images. New, N.A. | SRS.Allura.Func.SelectProcessingFocus ¦ SRS.Allura.Func.PredefinedAnnotation ¦ SRS.Allura.Func.ModifyStudy ¦ SRS.Allura.Func.SelectProcessingScope ¦ SRS.Allura.Func.ManualAnnotation ¦ SRS.Allura.Func.OverrideManualAdjustments | This function allows the user to select whether the processing applies to the images of both channels (in case of biplane runs) or only one. During manual annotation, a list can be displayed containing predefined and user defined text strings, which can be inserted within the current annotation block. The user shall be able to customize this list: The predefined annotations can be deleted and modified. The user defined annotations can be added, deleted and modified. This function enables to modify existing study data. The following data cannot be modified locally: Data originating from a WLM system (modify on WLM). Auto-stored data (e.g. run info); exception: the run description. The user shall be able to set the scope of the interactive post-processing operations within a run: the scope may be set to either the whole run or to the active image for annotation; the scope may be set to either the whole run, to a selection of images within the run, or to the active image for image enhancement, video invert, pan, zoom, and landmarking; the scope may be set to either the whole run, to a selection of images within the run, the active image, active image and all preceding images in the run, active image and all following images in the run for pixel-shift; for all other operations the scope is always the whole run. Image Enhancement This enables the user to add text and/or shapes on a certain position into the image on display. Annotations can be repositioned, modified, hidden, copied and deleted. Annotations can be added to x-ray images and as well on secondary capture images. This function resets settings (if altered during postprocessing) to “as during acquisition”. The user can select the scope of this function to all postprocessing settings, or a defined subset.
UNS.SmartSuite.AuthorizedUser | Use the system with or without authorization log on; Limit/restrict the use of the system by users; In case of an emergency, unblock the use of primary system functionality by any user. New, N.A. | SRS.Security.MD.PhysicalAccess ¦ SRS.Security.PHD.PhysicalAccess | The Azurion shall be protected against attacks via physical access. The Patient Health Data shall be protected against attacks via physical access.
UNS.SmartSuite.AutomaticPositionControl | Store and recall one table position; Automatically re-acquire the exact same projection as the reference image currently displayed; Switch between defined projections, mostly of different views of the same body part; Customize and store pre-defined projections and sequences as part of the EPX, but without going to the EPX level; Store temporally and recall up to 20 projections of the system during a procedure (APC Store/Recall); Move the geometry to predefined projections (patient oriented) and along a sequence of projections defined at EPX level (APC Sequence); Automatically set back the position of the system (beam carrier(s) and table) to the same position that was used during live, acquisition, reference images and/or manual storage, including the detector orientation, FOV, shutters and wedge positions; Move the position of the system (rotation and angulation) to the patient oriented projection defined in the interventional tools; Addition to APC with an APC along a path of positions of the longitudinal, transversal and z-rotation axes and of the table for the FlexArm configuration. New, N.A. | SRS.Allura.Func.APCReference ¦ SRS.Allura.Func.ModifyEPXParameters ¦ SRS.Allura.Func.APCTable ¦ SRS.Allura.Func.SystemCustomization ¦ SRS.Allura.Func.APCPredefined ¦ SRS.Allura.Func.APCPathway ¦ SRS.Allura.Func.APCStore/Recall ¦ SRS.Allura.Func.APCCommon ¦ SRS.Allura.Func.APC3DReference | The user shall be able to recall the geometry positions (i.e. position of frontal and lateral stand, table, shutters and wedges) plus the receptor field size which corresponds to an image that is displayed in a live or reference viewport. In a biplane system, the user shall be able to recall a position for both frontal and lateral channels independently or simultaneously. Note: Table position recall is not supported for TruSystem tables in combination with FlexMove. On AD7X tables recall of longitudinal and lateral table position requires appropriate motors which are optional APC reference requires a license. Focuses on changing the value of EPX parameters. Complex settings are protected against modification by insufficiently skilled engineers. (setting-specific authorization level versus service-key defined skill level). The scope of this functionality can be system-wide, and/or physician-specific. Patient And Beam Positioning This section deals with positioning of patient and/or beam in order to select the anatomical area of interest. Orientation Aspects After the user has brought a point-of-interest into the system’s iso-center, the user shall be able to store the corresponding table position (i.e. table longitudinal, lateral, height, tilt and cradle). The system shall provide a default point-of-interest (EPX defined). When a new acquisition study is selected the 'stored table position' is reset such that the default point-of-interest is in the iso-center. The user shall be able to recall the stored table position or only the table height. The system shall indicate to the user if the table is in the stored position. Note: Not applicable for Hillstrom TruSystem table with universal table top and Maquet Magnus table with universal table top. On TruSystem tables with the SQ14-Xtra table top only height can be recalled. On AD7X tables recall of longitudinal and lateral table position requires appropriate motors, which are optional. APC table requires a license. Allows the user to customize system settings, that are available for the system's configuration, such as: Date, Time and Regional Settings User Administration and Administration settings List of annotation and performing physician Procedure cards and RIS code mapping Settings for viewing and quantitative analysis Media and DICOM export settings APC positions Linking DICOM export settings, APC positions and procedure help text to EPX procedures Remote support settings Defining APC pathways Saved settings are persistent over system restarts. System customization settings can be imported and exported. Administration Functions This section focuses on administrative functions to deal with patient data and related information. General Administration Concepts A local database includes all information that is entered manually and generated in acquisition. It can contain multiple patients/studies/procedure steps comprising patient data, images, etc. Local Administration Functions The system shall provide pre-defined sets each comprising a sequence of projections (defined by rotation/angulation, SID and Field of View). The sets shall be EPX defined and a set can have max. 20 projections. The user shall be able to customize the sets. Specific sets can be customized for different users. The user shall be able to recall the projections of a set in sequential order, but alternatively the user shall be able to recall a projection directly. APC predefined requires a license. Only available on the FlexArm configuration. A user shall be able to define a pathway of longitudinal, transversal and z-rotation stand positions for the system. For tables with recall capabilities, also a table position can be given for the first and last positions. A user shall be able to recall the pathway, causing the stand and table to be moved to the positions of the pathway in the sequence defined for the pathway. Beam Limitation & Conditioning This section focuses on shaping and filtering the X-ray beam. Beam Shape Settings To restrict X-ray radiation to the clinical region of interest the user can apply shutters. Additionally, the user can apply wedges that reduce X-ray intensity to prevent over-exposure in parts of the image. The system shall allow the user to store and recall multiple geometry positions (i.e. position of stand, table, shutters and wedges) including the receptor field size. When a new acquisition study is selected, all stored positions shall be removed. The system shall indicate to the user if the system has not reached a selected position. Note: Table position recall is not supported for TruSystem tables in combination with FlexMove. On AD7X tables recall of longitudinal and lateral table position requires appropriate motors which are optional During APC movements the SID shall be maximized in order to reduce the chance of a collision. In case of a collision APC movement is stopped. The user shall be able to recall the frontal projection (rotation/angulation) that corresponds to a view on the Interventional Workspot (3DRA, XperCT, etc). Indicated by the Interventional Workspot, the SID shall go to minimum or to maximum (maximum to e.g. prevent hitting the needle) during the recall movement. APC 3D reference requires a license.
UNS.SmartSuite.AuxiliarySystems | Integrate and/or connect major brands of peripheral equipment and information systems; Connect the viewing system to a wide variety of third party systems (e.g. surgical knife, injector, video screens); View and interact with intra-vascular ultrasound (IVUS) images from table side; View and analyse images and volumes, originating from auxiliary systems, together with the images generated by the system; Transfer images and volumes, created by the system, to the auxiliary systems, as input for additional processing by these auxiliary systems; Position catheters, as well as display and analyse catheter measurements; Acquire and display physiological signals; Use the system in a Hybrid OR environment by: combining a special hybrid OR patient table with the system creating a clean ceiling layout with the ability to position the equipment; Couple an injector to align (timing) the contrast injection with the runs; Control the injection of contrast fluid. New, Have flexibility in where to connect external equipment (e.g. different corners in the room, on the MCS); Align (timing) the contrast injection with the runs. | SRS.Allura.Conf.XperFDSeriesConfigurations ¦ SRS.Allura.Intgr.FHDVideoSources ¦ SRS.Allura.Conf.RemainingComponents ¦ SRS.Allura.Intgr.OpenVideoInterface.LargeScreenDownscaled ¦ SRS.Allura.Intgr.DoseAwareXtend ¦ SRS.Allura.Intgr.Interfaces.ControlAuxiliarySystems ¦ SRS.Allura.Intgr.Multivision ¦ SRS.Allura.Intgr.Interfaces.RealTimeImageLink ¦ SRS.Allura.Func.PhysioAcquisition ¦ SRS.Allura.Intgr.Magnus ¦ SRS.Allura.Intgr.ImageExportImport ¦ SRS.Allura.Conf.UI-RelatedItems ¦ SRS.Allura.Intgr.TruSystem ¦ SRS.Allura.Intgr.OpenVideoInterface.LargeScreen ¦ SRS.Allura.Intgr.Flexspot.AdditionalWorkspot ¦ SRS.Allura.Intgr.Flexspot ¦ SRS.Allura.Intgr.Interfaces.AuxiliarySystemMounting ¦ SRS.Allura.Func.SetInjectorCoupling ¦ SRS.Allura.Intgr.Flexvision.KeyboardMouse ¦ SRS.Allura.Intgr.4KVideoInputOutput ¦ SRS.Allura.Intgr.VideoSlave ¦ SRS.Allura.Conf.Table-RelatedItems ¦ SRS.Allura.Intgr.Interfaces.ControlFromAuxiliarySystems ¦ SRS.Allura.Conf.MCS-RelatedItems ¦ SRS.Allura.Intgr.SwitchableMonitors ¦ SRS.Allura.Intgr.PhysioEquipment(ECG) ¦ SRS.Allura.Intgr.Flexvision ¦ SRS.Allura.Intgr.AblationEquipment ¦ SRS.Allura.Intgr.Flexspot.KeyboardMouse ¦ SRS.Allura.Intgr.Injectors ¦ SRS.Allura.Intgr.EquipmentRack ¦ SRS.Allura.Intgr.KVM-Switching ¦ SRS.Allura.Intgr.VideoDisplayInterface ¦ SRS.Allura.Intgr.Interfaces.DataExchange | Table: sub tag Systems: FD12 FD12/12 FD12 FD12/12 FD12 FD12/12 FD20 FD20 FD20 FD20 FD 15 FD20/12 FD20/15 FD20/12 FD20/15 item Ce Fl Bpl Ce XY fa XY fm Fl Fl Bpl NB notes .std Geo Stand -1 Poly-G2-C # Poly-G2-C is supported for remanufacturing -2 Poly-G3-C # -3 Poly-G2-F + + -4 Clea-2-C + + Clea-2-C is supported for remanufacturing (not XY fm) -5 Clea-3-C + -6 Clea-F # # # # -7 Clea-3-F # Only in combination with AD7 table -8 Larc-N 5) # -9 Larc-C # # # -10 Larc-3-N # Only in combination with AD7 table .pt Patient Table -1 AD7XT # # o o o o # # o o -2 AD7XNT # # o o o # # o o -3 Maquet Magnus c c c c c c -4 Hillrom TruSystem c c See SRS.Allura.Intgr.TruSystem .ms Monitor Ceiling Support -1 Rail suspended MCS % % % % % % % % % -2 Spring-arm boom % % % % % % % % % % -3 Spring-arm boom XL % % % % % % % % % % -4 3rd party boom % % % % % % % % % % MCS Monitor integration kit -5 3rd party boom XL % % % % % % % % % % MCS Monitor integration kit XL .xrc X-ray chain ++= frontal & lateral -1 Tube MRC 200+ 0508 + + ++ +1) +1) -2 Tube MRC 200+ 04072) + + + + + +1) +1) -3 Collimator Nicol3) + + ++ + + + + + ++ ++ -5 AEP dose-meter o o o o o o o o o o AEP calculated by default, optionally hardware based -7 Detector FD124) + + ++ # # -8 Detector FD15 + # # -9 Detector FD20 + + + + # # Notes: 1) Frontal channel MRC0407; lateral channel MRC0508 2) MRC 200+ 0407 ROT-GS1008 for FD20 XY-fa; ROT-GS1004 for all other configurations 3) Nicol v4 on Clea-3, L-arc 3-N and Poly-G3 Ceiling. All other stands Nicol v3 4) Monoplane configurations PX2121S or PX2121C/V; biplane configurations PX2121C/V only 5) Azurion FD20/12 Clea-Floor with FD20 & Larc-N with FD12 is not supported for NPI. This configuration is only supported for the installed base In combination with IP based video integration infrastructure, FlexVision exam room display configuration and FlexSpot + 2nd FlexSpot, or MultiView + Additional MultiView (i.e. no other CR/ER displays), the system shall support connecting up to 20 FHD auxiliary systems and up to 4 integrated workstations. In combination with DVI based video infrastructure the system shall support up to 16 FHD video sources, of which up to 11 video sources (total number from the integrated workstations and auxiliary systems). The system shall offer the flexibility to connect auxiliary systems in the control room, the examination room (on MSCs, walls or equipment racks) and the technical room. For VoIP configurations, the support of FHD video sources is licensed. This is a counted license for the number of video sources. The license does not actually limit the number of FHD video sources that can be connected. Table: sub tag item availability notes .ps Partner Systems See for details corresponding requirement in Room-Specific Auxiliary Systems section -1 DoseAware Xtend o .mi Miscellaneous -1 Ceiling Suspended radiation shield o -2 Examination light o -3 X-ray spacer o limits minimum Source Skin Distance (Health and Human Services (HHS) requirement). Mandatory for 21 CFR 1020.30. -4 Drip stand o -5 Medical DVD recorder o Requires DVI based video infrastructure -6 1 phase UPS o see SRS.Allura.Func.PowerFailureBehaviour -9 UPS control c see SRS.Allura.Intgr.ExternalThreePhaseUPS -10 VideoSlave o see SRS.Allura.Intgr.VideoSlave -11 Open video interface o see SRS.Allura.Intgr.OpenVideoInterface.LargeScreen, -12 DVD burner o see SRS.Allura.Func.StoreToMedia -13 HD output o see SRS.Allura.Intgr.OpenVideoInterface.LargeScreenDownscaled; in DVI based video infrastructure it requires FlexVision open video interface -14 Video scan converter o for 50/60 Hz SLR TV output, requires DVI based video infrastructure -15 Wall-connection boxes o see SRS.Allura.Intgr.FHDVideoSources, requires DVI based video infrastructure -16 Auxiliary System Input Connection Box o see SRS.Allura.Intgr.FHDVideoSources, requires IP based video infrastructure -17 4K Auxiliary System Input Connection Box o see SRS.Allura.Intgr.4KVideoInputOutput, not available in combination with DVI based video infrastructure Core Functionality Core functionality is defined as functionality which is not dependent on the presence of any license key. The following functions are in the core: Clinical core functionality is a subset of all functionality as defined in chapter FUNCTIONAL REQUIREMENTS. The subset is determined by subtracting all functionality that is dependent on any license key. All functionality needed to install a system and to perform basic checks to initially verify correct system working condition. User messages (see SRS.Allura.Func.DisplayUserMessages). Tooltip (balloon) help. Software-Keyed Clinical Options A Clinical Option gives additional clinical functions/features/characteristics and is enabled through a software license key. Two types of licenses exist: counted license: used to implement licenses that restrict the number of items that can be used, e.g. number of video outputs. uncounted license: used to implement licenses that restrict the usage of a certain feature (on/off) The system shall provide a downscaled copy of the large screen of FlexVision video output in FHD resolution (DP is the interface in combination with IP based video infrastructure and SL-DVI in combination with DVI based video infrastructure). In combination with IP based video infrastructure that video output shall be also available (configurable) for switching on Flexspot, (Additional,3rd, and 4th) Multiview and Switchable Monitors. In VoIP infrastructure, the availability of the downscaled copy requires a license. The system shall optionally provide integration with the DoseAware Family product, which means: In combination with DoseAware Xtend, video can be shown on the Azurion display. Azurion provides X-ray procedure related information to DoseAware Xtend. A reference dosimeter can be attached to the frontal stand. In combination with DoseAware the system shall provide electrical and mechanical integration of the DoseAware basestation on the MCS frame. The system shall provide communication interfaces to control compatible auxiliary systems from the systems user interface with wireless mouse on the auxiliary system’s graphical UI displayed on FlexVision with mouse and keyboard on the auxiliary system’s graphical UI displayed on Flexspot from TSM by a generic button based UI tailored by the auxiliary system or by a graphical UI streamed from the auxiliary system including (multi-)touch control. This requires a license control shall be established automatically when (mobile) compatible auxiliary system is connected MultiVision enables to display video on the exam room display originating from multiple video sources. The user can select the current video source for display. Video can originate from the system itself and/or from auxiliary systems (e.g. Interventional Workspot, IVUS). Not available in combination with Switchable Monitors or FlexVision. In a DVI configuration, the control room review monitor is selectable as video source and max. 4 other video sources are configurable. MultiVision requires a license The system shall provide communication interfaces to provide x-ray images of the frontal channel in real time to compatible auxiliary equipment during x-ray acquisition. This requires a license. The system can acquire physiological signals together with the X-ray images. This requires a license. Number of channels (up to 4) is configurable. Only compatible with electrical signals (e.g. ECG; no compatibility with pressure, etc.). EPX determined storage (on/off) of all inputs; recording only in parallel with exposure. Image Generation Scenarios This section focuses on acquisition of images during a controlled geometry movement. Dynamic Rotational Angiography In a DRA scan the X-ray beam rotates around the patient's area of interest placed within the iso center. Compatibility with Magnus OR table of Maquet. Move table via system’s UI modules and Magnus UI modules to position table top, tilt, cradle and set height. Move table for reset geometry, APC iso-center and FDPA. Stop table movements in case of emergency stop. Detect collisions and allow override table movements. Synchronize table lock setting, patient orientation setting, table positions, status and capabilities. The system can be connected to departmental image storage providers (e.g. Xcelera PACS) and/or storage users. Compatibility according to DICOM conformance statement. Support for Query (DICOM SCU), Export (DICOM SCU) and Import (DICOM SCP). Supported formats: XA-multi-frame, Secondary Capture, JPEG-lossless. Supported features: Grayscale and non Grayscale Standard Display Compatible, DICOM Presentation state, Storage Commit. Departmental Worklist Management (WLM) Systems Table: sub tag item monoplane biplane .xs X-ray switches -1 footswitch 3f + -2 footswitch 6f + -3 wireless footswitch 3f o -4 wireless footswitch 6f o -5 additional footswitch CR 3f o -6 additional footswitch CR 6f o -7 additional footswitch ER 3f o -8 additional footswitch ER 6f o -9 X-ray handswitch + + Table: sub tag item Monoplane Monoplane Monoplane XY (fm/fa) XY (fm/fa) XY (fm/fa) biplane biplane biplane sub tag item AD7 XT AD7 XNT Maq AD7 XT AD7 XNT Maq Tru AD7 XT AD7 XNT Maq ui UI-modules -1 ControlModule NT +1) -2 ControlModule Tilt +1) -3 ControlModule Flex +1) +1) -4 ControlModule ORT o2) o2) +1,3) -5 ControlModule ORT Flex o2) o2) +1,3) -6 ControlModule Geo NT7) +1) -7 ControlModule Geo Tilt7) +1) -8 ControlModule Geo ORT7) o2) o2) + -9 ControlModule Geo NT IBR +1) -10 ControlModule Geo Tilt IBR +1) -11 ControlModule Imaging o4) o4) o4) o4) o4) o4) + + + -12 speed controller o5) o5) o5) o5) o5) o5) -13 pan handle o o o o o o -14 TouchScreenModule +6) +6) +6) +6) +6) +6) +6) +6) +6) -15 review module + + + + + + + + + -16 viewpad (cardio) # # # # # # # # # -17 viewpad (mixed) # # # # # # # # # -18 ControlModule Geo ORT IBR o2) o2) + Notes: 1) Additional ControlModule optional for CR 2) Optional as additional ControlModule on pedestal for ER 3) On pedestal 4) ControlModule imaging optionally available for ER (not in combination with 2)) 5) Required for FDPA, not an option for TruSystem 6) Additional TSM optional for CR and (on pedestal) for ER 7) Not for NPI, applicable for modules already present in the installed base Table: sub item Table type Table type Table type notes tag AD 7XT AD 7XNT Maquet TruSystem .mi miscellaneous -1 TSM swingarm + + + -2 Pedestal o o + trolley to attach UI modules; for AD7X(N)T including Pedestal Wall Conn. Box; for ORT including Surgery Wall Conn. Box -3 Two-way Intercom o o o -4 Keyboard + Mouse in CR + + + in CR; additional second keyboard + mouse in case of additional FlexSpot -5 Mouse in ER o o o Requires a license -7 MultiVision Switch o o o see SRS.Allura.Intgr.Multivision -8 KVM-Switch o o o see SRS.Allura.Intgr.KVM-Switching, requires DVI based video infrastructure -10 MCS Auxiliary box + + + for viewpad-receiver and X-on light on 3rd party booms Compatibility with TruSystem 7500 OR table of Hillrom In combination with the FlexArm: Move table via system's UI modules and TruSystem UI modules to position table top, tilt, cradle and set height. Move table for reset geometry, APC iso-center. Stop table movements in case of emergency stop. Detect collisions and allow override table movements. Synchronize table lock settings, patient orientation setting, table positions, status and capabilities. Excluded are: FDPA, iso-centric tilt, and Syncra tilt. In combination with the FlexMove there is limited integration: Table can be locked by the system The system's emergency stop also stops all active table movements The system shall provide a large screen video output (DP1.2 or, in combination with DVI based video infrastructure, 2*DL-DVI) that contains a copy of the large screen of FlexVision. In VoIP infrastructure, this requires a license. A second workspot (keyboard, mouse, single monitor) next to FlexSpot shall be offered, which can be a workspot in the control room or exam room. One video source can be selected to be shown. Video can originate from the system itself and from auxiliary systems (configurable set, e.g. Interventional Workspot, IVUS). The additional FlexSpot requires a license. Provides the user with the same features as for FlexVision in the control room. Exception is that FlexSpot doesn't support display of 2k2 images. In combination with VoIP based infrastructure, FlexSpot requires a license. A second monitor for the FlexSpot workspot shall be offered. This second monitor requires a license. The system shall provide mechanical interfaces and cabling to mount external equipment to: Patient Table Monitor Ceiling Suspension Monitor Ceiling Carriage Video Integration Enables to couple motorized injection to acquisition. This requires a license. The user can select between coupled and uncoupled. At a new procedure step selection for acquisition, and customizable after every run, the injector is set to uncoupled. Timing of injection is based on a fixed "injector response time" and EPX defined image after which injection should start and “Contrast Arrival Time” (user overridable). In case the start of acquisition is delayed, a “Count-down bar” is shown representing remaining time. Type of injector coupling is customizable: One-knob: X-ray hand- or footswitch control both injection and X-ray. Two-knob: X-ray hand- or footswitch controls X-ray, but injection is enabled by injector handswitch. Acquisition Of Physiology Signals With FlexVision, the system shall allow for remote control of auxiliary systems via on-screen keyboard and touchpad (on TSM) and / or wireless mouse. This requires a license. It shall be possible to connect to the system (connection available on MCS or wall) QHD, UHD or 4K video sources (max 2) and output it: as downscaled FHD video source to the system monitor or external video system optionally directly (no switching possibility for a clinical user) to a 4K 3rd party monitor(s) (mounted on a 3rd party monitor boom or installed in the control room) Connecting one 4K video source and displaying its output on the 4K 3rd party monitor shall not impact the total number of FHD video sources of the auxiliary systems that can be connected to the system by more than 10. Connecting and outputting 4K external inputs are both licensed. These are counted licenses for the number of 4K auxiliary system input connection boxes and 4K outputs installed. The licenses are not enforced. VideoSlave enables showing a copy of an X-ray video source (e.g. Live/Ref1, Ref2/Ref3, Review for monoplane system) or auxiliary system video source on a static video output. The static video output can be connected to a slave monitor (powered by the system) or to an external monitor/video system. In combination with IP based video infrastructure: maximum of 2 slave monitors in CR (close to the main workspot) and additionally maximum of 10 slave monitors at flexible location (in CR/ER outside of MCSs/Booms, part of total pool of 20 FHD monitors) maximum of 40 outputs for use outside the system. Every video output requires a counted license. The license does not actually limit the number of video outputs. DP is the video output standard Duplicating a X-ray video source that is already displayed on another monitor requires a license. The licenses are counted licenses for the number of duplicated video outputs. In combination with DVI based video infrastructure: maximum of 5 slave monitors (from that max 3 in CR) maximum of 8 outputs for use outside the system DVI-D is the video output standard Table: sub AD7XT AD7XNT Maquet tag item TruSystem notes . ptf Features -1 neuro tabletop # # -2 cardio tabletop # # -3 universal tabletop % -4 carbonfiber & neuro tabletops % variants A0…A7 for Maquet variants Carbon Floatline, Carbon Spine, SQ14-Xtra and universal for TruSystem -5 tabletop brake option o o VA kit: when power off, friction brake is applied to prevent movement of tabletop -6 table pivot option o o + see SRS.Allura.Func.PivotRotationTableBase -7 table swivel option o o only for Clea-F, see SRS.Allura.Func.SwivelMovementTableBase -8 motorized -float o o + motorized longitudinal and lateral tabletop movements -9 table tilt option o + see SRS.Allura.Func.TableTilting -10 table cradle option o + see SRS.Allura.Func.TableCradle -13 floor spacer o o ‘retention plate’, as mitigation against spillage in OR .pta AD7XT / 7XNT accessories -1 table-mounted radiation shield o o -2 peripheral filters o o -3 arm supports o o -4 head support o o -5 patient straps + + Mandatory in combination with AD7XT and AD7NXT -6 rail accessory clamp o o -7 tabletop accessory clamp o o -8 mattress o o -9 table covers (non-sterile) o o in hybrid OR, table-base, for cleaning purposes -10 additional OP-rail o o extended positioning for UI modules -11 auxiliary OP-rail kit o o positioning for third party equipment -13 neuro wedge o o for patient posturing -14 table rail cable guide + + -15 ratchet compressor o o compression belts with Ratchet-winding mechanism for symmetrical compression -16 handgrip and clamp set o o table clamps and handgrips -17 neuro head holder c c -18 Leg support board o o For use in AngioCT The system shall provide communication interfaces to allow compatible auxiliary systems to select EPX procedures which may or may not be specific for the auxiliary system (i.e. they are selectable only by the auxiliary system) set APC target positions lock/unlock the table set specific acquisition parameters (e.g. injector coupling, patient orientation, biopsy mode, etc.) push a secondary capture image to a reference viewport Table: sub item notes tag Rail susp 3rd party boom 3rd party boom XL Springarm Springarm XL Up to two rail suspended MCS and in total up to four MCS of different type and configuration can be combined in one system .frm MCS frame -1 Frame + + + + -2 VxH= 1x2 % % % 2 x 27” -3 VxH= 1x2 + 1 % 2 x 19" + 1 x 27" -4 VxH= 2x2 % % % 4 x 27” -5 VxH= 2x2 + 2 % % 6 x 27” -7 Large Screen % % % 58” XL -8 Large Screen + 2 % % % 58” XL + 2 x 27” .dsp MCS displays max. number is configuration dependent; In combination with IP based video infrastructure the total number of all monitors installed on MCSs/booms in ER and FHD monitors installed outside MCSs/booms in CR and ER never exceeds 20. The maximum doesn’t include FlexVision, Slave FlexVision, (Second) FlexSpot, Basic CR, MultiView and Additional MultiView monitors -1 27” diagnostic color % % o % o 27" color monitor, min. 400 cd/m2, resolution 1920x1080, aspect ratio 16:9 -2 32” diagnostic color1) o 32" color monitor, min. 400 cd/m2, resolution 1920x1080, aspect ratio 16:9 -4 19” multi-purpose o 19" color monitor (only for VxH= 2x1 + 1) min. 250 cd/m2, resolution 1280x1024, aspect ratio 5:4 -5 58” XL % % % 58” color monitor, min. 400 cd/m2, resolution 3840x2160, aspect ratio 16:9 -9 Dummy fill o o o .swm Switchable Monitors -1 27” or 32” diagnostic color1) o o o Up to 16 switchable monitors, can also be offered without MCS Note: 1) 32" monitors are not part of NPI. The system shall provide an option to flexibly switch during the procedure all available video sources to monitors in the examination room (within or outside MCSs/booms) and control room. Video sources can be displayed on at most 16 monitors. Switchable Monitors requires a license. In combination with DVI based video infrastructure, if FlexVision is present at most 8 switchable monitors are supported. In VoIP based infrastructure the license is a counted license for the number of Switchable Monitors. Compatibility with ECG signals from haemo equipment (acquisition of physiological signals, see SRS.Allura.Func.PhysioAcquisition). FlexVision enables flexible viewing for displaying multiple video sources (live images but also video from other systems) on one large screen display in the exam room with which the user can determine the layout, the position and size of multiple video sources on one large screen display in the exam room. Layouts are predefined. A layout divides the large screen in viewports. The user shall be able to select a preset (a preset contains the layout and for each viewport the specific video source shown in the viewport) to be applied on the large screen. The user shall be able to create, delete, change, save and rename a preset. The user shall be able to select which source is displayed in which viewport. The user shall be able to hide/unhide the video source display in a viewport. The user shall be able to swap the video sources of two viewports in a preset. The user shall be able to enlarge a specific viewport and restore it to its original size. The user shall be able to display video sources in native resolution. Display of 2k2 images shall be supported. Video can originate from the system itself and from auxiliary systems (configurable set, e.g. Interventional Workspot, IVUS). In combination with VoIP based infra, FlexVision requires a license. The display of 2K2 images requires an additional license. The system shall be tolerant against the electromagnetic fields of ablation equipment and catheter tracking systems. With FlexSpot, the system shall allow for remote control of auxiliary systems via keyboard and mouse. The system shall provide interfaces to mount, power and control an injector. Departmental Image Storage Providers and Image Storage Users An equipment rack shall be provided for placing external equipment near the table. Provided connections are: optional and country-specific hospital mains power outlets, WCB (only in combination with DVI based video infrastructure), Auxiliary System Input Connection Box mounting and cabling, infusion extension rods, extension arm with Video Electronics Standards Association (VESA) interface, optional and country-specific gas outlets for O2, NO2 and vacuum, and optional hospital Ethernet connections. It shall be possible to position the equipment rack on both sides of the table-end perpendicular to the table. Switching display, keyboard, and mouse enables the control room user to access any functionality in auxiliary applications (running on dedicated Philips PC’s). There is no implicit integration of information. The system shall support the display of its own video sources and/or sources provided by auxiliary systems on its examination/control room monitors and large screen. The system shall provide communication interfaces to exchange data with compatible auxiliary systems Patient and examination related data Acquisition related data Dose data ECG data Geometry related data
UNS.SmartSuite.BiplaneContrast | Perform imaging from two angles simultaneously to obtain an impression of the anatomy with a single contrast injection. | SRS.Allura.Func.ContrastInjection ¦ SRS.Allura.Func.BiplaneAcquisition ¦ SRS.Allura.Func.SetInjectorCoupling | For fluoro, contrast medium is to be administered manually, and no synchronization or coupling to a motorized injector is provided. In a biplane system, the user can acquire a frontal and lateral projection simultaneously, saving procedure time and contrast medium. This is manually pre-selected before the start of a run (for exposure), or based on the activated pedal(s) (for fluoro). Enables to couple motorized injection to acquisition. This requires a license. The user can select between coupled and uncoupled. At a new procedure step selection for acquisition, and customizable after every run, the injector is set to uncoupled. Timing of injection is based on a fixed "injector response time" and EPX defined image after which injection should start and “Contrast Arrival Time” (user overridable). In case the start of acquisition is delayed, a “Count-down bar” is shown representing remaining time. Type of injector coupling is customizable: One-knob: X-ray hand- or footswitch control both injection and X-ray. Two-knob: X-ray hand- or footswitch controls X-ray, but injection is enabled by injector handswitch. Acquisition Of Physiology Signals
UNS.SmartSuite.BiplaneFocusSelection | Manually change the focus ( frontal, lateral or biplane) for imaging and acquisition, if the automatic focus selection does not apply. New, N.A. | SRS.Allura.Func.SelectActiveChannel ¦ SRS.Allura.Func.BiplaneAcquisition | In a biplane system this enables the user to select the acquisition channel(s): frontal, lateral or biplane. In a biplane system, the user can acquire a frontal and lateral projection simultaneously, saving procedure time and contrast medium. This is manually pre-selected before the start of a run (for exposure), or based on the activated pedal(s) (for fluoro).
UNS.SmartSuite.BiplaneHeadSpace | Free up the head space in a biplane system. | SRS.Allura.Func.Z-RotateStandBase(L-ArmRotation) | Enables the user to rotate the frontal stand for position at the side or head of the table for optimized patient imaging. Iso-center stand rotation around a vertical axis to accomplish improved access to the patient. Function does not apply to the Poly-G floor stand.
UNS.SmartSuite.BiplaneRealTime | Perform imaging from two angles simultaneously to obtain multi-dimensional impression of the anatomy in real-time. | SRS.Allura.Func.SelectActiveChannel ¦ SRS.Allura.Conf.UI-RelatedItems ¦ SRS.Allura.Func.BiplaneAcquisition ¦ SRS.Allura.Func.ActivateExposure ¦ SRS.Allura.Func.SingleShotFootpedal ¦ SRS.Allura.Func.AcquisitionMethods ¦ SRS.Allura.Func.ActivateFluoro ¦ SRS.Allura.Func.StoreFluoro | In a biplane system this enables the user to select the acquisition channel(s): frontal, lateral or biplane. Table: sub tag item monoplane biplane .xs X-ray switches -1 footswitch 3f + -2 footswitch 6f + -3 wireless footswitch 3f o -4 wireless footswitch 6f o -5 additional footswitch CR 3f o -6 additional footswitch CR 6f o -7 additional footswitch ER 3f o -8 additional footswitch ER 6f o -9 X-ray handswitch + + Table: sub tag item Monoplane Monoplane Monoplane XY (fm/fa) XY (fm/fa) XY (fm/fa) biplane biplane biplane sub tag item AD7 XT AD7 XNT Maq AD7 XT AD7 XNT Maq Tru AD7 XT AD7 XNT Maq ui UI-modules -1 ControlModule NT +1) -2 ControlModule Tilt +1) -3 ControlModule Flex +1) +1) -4 ControlModule ORT o2) o2) +1,3) -5 ControlModule ORT Flex o2) o2) +1,3) -6 ControlModule Geo NT7) +1) -7 ControlModule Geo Tilt7) +1) -8 ControlModule Geo ORT7) o2) o2) + -9 ControlModule Geo NT IBR +1) -10 ControlModule Geo Tilt IBR +1) -11 ControlModule Imaging o4) o4) o4) o4) o4) o4) + + + -12 speed controller o5) o5) o5) o5) o5) o5) -13 pan handle o o o o o o -14 TouchScreenModule +6) +6) +6) +6) +6) +6) +6) +6) +6) -15 review module + + + + + + + + + -16 viewpad (cardio) # # # # # # # # # -17 viewpad (mixed) # # # # # # # # # -18 ControlModule Geo ORT IBR o2) o2) + Notes: 1) Additional ControlModule optional for CR 2) Optional as additional ControlModule on pedestal for ER 3) On pedestal 4) ControlModule imaging optionally available for ER (not in combination with 2)) 5) Required for FDPA, not an option for TruSystem 6) Additional TSM optional for CR and (on pedestal) for ER 7) Not for NPI, applicable for modules already present in the installed base Table: sub item Table type Table type Table type notes tag AD 7XT AD 7XNT Maquet TruSystem .mi miscellaneous -1 TSM swingarm + + + -2 Pedestal o o + trolley to attach UI modules; for AD7X(N)T including Pedestal Wall Conn. Box; for ORT including Surgery Wall Conn. Box -3 Two-way Intercom o o o -4 Keyboard + Mouse in CR + + + in CR; additional second keyboard + mouse in case of additional FlexSpot -5 Mouse in ER o o o Requires a license -7 MultiVision Switch o o o see SRS.Allura.Intgr.Multivision -8 KVM-Switch o o o see SRS.Allura.Intgr.KVM-Switching, requires DVI based video infrastructure -10 MCS Auxiliary box + + + for viewpad-receiver and X-on light on 3rd party booms In a biplane system, the user can acquire a frontal and lateral projection simultaneously, saving procedure time and contrast medium. This is manually pre-selected before the start of a run (for exposure), or based on the activated pedal(s) (for fluoro). This causes X-ray radiation, images are acquired, displayed and stored. The user is warned when the free space is insufficient to make the next exposure run. The run is stopped (in case of max. 1k2) after 1000 images, or when no more image storage free space is available. This enables the user to acquire a single image in vascular/lock-in mode directly via a separate footpedal. Image generation (also called acquisition) can be done in a number of ways: Fluoro: intended to provide a real-time X-ray image. Exposure: intended for real-time diagnostic imaging and documentation purposes. Compared to fluoro a higher signal to noise ratio (S/N) is necessary for diagnostic imaging and the images are always stored, enabling reviewing. As long as the user activates fluoro, X-ray is generated and images are acquired and displayed. Exposure will temporarily interrupt fluoro; after the exposure switch is released, fluoro is automatically resumed (if still activated). After 10 minutes of uninterrupted fluoro, fluoro is automatically stopped. Fluoro images are cached with cyclic overwrite for an EPX defined time-span. During fluoroscopy, the user can grab (i.e. store) single fluoro images or an image sequence. Multiple grabs during a single fluoro run are stored for each grab in its own run. After fluoroscopy, auto-cycle displays the cached sequence. After fluoroscopy, the user can store this complete cached run as a normal run. The cache is cleared at the start of a next fluoro run.
UNS.SmartSuite.BolusChaseReconstruction | Automatically reconstruct Bolus Chase runs immediately after acquisition in order to create an overview of the peripheral arteries in one image; Export images to the hospital information system. New, N.A. | SRS.Allura.Func.BolusChaseReconstruction | Bolus Chase Reconstruction (BCR) enables to build a single Bolus Chase reconstructed image of the complete leg(s), by automatically ”stitching” successive images that have been acquired with a Bolus Chase protocol. This requires a license. A BCR run shall be made with a specific EPX procedure (Flexible Dynamic Peripheral Angiography (FDPA)). Only one channel is supported in biplane configurations (frontal or lateral channel (lateral channel only in case of a PolyG-Floor frontal stand)). A reconstructed image can be made from a single contrast run or with both a contrast and mask run (subtracted view is then available). BCR shall be started automatically when a new BCR run is acquired. It shall be possible to start reconstruction manually by selecting a previously acquired BCR run. The user shall be able to make more than one reconstruction within one study. It shall be possible to use the reconstructed image as a reference when navigating through the original images. It shall be possible to save the Bolus Chase reconstructed image as a secondary capture image. Basic image processing (brightness, contrast, invert, subtraction on/off and landmarking), annotation and measurement functions shall be supported on the original images. Basic image processing (brightness, contrast, invert, subtraction on/off and landmarking) shall be supported on the reconstructed image. Fixed anatomy navigation mode shall be supported where original images are displayed on the selected patient's fixed anatomy. Not supported for a TruSystem table with Azurion FlexArm option. Marker Functionality
UNS.SmartSuite.CardiacSwing | : Acquire images of the patient’s anatomy from multiple views in a single run by rotating the system along curved trajectories around the patient; Rotate the system along optimized predefined trajectories for adequate image quality; Rotate the system along dedicated trajectories for the left and the right coronary arteries; Rotate the system while staying automatically within safe boundaries in order to avoid any collisions between the patient/bystander and System (‘BodyGuard’). New, N.A. | SRS.Allura.Func.Dual-AxisAcquisition | A run is acquired while the X-ray beam is rotated and angulated according a curved movement profile. This function is only available for the frontal channel. Scans are performed automatically according pre-defined acquisition/geometry settings; cardiac swing or helical. Cardiac Swing requires a license. Digital File & Data Management Focus is on the handling of files with images and/or related data, as perceived by the user. File & Data Operations
UNS.SmartSuite.Carriage | To select best treatment position and pass the table at head end from nurse to doctor side and vice versa with the FlexArm system. New, N.A. | SRS.Allura.Conf.Stand-RelatedItems | Table: sub Poly-G Poly-G Poly-G Clea Clea Clea Clea L-arc tag Item Ce XY fm Fl Ce XY fm XY fa Fl Bpl notes .std -1 Extended rail o o o o o o FlexArm: 4.3m; standard: 6.0m or 7.8m. -2 L-arm long # # # # -3 L-arm short # # # # -4 Flexarm long # -5 Flexarm short # -6 Ceiling rail covers o See SRS.Allura.Phys.LaminarAirflowCompatibility -7 Optical imaging system o o o see SRS.Allura.Intgr.OpticalImagingSystem -8 Coverbags c c c c c c c c Arc / Tube / Detector / TSM
UNS.SmartSuite.CeilingCompatibility | Install the system in existing rooms; Reuse the existing Allura/Azurion family ceiling construction as much as possible to avoid reconstruction and save cost. New, N.A. | SRS.Allura.Conf.Stand-RelatedItems | Table: sub Poly-G Poly-G Poly-G Clea Clea Clea Clea L-arc tag Item Ce XY fm Fl Ce XY fm XY fa Fl Bpl notes .std -1 Extended rail o o o o o o FlexArm: 4.3m; standard: 6.0m or 7.8m. -2 L-arm long # # # # -3 L-arm short # # # # -4 Flexarm long # -5 Flexarm short # -6 Ceiling rail covers o See SRS.Allura.Phys.LaminarAirflowCompatibility -7 Optical imaging system o o o see SRS.Allura.Intgr.OpticalImagingSystem -8 Coverbags c c c c c c c c Arc / Tube / Detector / TSM
UNS.SmartSuite.ChangeDetectorOrientation | Image in portrait and landscape detector position independent of the gantry position; Align the detector to the anatomy of the patient independent of the gantry position*. *this option is specific to the FlexArm and Azurion 7 B20/15LN. New, N.A. | SRS.Allura.Func.ImageBeamRotation ¦ SRS.Allura.Func.RotateFDXDDetector | In FlexArm (XY-fa) and biplane (Clea3 & L-arc3 only) configurations, the frontal beam shall be rotatable i.e. collimator and detector rotate simultaneously and stay continuously aligned (IBR). Enables the user to rotate the frontal FD20 detector for optimized imaging (manual and motorized). User shall be able to rotate the detector by 90° in order to switch between landscape and portrait usage of the detector. Beam width is adapted to maintain the selected detector field size. User shall be able to rotate the detector by any angle between 0° and 90° in order to align with specific anatomy. In FlexArm (XY-fa) and biplane (Clea3 & L-arc3 only) configurations, the collimator rotates synchronously to maintain the selected detector field-size. In all other configurations, the collimator is fixed and the beam-width is limited such that the irradiated area is completely visible.
UNS.SmartSuite.ChangeFluoroExposure | Change the X-ray protocol during exposure and fluoroscopy. New, N.A. | SRS.Allura.Func.ContrastInjection ¦ SRS.Allura.Func.SelectActiveChannel ¦ SRS.Allura.Conf.UI-RelatedItems ¦ SRS.Allura.Func.ActivateExposure ¦ SRS.Allura.Func.SingleShotFootpedal ¦ SRS.Allura.Func.AcquisitionMethods ¦ SRS.Allura.Func.SetInjectorCoupling ¦ SRS.Allura.Func.ActivateFluoro ¦ SRS.Allura.Func.StoreFluoro | For fluoro, contrast medium is to be administered manually, and no synchronization or coupling to a motorized injector is provided. In a biplane system this enables the user to select the acquisition channel(s): frontal, lateral or biplane. Table: sub tag item monoplane biplane .xs X-ray switches -1 footswitch 3f + -2 footswitch 6f + -3 wireless footswitch 3f o -4 wireless footswitch 6f o -5 additional footswitch CR 3f o -6 additional footswitch CR 6f o -7 additional footswitch ER 3f o -8 additional footswitch ER 6f o -9 X-ray handswitch + + Table: sub tag item Monoplane Monoplane Monoplane XY (fm/fa) XY (fm/fa) XY (fm/fa) biplane biplane biplane sub tag item AD7 XT AD7 XNT Maq AD7 XT AD7 XNT Maq Tru AD7 XT AD7 XNT Maq ui UI-modules -1 ControlModule NT +1) -2 ControlModule Tilt +1) -3 ControlModule Flex +1) +1) -4 ControlModule ORT o2) o2) +1,3) -5 ControlModule ORT Flex o2) o2) +1,3) -6 ControlModule Geo NT7) +1) -7 ControlModule Geo Tilt7) +1) -8 ControlModule Geo ORT7) o2) o2) + -9 ControlModule Geo NT IBR +1) -10 ControlModule Geo Tilt IBR +1) -11 ControlModule Imaging o4) o4) o4) o4) o4) o4) + + + -12 speed controller o5) o5) o5) o5) o5) o5) -13 pan handle o o o o o o -14 TouchScreenModule +6) +6) +6) +6) +6) +6) +6) +6) +6) -15 review module + + + + + + + + + -16 viewpad (cardio) # # # # # # # # # -17 viewpad (mixed) # # # # # # # # # -18 ControlModule Geo ORT IBR o2) o2) + Notes: 1) Additional ControlModule optional for CR 2) Optional as additional ControlModule on pedestal for ER 3) On pedestal 4) ControlModule imaging optionally available for ER (not in combination with 2)) 5) Required for FDPA, not an option for TruSystem 6) Additional TSM optional for CR and (on pedestal) for ER 7) Not for NPI, applicable for modules already present in the installed base Table: sub item Table type Table type Table type notes tag AD 7XT AD 7XNT Maquet TruSystem .mi miscellaneous -1 TSM swingarm + + + -2 Pedestal o o + trolley to attach UI modules; for AD7X(N)T including Pedestal Wall Conn. Box; for ORT including Surgery Wall Conn. Box -3 Two-way Intercom o o o -4 Keyboard + Mouse in CR + + + in CR; additional second keyboard + mouse in case of additional FlexSpot -5 Mouse in ER o o o Requires a license -7 MultiVision Switch o o o see SRS.Allura.Intgr.Multivision -8 KVM-Switch o o o see SRS.Allura.Intgr.KVM-Switching, requires DVI based video infrastructure -10 MCS Auxiliary box + + + for viewpad-receiver and X-on light on 3rd party booms This causes X-ray radiation, images are acquired, displayed and stored. The user is warned when the free space is insufficient to make the next exposure run. The run is stopped (in case of max. 1k2) after 1000 images, or when no more image storage free space is available. This enables the user to acquire a single image in vascular/lock-in mode directly via a separate footpedal. Image generation (also called acquisition) can be done in a number of ways: Fluoro: intended to provide a real-time X-ray image. Exposure: intended for real-time diagnostic imaging and documentation purposes. Compared to fluoro a higher signal to noise ratio (S/N) is necessary for diagnostic imaging and the images are always stored, enabling reviewing. Enables to couple motorized injection to acquisition. This requires a license. The user can select between coupled and uncoupled. At a new procedure step selection for acquisition, and customizable after every run, the injector is set to uncoupled. Timing of injection is based on a fixed "injector response time" and EPX defined image after which injection should start and “Contrast Arrival Time” (user overridable). In case the start of acquisition is delayed, a “Count-down bar” is shown representing remaining time. Type of injector coupling is customizable: One-knob: X-ray hand- or footswitch control both injection and X-ray. Two-knob: X-ray hand- or footswitch controls X-ray, but injection is enabled by injector handswitch. Acquisition Of Physiology Signals As long as the user activates fluoro, X-ray is generated and images are acquired and displayed. Exposure will temporarily interrupt fluoro; after the exposure switch is released, fluoro is automatically resumed (if still activated). After 10 minutes of uninterrupted fluoro, fluoro is automatically stopped. Fluoro images are cached with cyclic overwrite for an EPX defined time-span. During fluoroscopy, the user can grab (i.e. store) single fluoro images or an image sequence. Multiple grabs during a single fluoro run are stored for each grab in its own run. After fluoroscopy, auto-cycle displays the cached sequence. After fluoroscopy, the user can store this complete cached run as a normal run. The cache is cleared at the start of a next fluoro run.
UNS.SmartSuite.Cleaning | Clean/disinfect system parts in the immediate vicinity of patients; Easily clean the cable covers. Receive or have direct access to instructions for appropriate cleaning/disinfection regimes of system parts (e.g. procedure, frequency, materials, and fluids). New, N.A. | SRS.Allura.Phys.Cleanability ¦ SRS.Allura.UI.Help ¦ SRS.Allura.SLS.patient-contamination ¦ SRS.Allura.UI.Help.Customize | System parts that can come in contact with the patients' secretions, excretions, other body fluids, or other fluids shall be designed to be cleaned and disinfected by disinfection agents/compounds. All other system parts shall be designed to be cleaned by water with detergents. The user shall be able to access integrated help and electronic Instructions for Use. This includes tooltips, task and procedure guidance and instruction material. The user shall be able to search for a specific topic within the Instructions for Use. The system shall be designed for: Cleaning and disinfection of system parts in the patient environment. Creating an aseptic environment. The user shall be able to delete and modify a help header and link a new help header to an imported hospital protocol/checklist. This requires a license.
UNS.SmartSuite.ClinicalImagesOnTSM | Process and view images on the TSM*); Navigate images on TSM without using remote control; Zoom/pan images on the TSM without using remote control; Reset the zoom/panning level on TSM; Display semi-live the acquired X-ray, IVUS and 3D (SmartCT) images on the TSM. *) Image display on TSM is not for diagnostic purposes. New, N.A. | SRS.Allura.UI.TSM.Image | The TSM shall display live X-ray images for image post-processing, image navigation and moving shutters and wedges on the image. Beside this live X-ray image, live images can be shown from applications running on other systems that are compatible with the multi-modality TSM. This requires a license
UNS.SmartSuite.CollisionDetection | Rely on a collision avoidance system that: Maintains rotation and angulation speeds until it comes close to the patient; Is compatible with released Philips native tables and released integrated ORT solutions. New, Rely on a collision avoidance system that: Prevents collision with the patient; Allows the user to override the stop and continue at a slow speed in close proximity to the patient and/or system parts. | SRS.Allura.SLS.movement-integrity ¦ SRS.Allura.SLS.collision-avoidance ¦ SRS.Allura.Intgr.Magnus ¦ SRS.Allura.Func.OverrideMovementsBlockage ¦ SRS.Allura.Intgr.TruSystem ¦ SRS.Allura.SLS.collision-harm-reduction ¦ SRS.Allura.SLS.careless-manual-operation ¦ SRS.Allura.SLS.collision-unintend-ui ¦ SRS.Allura.SLS.unexpected-manual-behavior | Out of control movements (e.g. spontaneous movements, movements with high speed than requested) shall be prevented in all conditions of the system (during start-up, normal use, power failure, power off). During motorized movements, the system shall help the operator to avoid collisions with a person. Compatibility with Magnus OR table of Maquet. Move table via system’s UI modules and Magnus UI modules to position table top, tilt, cradle and set height. Move table for reset geometry, APC iso-center and FDPA. Stop table movements in case of emergency stop. Detect collisions and allow override table movements. Synchronize table lock setting, patient orientation setting, table positions, status and capabilities. This function enables the user to override a stop for collision avoidance reasons, in order to reach positions that were otherwise not possible. During override specific measures apply. Compatibility with TruSystem 7500 OR table of Hillrom In combination with the FlexArm: Move table via system's UI modules and TruSystem UI modules to position table top, tilt, cradle and set height. Move table for reset geometry, APC iso-center. Stop table movements in case of emergency stop. Detect collisions and allow override table movements. Synchronize table lock settings, patient orientation setting, table positions, status and capabilities. Excluded are: FDPA, iso-centric tilt, and Syncra tilt. In combination with the FlexMove there is limited integration: Table can be locked by the system The system's emergency stop also stops all active table movements In cases where motorized movements (despite preventive measures) do lead to collisions of mechanical parts with a person, the collision forces shall be limited to prevent harm. During manual movements the system shall help the operator to prevent collisions with a person. Measures shall be taken to prevent unintended motorized movements related to user interface design and user interaction with the system. Measures shall be taken to prevent unintended and/or uncontrolled manual movement behavior.
UNS.SmartSuite.CompatibilityLAF | Fit FlexArm in ORs offering a free floor and such that Raumklasse 1A (DIN 1946) can be established in an examination room with a Laminar Air Flow system with the FlexArm in parking position and in working position (the latter preferred, but not required); The user should not experience any jet stream when carriage moves under LAF area. New, N.A. | SRS.Allura.Phys.LaminarAirflowCompatibility | FlexArm and FlexMove configurations (FD20 XY-fa and XY-fm) shall not impact compliance with hygiene standard (DIN-1946-4, norm Raum Klasse 1A) in operating rooms having a Laminar Airflow. For FlexArm configurations flow-optimized rail covers shall be optionally available.
UNS.SmartSuite.CompensatePatientMovements | Perform diagnosis on patients even when small patient and/or table unintended movements occur during acquisition of images (automatic image correction in case of small patient/table movements). Export motion correction parameters to peripheral systems (such as PACS, workstations); Do/undo automatic motion correction (toggle) to check and set quality of result. New, N.A. | SRS.Allura.Func.Roadmap.AutomaticPixelShift ¦ SRS.Allura.Func.NormalSubtraction ¦ SRS.Allura.Func.DICOM.Export ¦ SRS.Allura.Func.ExposureSubtract ¦ SRS.Allura.Func.X-rayImageProcessing | During the roadmap Vessel and Device phases, optionally automatic pixel-shift is applied to the live image (EPX controlled) to compensate for undesired movements. Minimal movement in acquisition is assumed, to obtain enhanced view with single mask subtraction. Enables that the user can step through the run to see the progressive filling of arteries or veins. The user can switch subtraction on/off, and/or select another mask from the run. Manually, registration errors due to patient movements can be corrected ("Pixel-shift" function). A background portion can be added manually for orientation purposes ("Landmarking" function). Subtracted imaging requires a license. Enables to export (part-of) a study to a remote DICOM workstation or to a DICOM archive. Auto-export at the end of run and/or finish study, according EPX defined destination (multiple) and contents (e.g. images, downscaling). The user shall be able to get an overview of which runs will be auto-exported upon finishing the study and to which destination. For manual export, the user can select the content and the destination. For manual export, the content can be de-identified and the user can define a fictive patient name for the de-identified content. The selected destination determines the image processing format (raw or processed) which will be used. A biplane run is exported as two separate DICOM objects that refer to each other. Displays the run subtracted after the mask is acquired. The mask is automatically selected (EPX procedure defined). The user can override subtraction on/off for the complete exposure run (before acquisition has started). Optionally automatic pixel-shift is applied to the live image (EPX controlled) to compensate for undesired movements. Subtracted imaging requires a license. Image processing is applied to enhance the diagnostic value and the user's appreciation of the images. Processing is EPX controlled to allow optimized images for each type of procedure and each user. Fluoro processing depends on the selected fluoro-flavour. The processing algorithm ClarityIQ is optionally available. The option OncoBlue enables ClarityIQ image processing restricted to abdominal examinations only. ClarityIQ and OncoBlue both require a separate license. Common Functionality For Fluoro And Exposure
UNS.SmartSuite.ControlModule | Protect the system from ingress of fluids (e.g. IPX4); Control the system at table side using a compact, simple and waterproof TSO, which allows for sufficient patient access and workspace. Have visibility and guidance in the dark (e.g. backlighting); Move the Control Module from the pedestal to the table and vice versa. New, N.A. | SRS.Allura.Func.RotateBeam ¦ SRS.Allura.Conf.UI-RelatedItems ¦ SRS.Allura.SLS.ListOfApplicableStandards ¦ SRS.Allura.UI.UIModule ¦ SRS.Allura.UI.Workspot ¦ SRS.Allura.Func.AngulateBeam | The user shall be able to rotate the beam around the longitudinal-table-axis (motorized). Intuitive control and display orientation rules (see SRS.Allura.Func.DisplayOrientation) also apply if rotation causes a reversed beam direction. In biplane systems, the frontal and lateral beam can be rotated simultaneously. During motorized movement the rotation position shall be displayed. Table: sub tag item monoplane biplane .xs X-ray switches -1 footswitch 3f + -2 footswitch 6f + -3 wireless footswitch 3f o -4 wireless footswitch 6f o -5 additional footswitch CR 3f o -6 additional footswitch CR 6f o -7 additional footswitch ER 3f o -8 additional footswitch ER 6f o -9 X-ray handswitch + + Table: sub tag item Monoplane Monoplane Monoplane XY (fm/fa) XY (fm/fa) XY (fm/fa) biplane biplane biplane sub tag item AD7 XT AD7 XNT Maq AD7 XT AD7 XNT Maq Tru AD7 XT AD7 XNT Maq ui UI-modules -1 ControlModule NT +1) -2 ControlModule Tilt +1) -3 ControlModule Flex +1) +1) -4 ControlModule ORT o2) o2) +1,3) -5 ControlModule ORT Flex o2) o2) +1,3) -6 ControlModule Geo NT7) +1) -7 ControlModule Geo Tilt7) +1) -8 ControlModule Geo ORT7) o2) o2) + -9 ControlModule Geo NT IBR +1) -10 ControlModule Geo Tilt IBR +1) -11 ControlModule Imaging o4) o4) o4) o4) o4) o4) + + + -12 speed controller o5) o5) o5) o5) o5) o5) -13 pan handle o o o o o o -14 TouchScreenModule +6) +6) +6) +6) +6) +6) +6) +6) +6) -15 review module + + + + + + + + + -16 viewpad (cardio) # # # # # # # # # -17 viewpad (mixed) # # # # # # # # # -18 ControlModule Geo ORT IBR o2) o2) + Notes: 1) Additional ControlModule optional for CR 2) Optional as additional ControlModule on pedestal for ER 3) On pedestal 4) ControlModule imaging optionally available for ER (not in combination with 2)) 5) Required for FDPA, not an option for TruSystem 6) Additional TSM optional for CR and (on pedestal) for ER 7) Not for NPI, applicable for modules already present in the installed base Table: sub item Table type Table type Table type notes tag AD 7XT AD 7XNT Maquet TruSystem .mi miscellaneous -1 TSM swingarm + + + -2 Pedestal o o + trolley to attach UI modules; for AD7X(N)T including Pedestal Wall Conn. Box; for ORT including Surgery Wall Conn. Box -3 Two-way Intercom o o o -4 Keyboard + Mouse in CR + + + in CR; additional second keyboard + mouse in case of additional FlexSpot -5 Mouse in ER o o o Requires a license -7 MultiVision Switch o o o see SRS.Allura.Intgr.Multivision -8 KVM-Switch o o o see SRS.Allura.Intgr.KVM-Switching, requires DVI based video infrastructure -10 MCS Auxiliary box + + + for viewpad-receiver and X-on light on 3rd party booms The standards as mentioned in chapter APPLICABLE STANDARDS are applicable. Environmental Care The review and control modules shall support backlighting to allow easy visibility of buttons in e.g. dark examination rooms and give user guidance. The UI modules at or near the table shall allow control under sterile conditions (e.g. under a sterile cover). The user shall be able to detach the UI modules from the pedestal resp. table and attach them to the table resp. pedestal. Movement and X-ray generating functions shall be controllable at or near the table (e.g on a pedestal). Changes the angle between beam and longitudinal-table-axis (motorized). Angulation can be performed simultaneously with beam rotation. During motorized movement the angle position is displayed.
UNS.SmartSuite.ControlRoomDesign | Look over the top of screens in control room to see the exam-room; Minimize the amount of monitors; Operate different system applications with harmonized keyboard and mouse interfaces. New, Have the same monitor configuration in the control room for the biplane and monoplane systems. | SRS.Allura.Intgr.AuxiliarySystems ¦ SRS.Allura.Intgr.Flexspot.KeyboardMouse ¦ SRS.Allura.Intgr.Flexspot ¦ SRS.Allura.UI.KeyboardMouse | With FlexSpot it shall be possible to install dedicated Philips auxiliary systems in the cabinet in the technical room. FlexSpot shall provide infrastructure to gain access from control room to the media devices available on these PCs; on DVI based video configurations the maximum number of PCs from dedicated Philips auxiliary systems installed in the cabinet in the technical room is three. on VoIP based video configuration the maximum number of PCs from dedicated Philips auxiliary systems installed in the cabinet in the technical room is four. The maximum number of RTO clients in the system (TR/CR/ER) is five for VoIP based configurations, and six for DVI based configurations. Without FlexSpot it shall be possible to install a single PC from dedicated Philips auxiliary systems in the cabinet in the technical room. However no further infrastructure except power and network access is provided. With FlexSpot, the system shall allow for remote control of auxiliary systems via keyboard and mouse. Provides the user with the same features as for FlexVision in the control room. Exception is that FlexSpot doesn't support display of 2k2 images. In combination with VoIP based infrastructure, FlexSpot requires a license. A second monitor for the FlexSpot workspot shall be offered. This second monitor requires a license. The keyboard and mouse button behavior shall be harmonized throughout the different system applications in the control room.
UNS.SmartSuite.CreateWorklist | Create a work list of scheduled patients and associated procedures; Manually create a work list and add patients; Automatically create a work list when the system is connected to the hospital information system (e.g. RIS/CIS). New, N.A. | SRS.Allura.Func.ShowStudyInfo ¦ SRS.Allura.SLS.IHE ¦ SRS.Allura.Func.AddStudy | The user shall be able to view the information of a selected study, comprising at least: Study info: overview of runs, applied EPXs, cumulative dose, fluoro time. Run/image info: exposure parameters, projections. The user shall be able to filter the runs within the selected study. The system implements all of the transactions specified in the Integrating the Healthcare Enterprise (IHE) Technical Framework to support the IHE Integration Profiles, Actors and Options listed below: Table: Integration Profiles Implemented Actors Implemented Options Implemented Cardiac Catheterization Workflow (CATH) Acquisition Modality Broad Worklist Query Scheduled Workflow (SWF) Acquisition Modality Patient Based Worklist Query, Broad Worklist Query Patient Information Reconciliation (PIR) Acquisition Modality none Consistent Presentation of Images (CPI) Acquisition Modality none Radiation Exposure Monitoring (REM) Acquisition Modality none Consistent Time (CT) Time Client none Audit Trail and Node Authentication (ATNA) Secure Node none This function enables to add a new entry and related information to the study list. If operational, the WLM system provides the data for scheduling (e.g. patient name/ID, scheduled procedure start date/time). Alternatively, it remains possible to add a study entry locally.
UNS.SmartSuite.DataHandler | Save reference images in specific destinations (e.g. reference file, Photo file, PACS, Information Systems) and in various standard formats. View/display small pictorials of image series (e.g. by using a data handler) coming from previous examinations (e.g. stored on the system database or retrieved from PACS) or the current examination; View/display small pictorials of image series arranged by filter settings (e.g. middle image; View/display relevant details of image series (e.g. pictorial pop-up); Easily access applications (e.g. Measurements, QA, BCR) from the DataHandler; Create/edit names for a single or for multiple fluoroscopy/exposure run(s). Have the last used series filter settings persistently available. New, N.A. | SRS.Allura.Func.ModifyStudy ¦ SRS.Allura.Func.StoreToMedia ¦ SRS.Allura.Func.DICOM.Export ¦ SRS.Allura.Func.ShowStudyInfo ¦ SRS.Allura.Func.DisplayStudyList ¦ SRS.Allura.Func.RunLogging ¦ SRS.Allura.Func.CopyImages | This function enables to modify existing study data. The following data cannot be modified locally: Data originating from a WLM system (modify on WLM). Auto-stored data (e.g. run info); exception: the run description. This function enables to store image data on media (CD/DVD/USB). The user selects DICOM format or PC format. With DICOM format, the user can select to store the DICOM viewer with the exported DICOM files. With PC format, the user determines the file name for creating a PNG (Portable Network Graphic) file or an MPEG4 (Moving Picture Experts Group) file. The content can be deidentified and the user can define a fictive patient name for the deidentified content. Enables to export (part-of) a study to a remote DICOM workstation or to a DICOM archive. Auto-export at the end of run and/or finish study, according EPX defined destination (multiple) and contents (e.g. images, downscaling). The user shall be able to get an overview of which runs will be auto-exported upon finishing the study and to which destination. For manual export, the user can select the content and the destination. For manual export, the content can be de-identified and the user can define a fictive patient name for the de-identified content. The selected destination determines the image processing format (raw or processed) which will be used. A biplane run is exported as two separate DICOM objects that refer to each other. The user shall be able to view the information of a selected study, comprising at least: Study info: overview of runs, applied EPXs, cumulative dose, fluoro time. Run/image info: exposure parameters, projections. The user shall be able to filter the runs within the selected study. This function lists the clinical (not field-service) studies present in the system's database. The presentation is filtered (e.g. schedule list only). The list can be sorted and searched for easy navigation to the targeted patient/study. Navigation within patient/study is enabled through the display of representative images ("pictorial index"). Information of stored fluoro and exposure runs is collected in the run log, as part of the study. This comprises: Figures on accumulated skin-dose and area-dose. Applied generator settings (e.g. KV/mA/mS) and X-ray spectral filter; per run. Applied geometry (e.g. rotation, angulation) settings; per run. Name of EPX procedure, time of acquisition, and number of images; per run. Fluoro Specific Functions The user shall be able to copy a single image, or a selection of images.
UNS.SmartSuite.Decommission | Decommission* the system in the image-guided lab. *, e.g. remove from PRS, data destruction, 'un-setting to work' of the system, Philips Remote Services, etc. | SRS.Allura.Serv.FSInstructions ¦ SRS.Allura.UI.Help | Instructions are available on how to unpack, install, de-install, decommission, and packing and disposal of the system and components. The user shall be able to access integrated help and electronic Instructions for Use. This includes tooltips, task and procedure guidance and instruction material. The user shall be able to search for a specific topic within the Instructions for Use.
UNS.SmartSuite.DeleteRun | Manually delete a complete run; Prevent accidental deletion of a complete run. New, N.A. | SRS.Allura.Func.DeleteRun | A complete run (containing one or more images) can be deleted manually. Measures have to be taken to prevent accidental deletion (e.g. request the user to confirm).
UNS.SmartSuite.DeleteStudy | Free up system storage space by deleting a study/studies and all related data from the database; Protect a study from automatic deletion by the system; Receive a message when the respective study cannot be deleted (e.g. study is protected or tasks are pending). New, N.A. | SRS.Allura.Func.DeleteStudy ¦ SRS.Allura.Func.Protect/UnprotectStudy | This function enables to delete a study and all related data from the database. Deletion can be done manually, or by the system to free storage space. The study is removed from the database and free storage space will increase. In case the respective study is protected or tasks (e.g. archiving) are pending, the user is prompted (e.g. to confirm deletion, or to remove the protection). The study selected for acquisition cannot be deleted. Protection safeguards a study from automatic deletion. The user shall be able to set the protection on/off. Unprotected studies can be deleted by the system (to free-up storage space), unless related pending tasks (e.g. archiving) are not completed yet. If sufficient storage space can no longer be created by the system the user shall be warned.
UNS.SmartSuite.Diagnosis | Diagnose problems on the system in the image-guided lab*. Use diagnosis to determine correct action to solve the problems. *All diagnostic capabilities and information that is needed to determine the corrective action to solve the (future) issue, e.g. diagnostic capabilities and information needed: correct actual lab configuration, fingerprint, logging, health status,image-guided system FRUs, tests, adjustments, EPX database, ServiceHub, heamo system, ultrasound (that can be moved to different lab), 3rd party devices, usage history (who used the device and for what reason), to enable condition/usage based preventive maintenance. | SRS.Allura.Qual.FSResponseTimes ¦ SRS.Allura.UI.ServiceUILanguage ¦ SRS.Allura.Serv.UserInterface ¦ SRS.Allura.Intgr.Interfaces.RemoteConnectionSettings ¦ SRS.Allura.Serv.HardwareTests ¦ SRS.Security.MD.RemoteAccessPoint ¦ SRS.Allura.Serv.NetworkDiagnosis ¦ SRS.Allura.Serv.FunctionalTests ¦ SRS.Allura.Serv.Maintenance.Predictive ¦ SRS.Allura.Serv.Adjust ¦ SRS.Allura.Serv.Logging ¦ SRS.Allura.Serv.SystemMonitoring&Alerting ¦ SRS.Allura.Serv.RemoteConnectivityStatus ¦ SRS.Allura.Serv.SystemInformation ¦ SRS.Allura.Serv.SystemInformation.Remote ¦ SRS.Allura.Serv.RemoteSFTP | Table: Categories/ items Max. notes Field-Service Tool sec Local start-up service framework: - from start-up to log-on screen 15 - from log-on screen to display framework 10 after entering service key password - start-up with already logged on service user 20 Remote start-up service framework 100 Local shut-down service framework: - from close to service conclusion screen 6 after removing service key or after closing the close confirmation screen - from service conclusion screen to shut-down 3 EPX Management sec Start-up EPX Management Tool 30 includes loading of database settings in editor Store modified settings for verification 30 up to modified settings active for try-out Batch verification (one modified EPX-Application/Procedure) 30 +10 sec per additional modified EPX-Application/Procedure Store modified changes to system 120 up to usable clinical system Calibration min Calibrate FD20 90 Calibrate FD15 90 Calibrate FD12 90 Frontal Stand calibration 340 calibrate geometry without X-ray Lateral Stand calibration 225 calibrate geometry without X-ray Table calibration 75 calibrate geometry without X-ray Generator calibration 90 Image transfer sec Image max. 5122 - 12 bit 150 over 40 Kbps link Image Quality This section describes the basic Image Quality requirements. The purpose is to describe the system in terms of Image Quality requirements such, that if a system meets the requirements in this specification, the image quality of that system is guaranteed (and acceptable for the end user). The service UI for service engineers is in English. QUALITY AND PERFORMANCE REQUIREMENTS In this chapter the requirements are given for the quality and performance requirements. X-ray Performance Figures The service UI shall provide access to the whole system. The service UI shall be based on a standardized service framework and the EPX management tool. The system shall provide on-line help information. The system shall provide communication interfaces to store and retrieve RSN remote connection settings. Room-Specific Auxiliary Systems Applied to diagnose the correct functioning of the Field Replaceable Units (FRU’s) and related interfaces. Power-On SelfTests (‘POST’) are performed automatically at system power-up. Build-In SelfTests (‘BIST’) enable the FSE to examine a particular FRU into greater detail. The Azurion shall be protected against attacks via a remote access point. The system shall provide means to detect and diagnose network connectivity issues for the (system internal) control network for the real time image link for the connection to external DICOM nodes on the hospital network for remote service connection The system shall provide tests to check the functional integrity of the system or part of the system. The tests shall provide information to enable the FSE to find a fault by reasoning, based on supplied documentation. Appropriate (Edoc) documentation shall be provided to guide the FSE in performing successive tests. The system shall provide remotely accessible indicators of the X-ray tube filament and video distribution components for wear-out analyzing purposes. Local Service Functionality This focuses on local (non-remote) service functionality, after the system has been assembled. Installation/Configuration/Adjustment/Customization Build-in correction measures shall provide adjustment functionality to bring the system within specified limits. The service framework shall provide adjustment criteria and procedures to the FSE and hospital technician. System information and events are stored and accessible e.g. for diagnostic/faultfinding purposes. Logged info comprises of system-info (e.g. configuration), system usage (e.g. user-actions, system-errors, wear-out indicators), and test results. FSE functions: read (select/filter/search), clear logfile, export logfile. Auxiliary Functions The system shall monitor the system status and pro-actively alert a service center accordingly. The alert shall be sent when ‘pro-active limits’ are exceeded (e.g. indicated by a wear-out sensor). The system shall inform the clinical user about the remote connection status The system shall provide a remote connection test for clinical user and hospital IT. For VPN connected systems it shall be possible to up- and download remote configuration settings from/to the system. The FSE shall be able to view system information. Note that no table information is available for the TruSystem table in a FlexMove configuration. The remote FSE shall be able to view system information. The system shall support the transfer of the system information to RSN. The system shall support data exchange with the RSN using the Secure File Transfer Protocol (sFTP) for both ISSL-link and VPN connected systems MANUFACTURING REQUIREMENTS This chapter focuses on additional requirements related to manufacturing and assembly activities. Requirements on installation, configuration, adjustments, customization, verification & report, diagnostics and faultfinding can be found in chapter SERVICE REQUIREMENTS; automated system test requirements can be found in chapter SYSTEM INTEGRATION AND VERIFICATION REQUIREMENTS.
UNS.SmartSuite.DicomNetworkExporting | Manually and automatically export images from the X-ray system to PACS and/or an external workstation in the hospital; Automatically transfer patient data (e.g. images, annotations) to the ISP; Schedule automatic transfer of images. Reproduce (grey levels, up/downscaling and (un)subtraction) exported images on a wide variety of image archives, such as external viewers as PACS. Automatically transfer cardiac data to cardio-PACS (with different settings, e.g. downscale, processed) and radiology data to radiology PACS (with different settings, e.g. no downscale, unprocessed) to optimally support mixed procedures. Configure transfer actions based on procedures used with different settings to destinations (workstations, PACS), customizable by the user. Limit/ restrict export settings configuration for unauthorized users; Show network and status of connected systems with modality related information (e.g. IP address, AE title). New, Group single images from a study and export as one series to a DICOM destination. Azurion reduces the DICOM export time from the system. The export time of both cardiac (multi-frame) and radiology data (single-frame) is at least 2x faster compared to Allura Xper FD R8.2* Optimized interpolation; IQ subtraction for paedriatic patients. *Multi-frame: export of 10 runs * 100 images of 1K. Single-frame: export of 40 images of 1K. | SRS.Allura.Qual.ClinicalResponseTimes.Arch ¦ SRS.Allura.Func.DICOM.Export/ImportJobManagement ¦ SRS.Allura.Intgr.ImageExportImport ¦ SRS.Allura.Func.SystemCustomization ¦ SRS.Allura.Func.Workflow.Flexibility ¦ SRS.Allura.Func.DICOM.Consistency ¦ SRS.Allura.Func.DICOM.Import ¦ SRS.Allura.Func.DICOM.Export ¦ SRS.Allura.Func.FinishStudy ¦ SRS.Allura.Func.DICOM.Query ¦ SRS.Allura.Func.Diagnostics ¦ SRS.Allura.SLS.IHE | performance values apply to transfer over a network with a speed of at least 1Gbps. Table: Categories/ items Max. notes: (all values in seconds) Export 10 runs * 100 images max. 1K2 200 multiframe, 8 bit, with Jpeg Import 40 images max. 1K2 900 singleframe 12bit, with & without Jpeg Export 40 images max. 1K2 100 singleframe 12bit, with & without Jpeg Printing 1 sheet with 12 images 60 Printer resolution satisfying source resol. Field-Service Response Times This function enables to observe a submitted import or export job. The user can abort, delete and retry the job. The system can be connected to departmental image storage providers (e.g. Xcelera PACS) and/or storage users. Compatibility according to DICOM conformance statement. Support for Query (DICOM SCU), Export (DICOM SCU) and Import (DICOM SCP). Supported formats: XA-multi-frame, Secondary Capture, JPEG-lossless. Supported features: Grayscale and non Grayscale Standard Display Compatible, DICOM Presentation state, Storage Commit. Departmental Worklist Management (WLM) Systems Allows the user to customize system settings, that are available for the system's configuration, such as: Date, Time and Regional Settings User Administration and Administration settings List of annotation and performing physician Procedure cards and RIS code mapping Settings for viewing and quantitative analysis Media and DICOM export settings APC positions Linking DICOM export settings, APC positions and procedure help text to EPX procedures Remote support settings Defining APC pathways Saved settings are persistent over system restarts. System customization settings can be imported and exported. Administration Functions This section focuses on administrative functions to deal with patient data and related information. General Administration Concepts A local database includes all information that is entered manually and generated in acquisition. It can contain multiple patients/studies/procedure steps comprising patient data, images, etc. Local Administration Functions The user shall be able to have a workflow according local preferences. Different studies/procedure steps can overlap in time, i.e. The user shall be able to suspend the current acquisition study/procedure step, select another study/procedure step for acquisition and later select the suspended study/procedure step for acquisition again. The user shall be able to select a study for viewing besides an already selected study for acquisition. Sequence of workflow steps is not fixed, i.e. the user shall be able to switch between administration, image generation, viewing, post-processing, printing, reporting, archiving and distribution for the same study. Customizable automated workflow, i.e. Automatic DICOM export. Automatic dose report creation. Automatic DICOM dose structured report. Automatic finishing of the current acquisition study when another study is selected for acquisition. Automatic setting a study to protected against deletion upon finishing of the study. Automatic selecting of procedure card based on RIS code. Concurrency It is always required to restore data consistency between the system and a connected node after the cause of an error has been removed (e.g. complete or restart an image transfer after power-down). Enables the user to manually import images from a remote destination (workstation or archive). The user selects a remote destination, determines the image data and imports it; or on a remote destination, the user determines the image data and pushes it to the system. An original biplane run is reconstructed from separate (frontal and lateral) DICOM objects. Enables to export (part-of) a study to a remote DICOM workstation or to a DICOM archive. Auto-export at the end of run and/or finish study, according EPX defined destination (multiple) and contents (e.g. images, downscaling). The user shall be able to get an overview of which runs will be auto-exported upon finishing the study and to which destination. For manual export, the user can select the content and the destination. For manual export, the content can be de-identified and the user can define a fictive patient name for the de-identified content. The selected destination determines the image processing format (raw or processed) which will be used. A biplane run is exported as two separate DICOM objects that refer to each other. With this function the user indicates that this study is finished on the system. Per procedure step within the study, EPX defined automation tasks are started (e.g. export). After pending jobs (e.g. archiving) are successfully finished, the procedure step is finished and (if MPPS is configured) reported according DICOM MPPS. Unless set to protected, the occupied storage space is released for automatic deletion. Examination Related Presetting EPX Concept Enables the user to obtain an administrative overview of the information present on a remote workstation or archive. The user can enter search criteria and start a query; the query can be aborted. The user can sort to optimize the presented query results. The user shall be able to assess the status and settings of the hospital network interface and all incoming and outgoing network connections. The system implements all of the transactions specified in the Integrating the Healthcare Enterprise (IHE) Technical Framework to support the IHE Integration Profiles, Actors and Options listed below: Table: Integration Profiles Implemented Actors Implemented Options Implemented Cardiac Catheterization Workflow (CATH) Acquisition Modality Broad Worklist Query Scheduled Workflow (SWF) Acquisition Modality Patient Based Worklist Query, Broad Worklist Query Patient Information Reconciliation (PIR) Acquisition Modality none Consistent Presentation of Images (CPI) Acquisition Modality none Radiation Exposure Monitoring (REM) Acquisition Modality none Consistent Time (CT) Time Client none Audit Trail and Node Authentication (ATNA) Secure Node none
UNS.SmartSuite.DisplayConfiguration | To see all relevant images for his tasks from his/her working position; In case of FlexVision (Pro), choose screen size (heamo live screen needs to remain in original resolution); View at selected screen location without having to lean towards the screen (images and UI should be clearly visible and readable); Slave screens for monitoring purposes; See a source on multiple monitors; Connect 3rd party screens; Connect the system to 3rd party video switching and streaming solutions; Activate or de-activate a general purpose on-screen stopwatch. New, Display and control up to 8 sources (one at a time) on the additional MultiSwitch without the dependency on the physical connection to the KVM switch; Increase switching flexibility when compared to the current DVI infrastructure, in the exam room and control room by enabling a flexible division between inputs, outputs and displays, with the routing flexibility of the network switch; Enable multiple signal layouts from external sources. Make sure some sources are always displayed (e.g. always see the live screen(s) and live heamo screen); Choose which image (live, ref) is shown on which screen in the exam room. Enables multiple signal layouts accordingly to the user preferences, on systems with FlexVision (Pro) and FlexSpot, due to the ability to connect up to 20 external video signals (not generated directly by the X-Ray system); Systems with 16 Switchable Monitors can connect up to 20 external video signals (not generated directly by the X-Ray system); Systems with FlexSpot can have up to 3 integrated workstations. | SRS.Allura.Func.Roadmap.DevicePhase.VesselUnSubtract ¦ SRS.Allura.Intgr.OpenVideoInterface.LargeScreenDownscaled ¦ SRS.Allura.Intgr.MultiView ¦ SRS.Allura.Conf.Control-RoomDisplays ¦ SRS.Allura.Intgr.Multivision ¦ SRS.Allura.Intgr.3rdAnd4thMultiView ¦ SRS.Allura.Func.Roadmap.DevicePhase.LiveSubtract ¦ SRS.Allura.Intgr.OpenVideoInterface.LargeScreen ¦ SRS.Allura.Intgr.Flexspot.AdditionalWorkspot ¦ SRS.Allura.Func.Select/DeselectStopwatch ¦ SRS.Allura.UI.TSM.Image ¦ SRS.Allura.Intgr.Flexspot ¦ SRS.Allura.Intgr.AdditionalMultiView ¦ SRS.Allura.Intgr.Flexvision.KeyboardMouse ¦ SRS.Allura.Intgr.VideoSlave ¦ SRS.Allura.Conf.MCS-RelatedItems ¦ SRS.Allura.Intgr.SwitchableMonitors ¦ SRS.Allura.UI.PositionMCSFrame ¦ SRS.Allura.Intgr.Flexvision ¦ SRS.Allura.Func.NavigateToImageData ¦ SRS.Allura.Func.ActivateFluoro ¦ SRS.Allura.Intgr.Flexspot.KeyboardMouse ¦ SRS.Allura.Func.Roadmap.DevicePhase ¦ SRS.Allura.Func.Roadmap.DevicePhase.VesselSubtract ¦ SRS.Allura.SLS.ListOfApplicableStandards ¦ SRS.Allura.Func.ActivateExposure ¦ SRS.Allura.Intgr.VideoDisplayInterface ¦ SRS.Allura.Func.DisplayInfo | In this mode, the images shall be displayed unsubtracted (no DeviceMask), but the images shall be displayed superimposed by the (faded) VesselMask. The system shall provide a downscaled copy of the large screen of FlexVision video output in FHD resolution (DP is the interface in combination with IP based video infrastructure and SL-DVI in combination with DVI based video infrastructure). In combination with IP based video infrastructure that video output shall be also available (configurable) for switching on Flexspot, (Additional,3rd, and 4th) Multiview and Switchable Monitors. In VoIP infrastructure, the availability of the downscaled copy requires a license. MultiView user can select what video source (1 of max 8, it is configurable what video sources shall be selectable) shall be displayed and controlled (keyboard/mouse support over USB) on the review monitor. Other type of data exchange between auxiliary equipment and X-Ray System is not supported. MultiView requires a license. This is a counted license for the number of configurable sources. Table: sub tag item DVI based infra VoIP based infra notes . crd control room displays -1 acquisition and review monitor # # Both 24" color monitors with native image resolution of 1920x1080 and aspect ratio 16:9. -2 FlexSpot primary monitor # # Instead of acquisition and review monitor, a 27" color monitor with native image resolution of 2560x1440 and aspect ratio 16:9. -3 FlexSpot secondary monitor o o Requires FlexSpot primary monitor, 27" color monitor with native image resolution of 2560x1440 and aspect ratio 16:9. -4 additional FlexSpot o Requires FlexSpot primary monitor, 27" color monitor with native image resolution of 2560x1440 and aspect ratio 16:9; shall be placed either in the control or exam room. -5 slave monitor1) o o 24” (desk version) or 27” (boom version) or 32" (boom version) color monitor with native image resolution of 1920x1080 and aspect ratio 16:9, e.g. extra live monitor. -6 MultiView o Requires acquisition and review 24" color FHD monitors, -7 Additional MultiView o Requires acquisition and review monitor, extra 24" color FHD monitor, mouse and keyboard; -8 Second FlexSpot primary monitor o Requires FlexSpot primary monitor, 27" QHD color monitor, mouse and keyboard; shall be placed either in the control or exam room. -9 Second FlexSpot secondary monitor o Requires Second FlexSpot primary monitor, 27" color QHD monitor; shall be placed at the same location as Second FlexSpot primary monitor -10 3rd MultiView o Requires extra 27" color QHD monitor, mouse and keyboard. Only in combination with FlexSpot, second FlexSpot, and 4th MultiView -11 4th MultiView o Requires extra 27" color QHD monitor, mouse and keyboard. Only in combination with 3rd MulitView Note: 1) 32" monitors are not part of NPI. MultiVision enables to display video on the exam room display originating from multiple video sources. The user can select the current video source for display. Video can originate from the system itself and/or from auxiliary systems (e.g. Interventional Workspot, IVUS). Not available in combination with Switchable Monitors or FlexVision. In a DVI configuration, the control room review monitor is selectable as video source and max. 4 other video sources are configurable. MultiVision requires a license In combination with FlexSpot and second FlexSpot, a 3rd and 4th workspot (keyboard, mouse, single monitor) shall be offered, which can be a workspot in the control room or exam room. One video source (out of max. 8, it is configurable what video sources shall be selectable) can be selected to be displayed and controlled. Video can originate from the system itself or from auxiliary systems. The 3rd and 4th MultiView require a separate license per workspot. This is a counted license for the number of configurable sources. In this mode, the images shall be displayed subtracted with the DeviceMask resulting into the so-called Device image. Optionally (EPX determined) the Device image shall be displayed superimposed by the LandmarkMask (containing the bones structure, derived from the DeviceMask). The system shall provide a large screen video output (DP1.2 or, in combination with DVI based video infrastructure, 2*DL-DVI) that contains a copy of the large screen of FlexVision. In VoIP infrastructure, this requires a license. A second workspot (keyboard, mouse, single monitor) next to FlexSpot shall be offered, which can be a workspot in the control room or exam room. One video source can be selected to be shown. Video can originate from the system itself and from auxiliary systems (configurable set, e.g. Interventional Workspot, IVUS). The additional FlexSpot requires a license. General purpose on-screen stopwatch; not coupled to X-ray generation events The TSM shall display live X-ray images for image post-processing, image navigation and moving shutters and wedges on the image. Beside this live X-ray image, live images can be shown from applications running on other systems that are compatible with the multi-modality TSM. This requires a license Provides the user with the same features as for FlexVision in the control room. Exception is that FlexSpot doesn't support display of 2k2 images. In combination with VoIP based infrastructure, FlexSpot requires a license. A second monitor for the FlexSpot workspot shall be offered. This second monitor requires a license. A second workspot (keyboard, mouse, single monitor) next to acquisition and review monitor shall be offered, which can be a workspot in the control room or exam room. One video source (out of max. 8, it is configurable what video sources shall be selectable) can be selected to be displayed and controlled (keyboard/mouse support over USB). Video can originate from the system itself or from auxiliary systems. Additional MultiView requires a license. This is a counted license for the number of configurable sources. With FlexVision, the system shall allow for remote control of auxiliary systems via on-screen keyboard and touchpad (on TSM) and / or wireless mouse. This requires a license. VideoSlave enables showing a copy of an X-ray video source (e.g. Live/Ref1, Ref2/Ref3, Review for monoplane system) or auxiliary system video source on a static video output. The static video output can be connected to a slave monitor (powered by the system) or to an external monitor/video system. In combination with IP based video infrastructure: maximum of 2 slave monitors in CR (close to the main workspot) and additionally maximum of 10 slave monitors at flexible location (in CR/ER outside of MCSs/Booms, part of total pool of 20 FHD monitors) maximum of 40 outputs for use outside the system. Every video output requires a counted license. The license does not actually limit the number of video outputs. DP is the video output standard Duplicating a X-ray video source that is already displayed on another monitor requires a license. The licenses are counted licenses for the number of duplicated video outputs. In combination with DVI based video infrastructure: maximum of 5 slave monitors (from that max 3 in CR) maximum of 8 outputs for use outside the system DVI-D is the video output standard Table: sub item notes tag Rail susp 3rd party boom 3rd party boom XL Springarm Springarm XL Up to two rail suspended MCS and in total up to four MCS of different type and configuration can be combined in one system .frm MCS frame -1 Frame + + + + -2 VxH= 1x2 % % % 2 x 27” -3 VxH= 1x2 + 1 % 2 x 19" + 1 x 27" -4 VxH= 2x2 % % % 4 x 27” -5 VxH= 2x2 + 2 % % 6 x 27” -7 Large Screen % % % 58” XL -8 Large Screen + 2 % % % 58” XL + 2 x 27” .dsp MCS displays max. number is configuration dependent; In combination with IP based video infrastructure the total number of all monitors installed on MCSs/booms in ER and FHD monitors installed outside MCSs/booms in CR and ER never exceeds 20. The maximum doesn’t include FlexVision, Slave FlexVision, (Second) FlexSpot, Basic CR, MultiView and Additional MultiView monitors -1 27” diagnostic color % % o % o 27" color monitor, min. 400 cd/m2, resolution 1920x1080, aspect ratio 16:9 -2 32” diagnostic color1) o 32" color monitor, min. 400 cd/m2, resolution 1920x1080, aspect ratio 16:9 -4 19” multi-purpose o 19" color monitor (only for VxH= 2x1 + 1) min. 250 cd/m2, resolution 1280x1024, aspect ratio 5:4 -5 58” XL % % % 58” color monitor, min. 400 cd/m2, resolution 3840x2160, aspect ratio 16:9 -9 Dummy fill o o o .swm Switchable Monitors -1 27” or 32” diagnostic color1) o o o Up to 16 switchable monitors, can also be offered without MCS Note: 1) 32" monitors are not part of NPI. The system shall provide an option to flexibly switch during the procedure all available video sources to monitors in the examination room (within or outside MCSs/booms) and control room. Video sources can be displayed on at most 16 monitors. Switchable Monitors requires a license. In combination with DVI based video infrastructure, if FlexVision is present at most 8 switchable monitors are supported. In VoIP based infrastructure the license is a counted license for the number of Switchable Monitors. The user shall be able to position the MCS frame(s) to have a clear view on the monitors from the workspot. FlexVision enables flexible viewing for displaying multiple video sources (live images but also video from other systems) on one large screen display in the exam room with which the user can determine the layout, the position and size of multiple video sources on one large screen display in the exam room. Layouts are predefined. A layout divides the large screen in viewports. The user shall be able to select a preset (a preset contains the layout and for each viewport the specific video source shown in the viewport) to be applied on the large screen. The user shall be able to create, delete, change, save and rename a preset. The user shall be able to select which source is displayed in which viewport. The user shall be able to hide/unhide the video source display in a viewport. The user shall be able to swap the video sources of two viewports in a preset. The user shall be able to enlarge a specific viewport and restore it to its original size. The user shall be able to display video sources in native resolution. Display of 2k2 images shall be supported. Video can originate from the system itself and from auxiliary systems (configurable set, e.g. Interventional Workspot, IVUS). In combination with VoIP based infra, FlexVision requires a license. The display of 2K2 images requires an additional license. Enables the user to zoom-in progressively on a particular file and/or a run, down to a specific image. After selection of a study for viewing or for acquisition, the first exposure or fluoro run is active for viewing. Exposure overrules the selection and makes the latest acquired run active for viewing. As long as the user activates fluoro, X-ray is generated and images are acquired and displayed. Exposure will temporarily interrupt fluoro; after the exposure switch is released, fluoro is automatically resumed (if still activated). After 10 minutes of uninterrupted fluoro, fluoro is automatically stopped. With FlexSpot, the system shall allow for remote control of auxiliary systems via keyboard and mouse. In this roadmap phase, the fluoro images shall be displayed according the current roadmap mode settings (EPX), meaning in one of the following Device phase modes: LiveSubtract; VesselSubtract; or VesselUnSubtract. Exposure runs can be made in-between Device phase runs without losing the related masks. In this mode, the images shall be displayed subtracted with the DeviceMask resulting into the so-called Device image, and the Device image shall be displayed superimposed by the (faded) VesselMask. Optionally (EPX determined) the Device image shall be displayed superimposed by the LandmarkMask (containing the bones structure, derived from the VesselMask). The standards as mentioned in chapter APPLICABLE STANDARDS are applicable. Environmental Care This causes X-ray radiation, images are acquired, displayed and stored. The user is warned when the free space is insufficient to make the next exposure run. The run is stopped (in case of max. 1k2) after 1000 images, or when no more image storage free space is available. The system shall support the display of its own video sources and/or sources provided by auxiliary systems on its examination/control room monitors and large screen. Displays detailed information about the current acquisition parameters. The display of information can be customized for live- and review viewport.
UNS.SmartSuite.DoorContact | Inhibit X-ray from the system while the door is open between Exam Room and adjacent rooms. New, N.A. | SRS.Allura.Func.DoorContact | The system shall be configurable to have a door contact connected to inhibit start of X-ray while the door is open between exam room and other adjacent rooms/corridor.
UNS.SmartSuite.DoseControl | Choose fluoroscopy flavours (e.g. Low, Medium, Normal) and different acquisition settings for same procedure (low vs high dose, low vs high quality) to adjust image quality; Perform acquisition with As-Low-As-Reasonably-Achievable (ALARA) X-ray dose to patient and staff, i.e. optimize balance between radiation dose and IQ by adjusting relevant settings (e.g. ProcedureCard, DAP, frame rate, FoV). Activate acquisition from all sides of table (head, side (left/right), arm, feet side, remotely (trolley, control room); Operate system at a safe distance/shielded (with regard to user exposure dose) during fluoroscopy and acquisition; See the pre-acquisition time; Connect multiple footswitches and hand switches in ER/CR to the system; receive dose information; both fluoroscopy and exposure. | SRS.Allura.Func.Move/ResetWedges ¦ SRS.Allura.Conf.UI-RelatedItems ¦ SRS.Allura.Conf.RemainingComponents ¦ SRS.Allura.Func.DisplayAndReportPatientDose ¦ SRS.Allura.Func.FluoroFlavours ¦ SRS.Allura.Func.Move/ResetShutters | Enables wedges to move in/out, or rotate. Reset puts the wedges outside the imaging area. The default position of wedges is EPX application/procedure determined. Table: sub tag item monoplane biplane .xs X-ray switches -1 footswitch 3f + -2 footswitch 6f + -3 wireless footswitch 3f o -4 wireless footswitch 6f o -5 additional footswitch CR 3f o -6 additional footswitch CR 6f o -7 additional footswitch ER 3f o -8 additional footswitch ER 6f o -9 X-ray handswitch + + Table: sub tag item Monoplane Monoplane Monoplane XY (fm/fa) XY (fm/fa) XY (fm/fa) biplane biplane biplane sub tag item AD7 XT AD7 XNT Maq AD7 XT AD7 XNT Maq Tru AD7 XT AD7 XNT Maq ui UI-modules -1 ControlModule NT +1) -2 ControlModule Tilt +1) -3 ControlModule Flex +1) +1) -4 ControlModule ORT o2) o2) +1,3) -5 ControlModule ORT Flex o2) o2) +1,3) -6 ControlModule Geo NT7) +1) -7 ControlModule Geo Tilt7) +1) -8 ControlModule Geo ORT7) o2) o2) + -9 ControlModule Geo NT IBR +1) -10 ControlModule Geo Tilt IBR +1) -11 ControlModule Imaging o4) o4) o4) o4) o4) o4) + + + -12 speed controller o5) o5) o5) o5) o5) o5) -13 pan handle o o o o o o -14 TouchScreenModule +6) +6) +6) +6) +6) +6) +6) +6) +6) -15 review module + + + + + + + + + -16 viewpad (cardio) # # # # # # # # # -17 viewpad (mixed) # # # # # # # # # -18 ControlModule Geo ORT IBR o2) o2) + Notes: 1) Additional ControlModule optional for CR 2) Optional as additional ControlModule on pedestal for ER 3) On pedestal 4) ControlModule imaging optionally available for ER (not in combination with 2)) 5) Required for FDPA, not an option for TruSystem 6) Additional TSM optional for CR and (on pedestal) for ER 7) Not for NPI, applicable for modules already present in the installed base Table: sub item Table type Table type Table type notes tag AD 7XT AD 7XNT Maquet TruSystem .mi miscellaneous -1 TSM swingarm + + + -2 Pedestal o o + trolley to attach UI modules; for AD7X(N)T including Pedestal Wall Conn. Box; for ORT including Surgery Wall Conn. Box -3 Two-way Intercom o o o -4 Keyboard + Mouse in CR + + + in CR; additional second keyboard + mouse in case of additional FlexSpot -5 Mouse in ER o o o Requires a license -7 MultiVision Switch o o o see SRS.Allura.Intgr.Multivision -8 KVM-Switch o o o see SRS.Allura.Intgr.KVM-Switching, requires DVI based video infrastructure -10 MCS Auxiliary box + + + for viewpad-receiver and X-on light on 3rd party booms Table: sub tag item availability notes .ps Partner Systems See for details corresponding requirement in Room-Specific Auxiliary Systems section -1 DoseAware Xtend o .mi Miscellaneous -1 Ceiling Suspended radiation shield o -2 Examination light o -3 X-ray spacer o limits minimum Source Skin Distance (Health and Human Services (HHS) requirement). Mandatory for 21 CFR 1020.30. -4 Drip stand o -5 Medical DVD recorder o Requires DVI based video infrastructure -6 1 phase UPS o see SRS.Allura.Func.PowerFailureBehaviour -9 UPS control c see SRS.Allura.Intgr.ExternalThreePhaseUPS -10 VideoSlave o see SRS.Allura.Intgr.VideoSlave -11 Open video interface o see SRS.Allura.Intgr.OpenVideoInterface.LargeScreen, -12 DVD burner o see SRS.Allura.Func.StoreToMedia -13 HD output o see SRS.Allura.Intgr.OpenVideoInterface.LargeScreenDownscaled; in DVI based video infrastructure it requires FlexVision open video interface -14 Video scan converter o for 50/60 Hz SLR TV output, requires DVI based video infrastructure -15 Wall-connection boxes o see SRS.Allura.Intgr.FHDVideoSources, requires DVI based video infrastructure -16 Auxiliary System Input Connection Box o see SRS.Allura.Intgr.FHDVideoSources, requires IP based video infrastructure -17 4K Auxiliary System Input Connection Box o see SRS.Allura.Intgr.4KVideoInputOutput, not available in combination with DVI based video infrastructure Core Functionality Core functionality is defined as functionality which is not dependent on the presence of any license key. The following functions are in the core: Clinical core functionality is a subset of all functionality as defined in chapter FUNCTIONAL REQUIREMENTS. The subset is determined by subtracting all functionality that is dependent on any license key. All functionality needed to install a system and to perform basic checks to initially verify correct system working condition. User messages (see SRS.Allura.Func.DisplayUserMessages). Tooltip (balloon) help. Software-Keyed Clinical Options A Clinical Option gives additional clinical functions/features/characteristics and is enabled through a software license key. Two types of licenses exist: counted license: used to implement licenses that restrict the number of items that can be used, e.g. number of video outputs. uncounted license: used to implement licenses that restrict the usage of a certain feature (on/off) This applies to Dose Area Product (DAP) and skin-entrance-dose (air-kerma). Dose info is displayed during and after acquisition. Generated and displayed values comply with legal demands for required accuracy. The user is warned if accumulated air-kerma within the current body zone is above a customizable threshold value (default value 2 Gray). The cumulative (zone-)dose figures are comprised in the study. The user can select between 3 EPX defined flavours, determining e.g. fluoro-frame-rate and X-ray dose. The flavour-determined frame-rate can be 30, 25, 20, 15, 12.5, 10, 7.5, 6.25, 5, 3.75, 3.125, 2.5, 1.875, 1.25, 1.0, 0.625, 0.5 img/sec. In biplane, frontal frame-rate equals lateral frame-rate (available range same as for monoplane). A fourth flavour can be selected by the system in situations where low-load is necessary (e.g. for graceful degradation). For legal reasons, a low-dose fluoro flavour is available with a ratio normal/low of at least a factor 2. Configurable, a high-dose flavour with a maximum of 20 R/min is offered. A flavour can also be selected during fluoro. Enables shutters to move in/out. Opposite shutters move in pairs; horizontal and vertical pairs can be moved simultaneously. Opening stops at the edge of the current detector field (so becomes again visible at a larger field). Closing must be immediately visible (i.e. fast moving to edge of detector field, if initially larger). In case a user-defined position is overruled (by switching towards a smaller field-size), enlarging the field-size will automatically widen the shutters until the original user-defined position is reached. Reset puts the shutters at the edge of the detector field-size. The default position of shutters is EPX application/procedure determined.
UNS.SmartSuite.DoseReporting | Create/obtain a report with dose data: Cumulative AK in skin area (in case of Cardiac/Thorax zones) Cumulative Air Kerma (AK) Cumulative DAP Total Fluoro Time Save/ export dose data (e.g. by using DICOM Radiation Dose Structured Report); Save/export the report on/to the information system or an external destination in a commonly used file formats. New, N.A. | SRS.Allura.Func.StoreToMedia ¦ SRS.Allura.Func.DICOM.DoseStructuredReport ¦ SRS.Allura.Func.DICOM.Export ¦ SRS.Allura.Func.DICOM.Import ¦ SRS.Allura.Func.Reporting.Dose | This function enables to store image data on media (CD/DVD/USB). The user selects DICOM format or PC format. With DICOM format, the user can select to store the DICOM viewer with the exported DICOM files. With PC format, the user determines the file name for creating a PNG (Portable Network Graphic) file or an MPEG4 (Moving Picture Experts Group) file. The content can be deidentified and the user can define a fictive patient name for the deidentified content. The system shall be able to generate and export a DICOM X-Ray Radiation Dose Structured Report object (RDSR) upon finishing a study. Local Media Storage Enables to export (part-of) a study to a remote DICOM workstation or to a DICOM archive. Auto-export at the end of run and/or finish study, according EPX defined destination (multiple) and contents (e.g. images, downscaling). The user shall be able to get an overview of which runs will be auto-exported upon finishing the study and to which destination. For manual export, the user can select the content and the destination. For manual export, the content can be de-identified and the user can define a fictive patient name for the de-identified content. The selected destination determines the image processing format (raw or processed) which will be used. A biplane run is exported as two separate DICOM objects that refer to each other. Enables the user to manually import images from a remote destination (workstation or archive). The user selects a remote destination, determines the image data and imports it; or on a remote destination, the user determines the image data and pushes it to the system. An original biplane run is reconstructed from separate (frontal and lateral) DICOM objects. The system shall be able to create a dose report which contains information for each run and accumulated dose for a study. The system shall be able to create the report upon finishing of the study (customizable). A preview function is available to observe the resulting report. The report can be printed or (DICOM) exported. Archiving & Distribution This section focuses on the archiving and distribution of image data. DICOM Networking Functions
UNS.SmartSuite.DualFluoroscopy | View the fluoroscopy images subtracted and un-subtracted simultaneously (‘Dual Fluoroscopy’); For biplane; View one reference image per channel at the same time with the live Frontal & Lateral image next to dual fluoroscopy; Display subtracted and un-subtracted exposures simultaneously; Digitally zoom in/out the subtracted and un-subtracted live image during Dual Fluoroscopy; View live image, reference image and Dual Fluoroscopy simultaneously. New, N.A. | SRS.Allura.Func.StoreReferenceRun/Image ¦ SRS.Allura.Func.DualFluoro ¦ SRS.Allura.Func.AcquisitionZoom | The user shall be able to add the active image or run to a reference viewport. This option requires a license. Enables the user to display the fluoro image simultaneously in subtracted and unsubtracted mode. This requires a license The display showing the subtracted image can be zoomed interactively also during live acquisition. Enables the user to zoom the live image during acquisition (center-zoom). Zoom value and zoom-center position can be set by the user during Last Image Hold (LIH). Acquisition zoom is only allowed if the unzoomed image is displayed also (e.g. in dual-fluoro)
UNS.SmartSuite.DynamicRotationalAngiography | Gain 3-D insight of the patient’s anatomy (by making a dynamic rotational angiography scan where the X-ray beam rotates around the patient's area of interest placed within the iso-center); Control a rotational scan from all table sides/pedestal (head, foot, nurse, physician, specific); Make a rotational scan with a wide range of motion to cover the complete patient’s anatomy (from head to toe); Make a rotational scan with a wide range of motion to cover the complete anatomy of obese (BMI>30) patients. New, N.A. | SRS.Allura.Func.DRAFor3D ¦ SRS.Allura.Func.ClassicDRA ¦ SRS.Allura.Func.DualPhaseXperCT ¦ SRS.Allura.Func.XperCT | DRA for 3D is a special case of DRA performed at predefined positions, to enable 3D reconstruction. Settings, including start- and end-positions, shall be defined in EPX. The system shall support DRA for 3D at least at the following positions: for FlexArm (XY-fa) configuration at head, doctor and nurse side for other FD20 and FD15 configurations at head and one side position (doctor or nurse) for FD12 only at head position SID is maximized to allow for 3D scans of obese patients. The user shall define the end- and start-position of the DRA scan, implicitly executing a test-scan to detect obstructions. The system shall be able to perform the DRA scan at the side or head of the table (C-arm roll and propellor respectively). DRA acquisition is only possible for the frontal channel. The user shall be able to acquire a second run to enable subtracted reviewing. Settings, including start- and end-positions, shall be defined in EPX, but if allowed in EPX, the user shall be able to choose another end- and/or start-position. DRA requires a license. This is a special case of XperCT where two consecutive DRA scans are being made by executing the scans in different phases of contrast propagation, which is known as Dual phase XperCT. The two scans are performed in opposite direction to reduce the overall procedure time to minimize the amount of contrast medium. For neuro applications, to allow the use of intravenous contrast injections in Dual phase XPerCT scans, the system shall provide a subtracted fluoroscopy flavor with enhanced contrast. It shall enable to detect the arrival of contrast agent in order to start the scan. This application shall be executable from the Control Room. Flexible Dynamic Peripheral Angiography (FDPA) This is a special case of DRA or DAR for 3D enabling reconstruction of soft-tissue. This function is only supported with an FD20 detector. The system shall offer dedicated EPX procedures and related settings for Neuro (DRA and DAR) and Abdomen (DRA only) applications. The system shall perform acquisition with up to 60 images/second to obtain acceptable scan time. The system shall be able to apply slit collimation to reduce patient dose.
UNS.SmartSuite.ECGTrigger | Synchronize fluoroscopy and acquisition timing with patient’s ECG signal (e.g. ‘ECG triggering’ to acquire images in the same phase of the heart cycle). New, N.A. | SRS.Allura.Intgr.PhysioEquipment(ECG) ¦ SRS.Allura.Func.EPTriggering | Compatibility with ECG signals from haemo equipment (acquisition of physiological signals, see SRS.Allura.Func.PhysioAcquisition). Acquisition timing can be synchronized to the patients ECG signal. The user can select triggering on/off and control trigger-delay. Not applicable to low frame-speed vascular exposure. Synchronization requires a license.
UNS.SmartSuite.EMFieldImageFiltering | Apply fluoroscopy and acquisition without artefacts that are introduced by electromagnetic tracking devices. New, N.A. | SRS.Allura.Func.StripeFilter | A function shall be provided to filter the images with a so called “stripe filter” in order to get rid of image disturbances caused by external equipment via magnetic interference (e.g. Biosense-EP). The filter can be enabled / disabled via EPX.
UNS.SmartSuite.EPXtree | Perform a simple procedure selection, coupled to the selection of the ProcedureCard. New, Perform a simple procedure selection, outside the content of a the ProcedureCard. | SRS.Allura.Func.PresettingFunctions ¦ SRS.Allura.Func.ProcedureCard ¦ SRS.Allura.Func.ModifySelectionTree ¦ SRS.Allura.Func.EPX ¦ SRS.Allura.Func.ModifyEPXParameters | Presetting is determined by the information in the study/procedure step that is current for acquisition. The user can modify the information. The procedure card to be used can be modified. The performing physician can be selected from a pre-defined physician list. The patient type can be entered/modified manually. The EPX application/procedure group/procedure can be selected from a list (which contains only items that can be performed within the system's current configuration). EPX application, procedure group and procedure can be selected in control room, as well as in exam room. EPX application, procedure group and procedure can be selected during x-ray. The selected settings will then be applied for the next exposure or fluoroscopy run. The user shall be offered a procedure based workflow by having procedure specific settings for a procedure step. A procedure card determines the procedure specific settings, being: The EPX selections the user can choose from during the procedure and a default EPX. The default preset for FlexVision. The default preset for FlexSpot. The default preset for Second FlexSpot. The patient orientation and surgical view Procedure guidance. Specific settings for compatible room specific auxiliary systems Predefined procedure cards shall be delivered with the system. The user can customize procedure cards and organize/group them. The user can set which procedure card shall be used when a procedure step is selected for acquisition. This can also be set automatically based on information from the WLM system (RIS code). Allows to modify the “Tree” of available EPX applications, procedure groups and procedures. Elements of the tree can be copied, renamed and/or deleted. It is not possible to create elements 'from scratch' (so only via copy of proven sets). The scope of this functionality is system-wide, not physician-specific. The system provides pre-optimized and/or preferred conditions in subsequent workflow phases via EPX selections, which adapt the system settings according following inputs: The kind of Examination (indicated by the application, procedure group). The selected Patient type. The selected X-ray operator (performing physician). The selected procedure. Patient type and X-ray operator can be based on information from the WLM system, and/or can be selected by the user. In a biplane system, settings can differ per channel. EPX selection examples: acquisition-settings, positioning-defaults, image processing, archiving preferences. Presetting Functions Focuses on changing the value of EPX parameters. Complex settings are protected against modification by insufficiently skilled engineers. (setting-specific authorization level versus service-key defined skill level). The scope of this functionality can be system-wide, and/or physician-specific. Patient And Beam Positioning This section deals with positioning of patient and/or beam in order to select the anatomical area of interest. Orientation Aspects
UNS.SmartSuite.EasyUSBExport | Export image data through a USB connected destination; To select the format (DICOM or standard multi-media formats) and anonymize the data before exporting; To include a DICOM viewer to the exported data; Give the exported data a new name. New, N.A. | SRS.Allura.Func.Import/ExportOfClinicalImages ¦ SRS.Allura.Func.StoreToMedia ¦ SRS.Allura.Intgr.Media | For marketing and educational purposes (medical fairs, Customer Visit Centre, development, brochures etc) marketing and application department must collect clinical data / images within constraints provided by Philips and customer privacy policies and controls. The collected data must be imported in our Image lab system, processed and exported again. The processed and exported data must then be imported again on the demo system. All privacy information (name, date of birth, physician, etc.) must be left out in demo image results. This feature supports storage of DICOM images (raw and processed images) to an USB-stick or external hard drive to improve the workflow and ease-of-use for application specialists (not intended for clinical users). General System Functions Power On / Off Functions This function enables to store image data on media (CD/DVD/USB). The user selects DICOM format or PC format. With DICOM format, the user can select to store the DICOM viewer with the exported DICOM files. With PC format, the user determines the file name for creating a PNG (Portable Network Graphic) file or an MPEG4 (Moving Picture Experts Group) file. The content can be deidentified and the user can define a fictive patient name for the deidentified content. In control room, support for storage on CD/DVD and USB according general PC format and DICOM. Also importing from CD/DVD and USB according DICOM is supported. UPGRADABILITY REQUIREMENTS Upgradability existing X-Ray systems This release will be available for initial deliveries (always in combination with IP based video distribution infrastructure) and as software only field upgrade for Azurion system releases < R3.0. Upgrade will not impact external interfaces of the upgraded system. FUNCTIONAL REQUIREMENTS In this chapter the functions are grouped according to their role in the normal workflow. Global Functional Behavior Workflow Flexibility
UNS.SmartSuite.Emergency | In case of an emergency, unblock the use of primary system functionality by any user; Operate the system without entering a user name and password; Starting imaging without having entered any patient data (e.g. set-up one of the X-ray protocols as default for emergency cases); Manually move the stand in XY out of way in emergency and/or power off situation; Access a system emergency stop from the ER; Perform CPR on the table without table top floating (lock table); Free up the C-arm space around the patient manually, even in case of a power down (excluding 15”and 20” detector floor system). New, N.A. | SRS.Allura.Func.MoveBeam ¦ SRS.Allura.Func.EmergencyStop ¦ SRS.Allura.Func.TableLock ¦ SRS.Allura.Qual.ClinicalResponseTimes.System ¦ SRS.Allura.Func.ImplicitSelectionDefaultStudy ¦ SRS.Allura.Func.EmergencyPowerDown ¦ SRS.Allura.Func.MoveBeamXY | The user shall be able to move the frontal and/or lateral beam longitudinal to select a clinical area for acquisition (manual and motorized). This function only applies to ceiling suspended configurations. The user shall be able to stop all motorized movements. Within operators reach, red-colored emergency stop buttons are provided. Movements to free the patient from the system remain possible. Normal operation is resumed after a system restart. The user shall be able to lock/unlock: The table pivot movements. The table lateral movements. All table movements. For the TruSystem table in combination with FlexMove, the lateral table lock/unlock is not supported via the system UI. Beam Positioning This section focuses on moving and positioning the stand under direct user control. Table: Categories/ items Max. notes: (all values in seconds) Regular System startup (power-up) 300 Time until functionality required for the intended use is available. System startup after total power down 300 (China only) 360 (Rest of the world) Measured after external power supply restored to the system. Warm restart 90 Cold restart 360 Full cycle from “off” button press to full functionality. Recovery geometry after emergency stop 105 Time until all geometry movements available. If a user didn't select a study/procedure step for acquisition and performs acquisition, the system creates a study/procedure step under which the acquired images are stored. An Emergency Power Down interface shall be offered to switch off all power to the system when the optional 1 phase UPS is installed. A recovery mechanism shall be available to resume normal operation. Display General Data The user shall be able to move the beam longitudinal and/or transversal (manual and motorized). To select a clinical area for acquisition as an alternative to positioning the tabletop. For standby parking (typically transversal only) during a surgical procedure. For normal parking when the stand is not used.
UNS.SmartSuite.EndExamination | Suspend an examination (or intervention) on one patient, continuing with an examination on another patient, and resume the examination on the previous patient a later moment in time; Close the Exam in the ER from the Additional FlexSpot. New, N.A. | SRS.Allura.Func.SelectStudyFromSuspendedList ¦ SRS.Allura.Intgr.Flexspot.AdditionalWorkspot ¦ SRS.Allura.Func.FinishStudy ¦ SRS.Allura.Func.ReopenFinishedStudy | This function lets the user select a study/procedure step for acquisition from the suspended study list. Actions are similar compared to selection from the scheduled list, except: The last used settings are used to continue acquisition (e.g. dose data, clinical files). Newly acquired images are appended to the study/procedure step data. A second workspot (keyboard, mouse, single monitor) next to FlexSpot shall be offered, which can be a workspot in the control room or exam room. One video source can be selected to be shown. Video can originate from the system itself and from auxiliary systems (configurable set, e.g. Interventional Workspot, IVUS). The additional FlexSpot requires a license. With this function the user indicates that this study is finished on the system. Per procedure step within the study, EPX defined automation tasks are started (e.g. export). After pending jobs (e.g. archiving) are successfully finished, the procedure step is finished and (if MPPS is configured) reported according DICOM MPPS. Unless set to protected, the occupied storage space is released for automatic deletion. Examination Related Presetting EPX Concept The user is able to re-open a finished study/procedure step and select it for acquisition. The study/procedure step is re-initialized in such way that a complete workflow scenario may again be applied on the study/procedure step. Actions are similar compared to selection from the scheduled list, except: When a finished study/procedure step is selected, new images acquired are appended to the study/procedure step data. The procedure card, clinical-file structure, dose report data and other relevant data already attached to this study/procedure step are used as a starting point for any additional X-ray acquisitions. Finishing A Study
UNS.SmartSuite.EnvironmentalCare | Use a system that complies with the international environmental standards. New, N.A. | SRS.Allura.SLS.ECO.ProductLabeling ¦ SRS.Allura.SLS.ListOfApplicableStandards | The product shall provide information on labels and in the IfU according the Philips QMS standard 'Product environmental labeling requirements' [REF-2] and "Environmental requirements for IFU and DFU" [REF-3]. IHE, VA and DoD Requirements The standards as mentioned in chapter APPLICABLE STANDARDS are applicable. Environmental Care
UNS.SmartSuite.Escalation | Collect data and send to Philips Remote Service to enable remote support to be able to escalate problems to a higher support level (tier-1 to 4). | SRS.Allura.Serv.Field-ServiceAccessPoints ¦ SRS.Allura.Serv.RemoteAccessAndConnection ¦ SRS.Allura.Serv.RetrieveInformation ¦ SRS.Allura.Serv.SystemInformation.Remote ¦ SRS.Allura.Serv.RemoteImageTransfer ¦ SRS.Allura.Serv.RemoteDesktop.ServiceMode ¦ SRS.Allura.Serv.RemoteSFTP | The system shall provide the following access points to the Field-Service Engineer (FSE): ‘Local’ access: the FSE shall be able to log-on to the system in the control room. ‘Far’ access: offering non-safety critical functionality via a remote connection, i.e. it will not be possible to remotely activate X-rays or mechanical movements. The system shall be accessible from a workspot on the Remote Service Network (RSN). The system shall be able to connect to the RSN via an Integrated Secure Socket Layer (ISSL)-link or via a Virtual Private Network (VPN) established by a specific RSN Router. ISSL allows establishing a secure tunnel directly from the system over a single outgoing port. This 'far-access' mode shall be possible concurrent to local access. The key (local or remote) with the highest authorization level shall determine the access rights. The remote FSE shall be able to initiate collecting of recent log information from the system (without local user interaction). With local user interaction, the remote FSE shall be able to retrieve log and trace information (e.g. for offline analysis). Remote Software Installation The remote FSE shall be able to view system information. The system shall support the transfer of the system information to RSN. The local user shall be able to select an image(s) and transfer the image(s) and related acquisition information (e.g. KV) towards a remote center. It shall be possible to provide support to an on-site service engineer with remote desktop. In field-service mode, a remote expert shall be able to observe and assist, e.g. in fault finding/repair. Remote Miscellaneous The system shall support data exchange with the RSN using the Secure File Transfer Protocol (sFTP) for both ISSL-link and VPN connected systems MANUFACTURING REQUIREMENTS This chapter focuses on additional requirements related to manufacturing and assembly activities. Requirements on installation, configuration, adjustments, customization, verification & report, diagnostics and faultfinding can be found in chapter SERVICE REQUIREMENTS; automated system test requirements can be found in chapter SYSTEM INTEGRATION AND VERIFICATION REQUIREMENTS.
UNS.SmartSuite.ExaminationSelectionAndPresetting | Select the settings for an examination depending on procedure type, patient type and X-ray operator. Change the procedure type during the examination at tableside without leaving the ER. New, N.A. | SRS.Allura.Qual.ClinicalResponseTimes.Admin ¦ SRS.Allura.Func.PhysicianSpecificSettings ¦ SRS.Allura.Func.EPX ¦ SRS.Allura.Func.PresettingFunctions | Table: Categories/ items Max. notes: (all values in seconds) General response on UI requests e.g. switch exposure channel mono/biplane switch roadmap-on multiphase frame-rate subtraction on/off 1.0 one action; e.g. display a list, menu. Delete study 1.0 Select study for acquisition 5.0 Select study for viewing 2.8 Select EPX application for current study 4.0 Select EPX procedure for current study 1.2 in idle state, not valid when EPX is selected during an active fluoroscopy run Preferred settings can be obtained for a specific physician; EPX parameters predefined APC positions The system provides pre-optimized and/or preferred conditions in subsequent workflow phases via EPX selections, which adapt the system settings according following inputs: The kind of Examination (indicated by the application, procedure group). The selected Patient type. The selected X-ray operator (performing physician). The selected procedure. Patient type and X-ray operator can be based on information from the WLM system, and/or can be selected by the user. In a biplane system, settings can differ per channel. EPX selection examples: acquisition-settings, positioning-defaults, image processing, archiving preferences. Presetting Functions Presetting is determined by the information in the study/procedure step that is current for acquisition. The user can modify the information. The procedure card to be used can be modified. The performing physician can be selected from a pre-defined physician list. The patient type can be entered/modified manually. The EPX application/procedure group/procedure can be selected from a list (which contains only items that can be performed within the system's current configuration). EPX application, procedure group and procedure can be selected in control room, as well as in exam room. EPX application, procedure group and procedure can be selected during x-ray. The selected settings will then be applied for the next exposure or fluoroscopy run.
UNS.SmartSuite.ExceptionHandling |  | SRS.Allura.SLS.movements-failure-detection-and-fault-tolerant ¦ SRS.Allura.Func.TubeLoadProtectionAndIndication ¦ SRS.Allura.Func.DisplayServiceGuidance ¦ SRS.Allura.SLS.maintain-minimal-functionality ¦ SRS.Allura.Qual.ClinicalResponseTimes.System ¦ SRS.Allura.Func.GracefulDegradation ¦ SRS.Allura.Func.ExceptionHandling ¦ SRS.Allura.Func.DisplayUserMessages ¦ SRS.Allura.SLS.xray-imaging-failure-detection-and-fault-tolerant | Measures shall be taken to detect errors in the control of motorized movements. When errors are detected, that not directly result in a hazardous situation: The system shall allow movements (possibly with reduced performance). The user is informed. The system shall safeguard X-ray-generation components and indirectly protect the user and the patient against potential harmful consequences from overloading the tube. Low-load fluoro shall remain available. Error events initiate a service guidance. The service guidance is logged via the logging capabilities provided by the system. Additionally the error related service guidance is reported to the user. When the system becomes unavailable for the intended treatment or diagnosis, minimal functionality is maintained as long as possible to terminate the procedure and release the patient from the system. Table: Categories/ items Max. notes: (all values in seconds) Regular System startup (power-up) 300 Time until functionality required for the intended use is available. System startup after total power down 300 (China only) 360 (Rest of the world) Measured after external power supply restored to the system. Warm restart 90 Cold restart 360 Full cycle from “off” button press to full functionality. Recovery geometry after emergency stop 105 Time until all geometry movements available. In case of a system failure the available graceful degradation mode is entered, in which the system operates with reduced functionality and/or performance; the following rules apply: Via user/service guidance the graceful degradation situation is reported to the user. In a biplane system a defect in one channel must not prevent use of the other channel. The system will not try to automatically find alternate ways to realize the requested functionality. It is of the utmost importance that service functionality remains available for diagnosis. Demo System Behaviour A number of exceptional situations can be distinguished: A collision shall be indicated, also if other user messages are active. After a software error recovery, the system shall guarantee that previously stored clinical information is still available. If the user requests a function that cannot be executed in the current system state, for a reason which is not obvious to the user, the system shall display an appropriate user message. If the user requests a function and the system cannot execute the function because the system has run out of resources (e.g. storage full), the system shall display an appropriate user message. Graceful Degradation Comprises user guidance or error messages, which is to be displayed at the relevant workspot(s). Static messages are displayed as long as a cause exists (e.g. a system error). Triggered messages disappear automatically after a predefined time period. Confirmation by the user is not required. Measures shall be taken to detect errors in X-ray imaging functionality. When errors are detected that not directly result in loss of imaging functionality: The system shall allow imaging (possibly with reduced performance). The user is informed.
UNS.SmartSuite.ExportApp | Manage images/runs selected for export (including selection at table side) at table side; Export selected images/runs to DICOM destinations; Export in original raw format, in DICOM processed/unprocessed format; Export imported data. New, N.A. | SRS.Allura.Func.DICOM.Export ¦ SRS.Allura.Func.Flagging | Enables to export (part-of) a study to a remote DICOM workstation or to a DICOM archive. Auto-export at the end of run and/or finish study, according EPX defined destination (multiple) and contents (e.g. images, downscaling). The user shall be able to get an overview of which runs will be auto-exported upon finishing the study and to which destination. For manual export, the user can select the content and the destination. For manual export, the content can be de-identified and the user can define a fictive patient name for the de-identified content. The selected destination determines the image processing format (raw or processed) which will be used. A biplane run is exported as two separate DICOM objects that refer to each other. The user can flag/un-flag selected images or entire image-runs. This comprises adding a flag marker to the selected images, without implying any further action. Flagging groups the selected images in a set which can be treated as a single entity by other clinical functions. Example: archive the flagged images. Flagging does not imply any subsequent action. Only one type of flags exists. The flagged data is indicated to the user by a flag marker. In case of a biplane image/run the user can flag the images from one or both planes. Note: In combination with e.g. archiving, the flagging mechanism may result in the transference of an incomplete run. As a result, during cine display its dynamic appearance may look peculiar. Digital Viewing Emphasis is in this section on the selection of runs and images and the way this info is displayed. Viewing Navigation
UNS.SmartSuite.EyeHandCoordination | Perceive that the fluoroscopy images and exposure are displayed in real time and without noticeable delay. New, N.A. | SRS.Allura.SLS.image-latency ¦ SRS.Allura.Qual.ClinicalResponseTimes.Acquisition | To enable procedures which require hand-eye coordination, the perceived image latency shall not exceed 230 ms. Table: Categories/ items Max. notes: (all values in seconds) Switch viewing to fluoroscopy 1.8 from pedal press till the first image shown Switch viewing to exposure 1.8 from pedal press till the first image shown, For non-DSA cine Switch viewing to exposure 2.5 from pedal press till the first image shown, For acquisition types with test-shot lock-in and frame-speeds greater or equal to 4 fr/s Switch fluoro to viewing 1.8 from pedal release till display of viewing image on screen Switch exposure to viewing 1.8 from pedal release till display of viewing image on screen Switch between exposure/fluoro 1.8 For acquisition types without test-shot lock-in and frame-speeds greater or equal to 4 fr/s Switch between exposure/fluoro 2.5 For acquisition types with test-shot lock-in and frame-speeds greater or equal to 4 fr/s Digital Subtraction Angiography (DSA): delay after test-image 1.5 Change live fluoro settings 1.5 Fluo flavour, field-size (incl FD-mode). Change live exposure settings 0.5 Frame-rates for multiphase acquisitions Delay between Xray-OFF and Indicators-OFF 1.0 Fluoro pipeline perception latency 0.200 Essential for eye-hand coordination.
UNS.SmartSuite.FD12DetectorSize | Achieve a larger field of view when compared to with 10”detectors: at a selected SID while maintaining projection flexibility; to support viewing of the aortic valve and a good part of the aortic arch, preferably in one view with minimal table panning, so that one can follow the device when navigating the aortic arch and see the valve at the same time; to support viewing of the whole coronary tree in one view with minimal need for table panning. This can be challenging especially during procedures of patient with a dilated heart (cardiac myopathy); to view both the needle insertion point (subclavian or axillary vein) and the whole wire/lead towards the right atrium in one view with minimal table panning and C-arm repositioning. This can be challenging especially for patient with a dilated heart and left ventricle; See small details by having a small pixel size, resulting in high limiting resolution. New, N.A. | SRS.Allura.Conf.SupportedConfigurations ¦ SRS.Allura.Conf.XperFDSeriesConfigurations | The following FD configurations shall be supported: Table: Configuration Characteristic discriminators Azurion FD12 Poly-G Ceiling/Floor stand with FD12 detector Azurion FD15 Clea-Floor stand with FD15 detector Azurion FD20 Clea-Ceiling/Floor stand with FD20 detector Azurion FD12/12 Poly-G Floor with FD12 & Larc-C with FD12 Azurion FD20/12 Clea-Floor with FD20 & Larc-N1) or Larc-C with FD12 Azurion FD20/15 Clea-Floor with FD20 & Larc-N or Larc-C with FD15 Azurion FD20 OR-table Clea-Ceiling stand with FD20 detector Azurion FD12/12 OR-table Poly-G Floor with FD12 & Larc-C with FD12 Azurion FD20/12 OR-table Clea-Floor with FD20 & Larc-C with FD12 Azurion FD20/15 OR-table Clea-Floor with FD20 & Larc-N or Larc-C with FD15 Note 1): Azurion FD20/12 Clea-Floor with FD20 & Larc-N with FD12 is not supported for NPI. This configuration is only supported for the installed base. The following configurations are also supported as remanufactured systems: Table: Configuration Characteristic discriminators Azurion FD12 Poly-G Ceiling stand with FD12 detector and AD7 table Azurion FD20 Clea Floor stand with FD20 detector and AD7 table (no FlexMove) Detailed Configuration requirements Applied symbols: 1) In subsequent tables, additional detailing of configurations uses following abbreviations and indications: Mpl = Monoplane system Ce = Monoplane Ceiling system, suspended by longitudinal rail XY-fa = Monoplane Ceiling system, suspended by double-pivot arm (FlexArm) XY-fm = Monoplane Ceiling system, suspended by X-Y rails (FlexMove) Fl = Floor mounted Monoplane system Bpl = Biplane system (FD12/12 and FD20/12) NB = Neuro Biplane (FD20/15) Mag = Magnus Maquet Tru = Hillrom TruSystem 7500 2) Symbols are used in the tables have following meaning + = included in a system, not an option ++ = twice included in a system, not an option (in case of a Biplane system) # = each system comprises exactly one of the items marked with # in one cell % = each system comprises at least one of the items marked with % in one cell. o = option; is an extension based on extra hardware and/or based on a (software) license key c = compatibility based upon interfacing with the system 3) Use of sub-tags To deal with limited table space, sub-tags are used; to be appended to the specified table tag. Example: Table tag = SRS.Allura.Conf.XperFDSeriesConfigurations; sub-tag = .std-1 Resulting complete tag = SRS.Allura.Conf.XperFDSeriesConfigurations.std-1 Table: sub tag Systems: FD12 FD12/12 FD12 FD12/12 FD12 FD12/12 FD20 FD20 FD20 FD20 FD 15 FD20/12 FD20/15 FD20/12 FD20/15 item Ce Fl Bpl Ce XY fa XY fm Fl Fl Bpl NB notes .std Geo Stand -1 Poly-G2-C # Poly-G2-C is supported for remanufacturing -2 Poly-G3-C # -3 Poly-G2-F + + -4 Clea-2-C + + Clea-2-C is supported for remanufacturing (not XY fm) -5 Clea-3-C + -6 Clea-F # # # # -7 Clea-3-F # Only in combination with AD7 table -8 Larc-N 5) # -9 Larc-C # # # -10 Larc-3-N # Only in combination with AD7 table .pt Patient Table -1 AD7XT # # o o o o # # o o -2 AD7XNT # # o o o # # o o -3 Maquet Magnus c c c c c c -4 Hillrom TruSystem c c See SRS.Allura.Intgr.TruSystem .ms Monitor Ceiling Support -1 Rail suspended MCS % % % % % % % % % -2 Spring-arm boom % % % % % % % % % % -3 Spring-arm boom XL % % % % % % % % % % -4 3rd party boom % % % % % % % % % % MCS Monitor integration kit -5 3rd party boom XL % % % % % % % % % % MCS Monitor integration kit XL .xrc X-ray chain ++= frontal & lateral -1 Tube MRC 200+ 0508 + + ++ +1) +1) -2 Tube MRC 200+ 04072) + + + + + +1) +1) -3 Collimator Nicol3) + + ++ + + + + + ++ ++ -5 AEP dose-meter o o o o o o o o o o AEP calculated by default, optionally hardware based -7 Detector FD124) + + ++ # # -8 Detector FD15 + # # -9 Detector FD20 + + + + # # Notes: 1) Frontal channel MRC0407; lateral channel MRC0508 2) MRC 200+ 0407 ROT-GS1008 for FD20 XY-fa; ROT-GS1004 for all other configurations 3) Nicol v4 on Clea-3, L-arc 3-N and Poly-G3 Ceiling. All other stands Nicol v3 4) Monoplane configurations PX2121S or PX2121C/V; biplane configurations PX2121C/V only 5) Azurion FD20/12 Clea-Floor with FD20 & Larc-N with FD12 is not supported for NPI. This configuration is only supported for the installed base
UNS.SmartSuite.FD15DetectorSize | Achieve a 154% larger field of view than the 10” detector, to support viewing of larger vascular structures in routine procedures such as renal artery procedures, carotid artery procedures or iliac procedures. New, N.A. | SRS.Allura.Conf.SupportedConfigurations ¦ SRS.Allura.Conf.XperFDSeriesConfigurations | The following FD configurations shall be supported: Table: Configuration Characteristic discriminators Azurion FD12 Poly-G Ceiling/Floor stand with FD12 detector Azurion FD15 Clea-Floor stand with FD15 detector Azurion FD20 Clea-Ceiling/Floor stand with FD20 detector Azurion FD12/12 Poly-G Floor with FD12 & Larc-C with FD12 Azurion FD20/12 Clea-Floor with FD20 & Larc-N1) or Larc-C with FD12 Azurion FD20/15 Clea-Floor with FD20 & Larc-N or Larc-C with FD15 Azurion FD20 OR-table Clea-Ceiling stand with FD20 detector Azurion FD12/12 OR-table Poly-G Floor with FD12 & Larc-C with FD12 Azurion FD20/12 OR-table Clea-Floor with FD20 & Larc-C with FD12 Azurion FD20/15 OR-table Clea-Floor with FD20 & Larc-N or Larc-C with FD15 Note 1): Azurion FD20/12 Clea-Floor with FD20 & Larc-N with FD12 is not supported for NPI. This configuration is only supported for the installed base. The following configurations are also supported as remanufactured systems: Table: Configuration Characteristic discriminators Azurion FD12 Poly-G Ceiling stand with FD12 detector and AD7 table Azurion FD20 Clea Floor stand with FD20 detector and AD7 table (no FlexMove) Detailed Configuration requirements Applied symbols: 1) In subsequent tables, additional detailing of configurations uses following abbreviations and indications: Mpl = Monoplane system Ce = Monoplane Ceiling system, suspended by longitudinal rail XY-fa = Monoplane Ceiling system, suspended by double-pivot arm (FlexArm) XY-fm = Monoplane Ceiling system, suspended by X-Y rails (FlexMove) Fl = Floor mounted Monoplane system Bpl = Biplane system (FD12/12 and FD20/12) NB = Neuro Biplane (FD20/15) Mag = Magnus Maquet Tru = Hillrom TruSystem 7500 2) Symbols are used in the tables have following meaning + = included in a system, not an option ++ = twice included in a system, not an option (in case of a Biplane system) # = each system comprises exactly one of the items marked with # in one cell % = each system comprises at least one of the items marked with % in one cell. o = option; is an extension based on extra hardware and/or based on a (software) license key c = compatibility based upon interfacing with the system 3) Use of sub-tags To deal with limited table space, sub-tags are used; to be appended to the specified table tag. Example: Table tag = SRS.Allura.Conf.XperFDSeriesConfigurations; sub-tag = .std-1 Resulting complete tag = SRS.Allura.Conf.XperFDSeriesConfigurations.std-1 Table: sub tag Systems: FD12 FD12/12 FD12 FD12/12 FD12 FD12/12 FD20 FD20 FD20 FD20 FD 15 FD20/12 FD20/15 FD20/12 FD20/15 item Ce Fl Bpl Ce XY fa XY fm Fl Fl Bpl NB notes .std Geo Stand -1 Poly-G2-C # Poly-G2-C is supported for remanufacturing -2 Poly-G3-C # -3 Poly-G2-F + + -4 Clea-2-C + + Clea-2-C is supported for remanufacturing (not XY fm) -5 Clea-3-C + -6 Clea-F # # # # -7 Clea-3-F # Only in combination with AD7 table -8 Larc-N 5) # -9 Larc-C # # # -10 Larc-3-N # Only in combination with AD7 table .pt Patient Table -1 AD7XT # # o o o o # # o o -2 AD7XNT # # o o o # # o o -3 Maquet Magnus c c c c c c -4 Hillrom TruSystem c c See SRS.Allura.Intgr.TruSystem .ms Monitor Ceiling Support -1 Rail suspended MCS % % % % % % % % % -2 Spring-arm boom % % % % % % % % % % -3 Spring-arm boom XL % % % % % % % % % % -4 3rd party boom % % % % % % % % % % MCS Monitor integration kit -5 3rd party boom XL % % % % % % % % % % MCS Monitor integration kit XL .xrc X-ray chain ++= frontal & lateral -1 Tube MRC 200+ 0508 + + ++ +1) +1) -2 Tube MRC 200+ 04072) + + + + + +1) +1) -3 Collimator Nicol3) + + ++ + + + + + ++ ++ -5 AEP dose-meter o o o o o o o o o o AEP calculated by default, optionally hardware based -7 Detector FD124) + + ++ # # -8 Detector FD15 + # # -9 Detector FD20 + + + + # # Notes: 1) Frontal channel MRC0407; lateral channel MRC0508 2) MRC 200+ 0407 ROT-GS1008 for FD20 XY-fa; ROT-GS1004 for all other configurations 3) Nicol v4 on Clea-3, L-arc 3-N and Poly-G3 Ceiling. All other stands Nicol v3 4) Monoplane configurations PX2121S or PX2121C/V; biplane configurations PX2121C/V only 5) Azurion FD20/12 Clea-Floor with FD20 & Larc-N with FD12 is not supported for NPI. This configuration is only supported for the installed base
UNS.SmartSuite.FD20DetectorSize | Make images up to 48cm diagonal. New, N.A. | SRS.Allura.Conf.SupportedConfigurations ¦ SRS.Allura.Conf.XperFDSeriesConfigurations | The following FD configurations shall be supported: Table: Configuration Characteristic discriminators Azurion FD12 Poly-G Ceiling/Floor stand with FD12 detector Azurion FD15 Clea-Floor stand with FD15 detector Azurion FD20 Clea-Ceiling/Floor stand with FD20 detector Azurion FD12/12 Poly-G Floor with FD12 & Larc-C with FD12 Azurion FD20/12 Clea-Floor with FD20 & Larc-N1) or Larc-C with FD12 Azurion FD20/15 Clea-Floor with FD20 & Larc-N or Larc-C with FD15 Azurion FD20 OR-table Clea-Ceiling stand with FD20 detector Azurion FD12/12 OR-table Poly-G Floor with FD12 & Larc-C with FD12 Azurion FD20/12 OR-table Clea-Floor with FD20 & Larc-C with FD12 Azurion FD20/15 OR-table Clea-Floor with FD20 & Larc-N or Larc-C with FD15 Note 1): Azurion FD20/12 Clea-Floor with FD20 & Larc-N with FD12 is not supported for NPI. This configuration is only supported for the installed base. The following configurations are also supported as remanufactured systems: Table: Configuration Characteristic discriminators Azurion FD12 Poly-G Ceiling stand with FD12 detector and AD7 table Azurion FD20 Clea Floor stand with FD20 detector and AD7 table (no FlexMove) Detailed Configuration requirements Applied symbols: 1) In subsequent tables, additional detailing of configurations uses following abbreviations and indications: Mpl = Monoplane system Ce = Monoplane Ceiling system, suspended by longitudinal rail XY-fa = Monoplane Ceiling system, suspended by double-pivot arm (FlexArm) XY-fm = Monoplane Ceiling system, suspended by X-Y rails (FlexMove) Fl = Floor mounted Monoplane system Bpl = Biplane system (FD12/12 and FD20/12) NB = Neuro Biplane (FD20/15) Mag = Magnus Maquet Tru = Hillrom TruSystem 7500 2) Symbols are used in the tables have following meaning + = included in a system, not an option ++ = twice included in a system, not an option (in case of a Biplane system) # = each system comprises exactly one of the items marked with # in one cell % = each system comprises at least one of the items marked with % in one cell. o = option; is an extension based on extra hardware and/or based on a (software) license key c = compatibility based upon interfacing with the system 3) Use of sub-tags To deal with limited table space, sub-tags are used; to be appended to the specified table tag. Example: Table tag = SRS.Allura.Conf.XperFDSeriesConfigurations; sub-tag = .std-1 Resulting complete tag = SRS.Allura.Conf.XperFDSeriesConfigurations.std-1 Table: sub tag Systems: FD12 FD12/12 FD12 FD12/12 FD12 FD12/12 FD20 FD20 FD20 FD20 FD 15 FD20/12 FD20/15 FD20/12 FD20/15 item Ce Fl Bpl Ce XY fa XY fm Fl Fl Bpl NB notes .std Geo Stand -1 Poly-G2-C # Poly-G2-C is supported for remanufacturing -2 Poly-G3-C # -3 Poly-G2-F + + -4 Clea-2-C + + Clea-2-C is supported for remanufacturing (not XY fm) -5 Clea-3-C + -6 Clea-F # # # # -7 Clea-3-F # Only in combination with AD7 table -8 Larc-N 5) # -9 Larc-C # # # -10 Larc-3-N # Only in combination with AD7 table .pt Patient Table -1 AD7XT # # o o o o # # o o -2 AD7XNT # # o o o # # o o -3 Maquet Magnus c c c c c c -4 Hillrom TruSystem c c See SRS.Allura.Intgr.TruSystem .ms Monitor Ceiling Support -1 Rail suspended MCS % % % % % % % % % -2 Spring-arm boom % % % % % % % % % % -3 Spring-arm boom XL % % % % % % % % % % -4 3rd party boom % % % % % % % % % % MCS Monitor integration kit -5 3rd party boom XL % % % % % % % % % % MCS Monitor integration kit XL .xrc X-ray chain ++= frontal & lateral -1 Tube MRC 200+ 0508 + + ++ +1) +1) -2 Tube MRC 200+ 04072) + + + + + +1) +1) -3 Collimator Nicol3) + + ++ + + + + + ++ ++ -5 AEP dose-meter o o o o o o o o o o AEP calculated by default, optionally hardware based -7 Detector FD124) + + ++ # # -8 Detector FD15 + # # -9 Detector FD20 + + + + # # Notes: 1) Frontal channel MRC0407; lateral channel MRC0508 2) MRC 200+ 0407 ROT-GS1008 for FD20 XY-fa; ROT-GS1004 for all other configurations 3) Nicol v4 on Clea-3, L-arc 3-N and Poly-G3 Ceiling. All other stands Nicol v3 4) Monoplane configurations PX2121S or PX2121C/V; biplane configurations PX2121C/V only 5) Azurion FD20/12 Clea-Floor with FD20 & Larc-N with FD12 is not supported for NPI. This configuration is only supported for the installed base
UNS.SmartSuite.FDPA | Acquire (subtracted and non-subtracted) images from the lower abdomen to the feet for showing a survey of the contrast bolus as it progresses over time (‘bolus tracking’); Perform a single injection run off in Flexible Dynamic Peripheral Angiography (FDPA): Perform multiple runs for subtraction; Review the whole peripheral vasculature in one image (e.g. image stitching); Acquire the FDPA images with the frontal stand or lateral stand in case of PolyG floor frontal stand; New, N.A. | SRS.Allura.Func.RunSubtract ¦ SRS.Allura.Func.FreeInteractiveFDPA ¦ SRS.Allura.Func.BolusChaseReconstruction | Subtracted view is obtained for an acquisition run exhibiting intended movement (e.g. DRA). This requires a license. Subtraction is based on an acquired mask run; the matching run is automatically selected. View-Trace An image sequence is composed into a single image, giving an overview of the filled vessel-tree structure. The composite image can be viewed in subtracted mode. A run is acquired while the user directly controls the longitudinal movement of the table. FDPA requires a license. Optionally a mask run with identical auto-movement profile can be made for subtracted review. Not supported for a TruSystem table with Azurion FlexArm option. Dual-Axis Rotation (DAR) Bolus Chase Reconstruction (BCR) enables to build a single Bolus Chase reconstructed image of the complete leg(s), by automatically ”stitching” successive images that have been acquired with a Bolus Chase protocol. This requires a license. A BCR run shall be made with a specific EPX procedure (Flexible Dynamic Peripheral Angiography (FDPA)). Only one channel is supported in biplane configurations (frontal or lateral channel (lateral channel only in case of a PolyG-Floor frontal stand)). A reconstructed image can be made from a single contrast run or with both a contrast and mask run (subtracted view is then available). BCR shall be started automatically when a new BCR run is acquired. It shall be possible to start reconstruction manually by selecting a previously acquired BCR run. The user shall be able to make more than one reconstruction within one study. It shall be possible to use the reconstructed image as a reference when navigating through the original images. It shall be possible to save the Bolus Chase reconstructed image as a secondary capture image. Basic image processing (brightness, contrast, invert, subtraction on/off and landmarking), annotation and measurement functions shall be supported on the original images. Basic image processing (brightness, contrast, invert, subtraction on/off and landmarking) shall be supported on the reconstructed image. Fixed anatomy navigation mode shall be supported where original images are displayed on the selected patient's fixed anatomy. Not supported for a TruSystem table with Azurion FlexArm option. Marker Functionality
UNS.SmartSuite.FactoriumLogging | Automatically log events that happened while operating the system; Reuse this data for data-mining. New, N.A. | SRS.Allura.Serv.Logging | System information and events are stored and accessible e.g. for diagnostic/faultfinding purposes. Logged info comprises of system-info (e.g. configuration), system usage (e.g. user-actions, system-errors, wear-out indicators), and test results. FSE functions: read (select/filter/search), clear logfile, export logfile. Auxiliary Functions
UNS.SmartSuite.Flagging | Flag images/runs and sort on flagged images; New, N.A. | SRS.Allura.Func.Flagging | The user can flag/un-flag selected images or entire image-runs. This comprises adding a flag marker to the selected images, without implying any further action. Flagging groups the selected images in a set which can be treated as a single entity by other clinical functions. Example: archive the flagged images. Flagging does not imply any subsequent action. Only one type of flags exists. The flagged data is indicated to the user by a flag marker. In case of a biplane image/run the user can flag the images from one or both planes. Note: In combination with e.g. archiving, the flagging mechanism may result in the transference of an incomplete run. As a result, during cine display its dynamic appearance may look peculiar. Digital Viewing Emphasis is in this section on the selection of runs and images and the way this info is displayed. Viewing Navigation
UNS.SmartSuite.FlexMove | Move the stand (instead of the table) in XY direction (lateral & longitudinal movements) to position for imaging; Store and recall multiple XY projections. New, N.A. | SRS.Allura.Func.APCStore/Recall ¦ SRS.Allura.Func.MoveBeamXY | The system shall allow the user to store and recall multiple geometry positions (i.e. position of stand, table, shutters and wedges) including the receptor field size. When a new acquisition study is selected, all stored positions shall be removed. The system shall indicate to the user if the system has not reached a selected position. Note: Table position recall is not supported for TruSystem tables in combination with FlexMove. On AD7X tables recall of longitudinal and lateral table position requires appropriate motors which are optional The user shall be able to move the beam longitudinal and/or transversal (manual and motorized). To select a clinical area for acquisition as an alternative to positioning the tabletop. For standby parking (typically transversal only) during a surgical procedure. For normal parking when the stand is not used.
UNS.SmartSuite.FlexMoveControl | Control combined XY movement with a single hand in an intuitive way; Stay clear of the head-end area for anesthesia; Move from patient left to patient right with the stand, without moving anesthesia equipment; Stay and work in between table and stand in lateral position; Control all geometry movements from a single control panel; Operate the system with a XY joystick with proportional speed-control in concurrent X-Y directions. New, N.A. | SRS.Allura.Conf.UI-RelatedItems ¦ SRS.Allura.Func.MoveBeamXY ¦ SRS.Allura.Conf.Stand-RelatedItems | Table: sub tag item monoplane biplane .xs X-ray switches -1 footswitch 3f + -2 footswitch 6f + -3 wireless footswitch 3f o -4 wireless footswitch 6f o -5 additional footswitch CR 3f o -6 additional footswitch CR 6f o -7 additional footswitch ER 3f o -8 additional footswitch ER 6f o -9 X-ray handswitch + + Table: sub tag item Monoplane Monoplane Monoplane XY (fm/fa) XY (fm/fa) XY (fm/fa) biplane biplane biplane sub tag item AD7 XT AD7 XNT Maq AD7 XT AD7 XNT Maq Tru AD7 XT AD7 XNT Maq ui UI-modules -1 ControlModule NT +1) -2 ControlModule Tilt +1) -3 ControlModule Flex +1) +1) -4 ControlModule ORT o2) o2) +1,3) -5 ControlModule ORT Flex o2) o2) +1,3) -6 ControlModule Geo NT7) +1) -7 ControlModule Geo Tilt7) +1) -8 ControlModule Geo ORT7) o2) o2) + -9 ControlModule Geo NT IBR +1) -10 ControlModule Geo Tilt IBR +1) -11 ControlModule Imaging o4) o4) o4) o4) o4) o4) + + + -12 speed controller o5) o5) o5) o5) o5) o5) -13 pan handle o o o o o o -14 TouchScreenModule +6) +6) +6) +6) +6) +6) +6) +6) +6) -15 review module + + + + + + + + + -16 viewpad (cardio) # # # # # # # # # -17 viewpad (mixed) # # # # # # # # # -18 ControlModule Geo ORT IBR o2) o2) + Notes: 1) Additional ControlModule optional for CR 2) Optional as additional ControlModule on pedestal for ER 3) On pedestal 4) ControlModule imaging optionally available for ER (not in combination with 2)) 5) Required for FDPA, not an option for TruSystem 6) Additional TSM optional for CR and (on pedestal) for ER 7) Not for NPI, applicable for modules already present in the installed base Table: sub item Table type Table type Table type notes tag AD 7XT AD 7XNT Maquet TruSystem .mi miscellaneous -1 TSM swingarm + + + -2 Pedestal o o + trolley to attach UI modules; for AD7X(N)T including Pedestal Wall Conn. Box; for ORT including Surgery Wall Conn. Box -3 Two-way Intercom o o o -4 Keyboard + Mouse in CR + + + in CR; additional second keyboard + mouse in case of additional FlexSpot -5 Mouse in ER o o o Requires a license -7 MultiVision Switch o o o see SRS.Allura.Intgr.Multivision -8 KVM-Switch o o o see SRS.Allura.Intgr.KVM-Switching, requires DVI based video infrastructure -10 MCS Auxiliary box + + + for viewpad-receiver and X-on light on 3rd party booms The user shall be able to move the beam longitudinal and/or transversal (manual and motorized). To select a clinical area for acquisition as an alternative to positioning the tabletop. For standby parking (typically transversal only) during a surgical procedure. For normal parking when the stand is not used. Table: sub Poly-G Poly-G Poly-G Clea Clea Clea Clea L-arc tag Item Ce XY fm Fl Ce XY fm XY fa Fl Bpl notes .std -1 Extended rail o o o o o o FlexArm: 4.3m; standard: 6.0m or 7.8m. -2 L-arm long # # # # -3 L-arm short # # # # -4 Flexarm long # -5 Flexarm short # -6 Ceiling rail covers o See SRS.Allura.Phys.LaminarAirflowCompatibility -7 Optical imaging system o o o see SRS.Allura.Intgr.OpticalImagingSystem -8 Coverbags c c c c c c c c Arc / Tube / Detector / TSM
UNS.SmartSuite.FlexMoveDesign | Prevent interference with monitors/lamps as much as possible (integrated cabling); Obtain a high level of sterility/hygiene in a Hybrid OR (e.g. minimum open cabling at monitor connection); Minimize the interruption of laminar airflow caused by the system. New, N.A. | SRS.Allura.SLS.patient-contamin-airborne ¦ SRS.Allura.SLS.patient-contamination ¦ SRS.Allura.Func.MoveBeamXY | The system shall be designed to minimize its contribution to airborne contaminants reaching the patient. The system shall be designed for: Cleaning and disinfection of system parts in the patient environment. Creating an aseptic environment. The user shall be able to move the beam longitudinal and/or transversal (manual and motorized). To select a clinical area for acquisition as an alternative to positioning the tabletop. For standby parking (typically transversal only) during a surgical procedure. For normal parking when the stand is not used.
UNS.SmartSuite.FlexMovePark | Park the stand everywhere where the user wants to, within the system’s limitations. New, N.A. | SRS.Allura.Func.ParkorStandbyBeam ¦ SRS.Allura.Func.MoveBeamXY | Enables the user to move the frontal and/or lateral stand out of the way to a standby or park position to improve access to the patient if the beam is temporarily not used. The different ways to move the stand to a park / standby position depend on the particular configuration (examples: lateral movement for Flexarm standby, longitudinal movement in ceiling stands, swing stand base in a Poly-G floor stand). The user shall be able to move the beam longitudinal and/or transversal (manual and motorized). To select a clinical area for acquisition as an alternative to positioning the tabletop. For standby parking (typically transversal only) during a surgical procedure. For normal parking when the stand is not used.
UNS.SmartSuite.FlexSpot | Seamlessly control (view and manipulate) all viewports within a single view and with a single mouse and keyboard; Quickly access applications without the need to switch between applications and without different mouse control; Fully customize the layout (e.g. which applications on screen, size of viewports, number of viewports); Create presets, adjust these on-the-go and quickly save changes; Organize presets in groups for easy storage and retrieval; Display any (3rd party) system; Connect at least 11 external video signals (that are not generated directly by the X-ray system); Choose from at least 4 different FlexSpot configurations to accommodate the hospital workflow (single work spot with single monitor, single work spot with two monitors, two work spots with one monitor each, two work spots (one with two monitors and the other with one monitor); Reduce the footprint of display monitors in the Control Room; Change the layout of the FlexVision (Pro) from the Control Room; Create a snapshot of a viewport/all displayed video source with one click, export it as secondary capture and store the snapshot in the patient file; Visualize the references monitors in separate window (parallel to live). New, Have flexibility in positioning workspots (e.g. between CR/ER and multiple workspots in the CR); Layouts can be changed during the procedure without having to go into the system settings; | SRS.Allura.Intgr.Flexspot.KeyboardMouse ¦ SRS.Allura.Intgr.VideoInputInterface ¦ SRS.Allura.Intgr.VideoSourcesMaintainIQ ¦ SRS.Allura.Intgr.OpticalImagingSystem ¦ SRS.Allura.Intgr.Flexspot.AdditionalWorkspot ¦ SRS.Allura.Intgr.Flexspot ¦ SRS.Allura.Conf.Control-RoomDisplays ¦ SRS.Allura.Intgr.AuxiliarySystems ¦ SRS.Allura.Intgr.Flexspot.SecondFlexspot ¦ SRS.Allura.Intgr.VideoDisplayInterface | With FlexSpot, the system shall allow for remote control of auxiliary systems via keyboard and mouse. The following types of video interfaces shall be supported for connecting auxiliary systems: DP (only in combination with IP based video infrastructure) DVI-I including non-interlaced VGA The system shall provide set of approved measures to display video and/or optimize image quality of compatible auxiliary system video source, when displaying it on the system’s monitors or providing for use outside the system. The system shall optionally provide an optical imaging system to support auxiliary systems (e.g. compatible surgical navigation systems). The optical imaging system consists of video cameras positioned around the detector. The video stream of each camera is available for auxiliary systems. Parameters of the cameras are accessible and controllable by auxiliary systems. Peripheral Devices A second workspot (keyboard, mouse, single monitor) next to FlexSpot shall be offered, which can be a workspot in the control room or exam room. One video source can be selected to be shown. Video can originate from the system itself and from auxiliary systems (configurable set, e.g. Interventional Workspot, IVUS). The additional FlexSpot requires a license. Provides the user with the same features as for FlexVision in the control room. Exception is that FlexSpot doesn't support display of 2k2 images. In combination with VoIP based infrastructure, FlexSpot requires a license. A second monitor for the FlexSpot workspot shall be offered. This second monitor requires a license. Table: sub tag item DVI based infra VoIP based infra notes . crd control room displays -1 acquisition and review monitor # # Both 24" color monitors with native image resolution of 1920x1080 and aspect ratio 16:9. -2 FlexSpot primary monitor # # Instead of acquisition and review monitor, a 27" color monitor with native image resolution of 2560x1440 and aspect ratio 16:9. -3 FlexSpot secondary monitor o o Requires FlexSpot primary monitor, 27" color monitor with native image resolution of 2560x1440 and aspect ratio 16:9. -4 additional FlexSpot o Requires FlexSpot primary monitor, 27" color monitor with native image resolution of 2560x1440 and aspect ratio 16:9; shall be placed either in the control or exam room. -5 slave monitor1) o o 24” (desk version) or 27” (boom version) or 32" (boom version) color monitor with native image resolution of 1920x1080 and aspect ratio 16:9, e.g. extra live monitor. -6 MultiView o Requires acquisition and review 24" color FHD monitors, -7 Additional MultiView o Requires acquisition and review monitor, extra 24" color FHD monitor, mouse and keyboard; -8 Second FlexSpot primary monitor o Requires FlexSpot primary monitor, 27" QHD color monitor, mouse and keyboard; shall be placed either in the control or exam room. -9 Second FlexSpot secondary monitor o Requires Second FlexSpot primary monitor, 27" color QHD monitor; shall be placed at the same location as Second FlexSpot primary monitor -10 3rd MultiView o Requires extra 27" color QHD monitor, mouse and keyboard. Only in combination with FlexSpot, second FlexSpot, and 4th MultiView -11 4th MultiView o Requires extra 27" color QHD monitor, mouse and keyboard. Only in combination with 3rd MulitView Note: 1) 32" monitors are not part of NPI. With FlexSpot it shall be possible to install dedicated Philips auxiliary systems in the cabinet in the technical room. FlexSpot shall provide infrastructure to gain access from control room to the media devices available on these PCs; on DVI based video configurations the maximum number of PCs from dedicated Philips auxiliary systems installed in the cabinet in the technical room is three. on VoIP based video configuration the maximum number of PCs from dedicated Philips auxiliary systems installed in the cabinet in the technical room is four. The maximum number of RTO clients in the system (TR/CR/ER) is five for VoIP based configurations, and six for DVI based configurations. Without FlexSpot it shall be possible to install a single PC from dedicated Philips auxiliary systems in the cabinet in the technical room. However no further infrastructure except power and network access is provided. A second workspot (keyboard, mouse, one or two monitors) next to FlexSpot shall be offered, which can be a workspot in the control room or exam room. It shall provide the user with the same features as FlexSpot. Second FlexSpot requires a license. A second monitor for the second workspot shall be offered. This second monitor requires a license. The system shall support the display of its own video sources and/or sources provided by auxiliary systems on its examination/control room monitors and large screen.
UNS.SmartSuite.FlexVision(Pro) | Define and customize the ER monitor configuration and layout of images on screen; Add and connect a second/backup screen for live/heamo when a single large monitor is used; Simultaneously see: Live X-ray and Ref X-ray in same size; Heamo in proper size, resolution, non-scaled, pixel for pixel; Scaled 3rd party/Interventional Tools/QA input (optional and switchable, based on operator request); Start the system with Heamo in proper size/resolution/non scaled/pixel for pixel and have the option to scale the Heamo back and up again; Simultaneously see (optional) 3rd party monitors in addition; Rescale/rezoom display container to its demand; Seamlessly control (view and manipulate) all viewports at table side with a single means / mouse; Fully customize the layout (e.g. which applications on screen, size of viewports, number of viewports, position of the status bar) using the table side mouse, TSM or from the FlexSpot; Create presets, adjust these on-the-go and quickly save changes; Organize presets in groups for easy storage and retrieval; Display any (3rd party) system; Connect at least 11 external video signals (that are not generated directly by the X-ray system; View X-ray images in native resolution without downscaling; View external sources in their native resolution to prevent signal distortion due to scaling (up/down); Create a snapshot of a viewport/all displayed video source with one click, export it as secondary capture and store the snapshot in the patient file; Fully customize the layout (e.g. which applications on screen, size of viewports, number of viewports, position of status bar) using the table side mouse, TSM or from the FlexSpot; Visualize two full HD video sources in original pixel format, at the same time, mainly for EP procedures. New, Improved image quality when downscaling EP sources. | SRS.Allura.Intgr.Interfaces.RealTimeImageLink ¦ SRS.Allura.Intgr.Flexspot.KeyboardMouse ¦ SRS.Allura.Intgr.VideoInputInterface ¦ SRS.Allura.Intgr.VideoSourcesMaintainIQ ¦ SRS.Allura.Intgr.OpticalImagingSystem ¦ SRS.Allura.Intgr.Interfaces.ControlFromAuxiliarySystems ¦ SRS.Allura.Intgr.OpenVideoInterface.LargeScreen ¦ SRS.Allura.Intgr.NativeIQ ¦ SRS.Allura.Intgr.Flexspot ¦ SRS.Allura.Intgr.OpenVideoInterface.LargeScreenDownscaled ¦ SRS.Allura.Intgr.Flexvision ¦ SRS.Allura.Intgr.VideoDisplayInterface ¦ SRS.Allura.Intgr.Flexvision.KeyboardMouse ¦ SRS.Allura.UI.ComfortThemes ¦ SRS.Allura.Intgr.Flexvision.Snapshot ¦ SRS.Allura.Intgr.Interfaces.ControlAuxiliarySystems | The system shall provide communication interfaces to provide x-ray images of the frontal channel in real time to compatible auxiliary equipment during x-ray acquisition. This requires a license. With FlexSpot, the system shall allow for remote control of auxiliary systems via keyboard and mouse. The following types of video interfaces shall be supported for connecting auxiliary systems: DP (only in combination with IP based video infrastructure) DVI-I including non-interlaced VGA The system shall provide set of approved measures to display video and/or optimize image quality of compatible auxiliary system video source, when displaying it on the system’s monitors or providing for use outside the system. The system shall optionally provide an optical imaging system to support auxiliary systems (e.g. compatible surgical navigation systems). The optical imaging system consists of video cameras positioned around the detector. The video stream of each camera is available for auxiliary systems. Parameters of the cameras are accessible and controllable by auxiliary systems. Peripheral Devices The system shall provide communication interfaces to allow compatible auxiliary systems to select EPX procedures which may or may not be specific for the auxiliary system (i.e. they are selectable only by the auxiliary system) set APC target positions lock/unlock the table set specific acquisition parameters (e.g. injector coupling, patient orientation, biopsy mode, etc.) push a secondary capture image to a reference viewport The system shall provide a large screen video output (DP1.2 or, in combination with DVI based video infrastructure, 2*DL-DVI) that contains a copy of the large screen of FlexVision. In VoIP infrastructure, this requires a license. When the video signal is not displayed on the system monitor the system shall be able to offer a transparent image path such that it can be used by a 3rd party system. Transparent image path between the system input and output is a point to point 3rd party video signal distribution that doesn’t make any change to the signal being distributed. Auxiliary Systems Data Exchange & Control Provides the user with the same features as for FlexVision in the control room. Exception is that FlexSpot doesn't support display of 2k2 images. In combination with VoIP based infrastructure, FlexSpot requires a license. A second monitor for the FlexSpot workspot shall be offered. This second monitor requires a license. The system shall provide a downscaled copy of the large screen of FlexVision video output in FHD resolution (DP is the interface in combination with IP based video infrastructure and SL-DVI in combination with DVI based video infrastructure). In combination with IP based video infrastructure that video output shall be also available (configurable) for switching on Flexspot, (Additional,3rd, and 4th) Multiview and Switchable Monitors. In VoIP infrastructure, the availability of the downscaled copy requires a license. FlexVision enables flexible viewing for displaying multiple video sources (live images but also video from other systems) on one large screen display in the exam room with which the user can determine the layout, the position and size of multiple video sources on one large screen display in the exam room. Layouts are predefined. A layout divides the large screen in viewports. The user shall be able to select a preset (a preset contains the layout and for each viewport the specific video source shown in the viewport) to be applied on the large screen. The user shall be able to create, delete, change, save and rename a preset. The user shall be able to select which source is displayed in which viewport. The user shall be able to hide/unhide the video source display in a viewport. The user shall be able to swap the video sources of two viewports in a preset. The user shall be able to enlarge a specific viewport and restore it to its original size. The user shall be able to display video sources in native resolution. Display of 2k2 images shall be supported. Video can originate from the system itself and from auxiliary systems (configurable set, e.g. Interventional Workspot, IVUS). In combination with VoIP based infra, FlexVision requires a license. The display of 2K2 images requires an additional license. The system shall support the display of its own video sources and/or sources provided by auxiliary systems on its examination/control room monitors and large screen. With FlexVision, the system shall allow for remote control of auxiliary systems via on-screen keyboard and touchpad (on TSM) and / or wireless mouse. This requires a license. With FlexVision, the user shall be able to play movies or image slideshows to relax the patient. This requires a license. With FlexVision, a single viewport or all viewports at once can be grabbed and stored as image under the current acquisition study. The system shall provide communication interfaces to control compatible auxiliary systems from the systems user interface with wireless mouse on the auxiliary system’s graphical UI displayed on FlexVision with mouse and keyboard on the auxiliary system’s graphical UI displayed on Flexspot from TSM by a generic button based UI tailored by the auxiliary system or by a graphical UI streamed from the auxiliary system including (multi-)touch control. This requires a license control shall be established automatically when (mobile) compatible auxiliary system is connected
UNS.SmartSuite.FlexibleAngles | To position FlexArm C-arc in any angle for 2D imaging between -135 and +135 degrees with respect to the standard table position to facilitate access to the patient New, N.A. | SRS.Allura.Func.Z-RotateStandBase(L-ArmRotation) ¦ SRS.Allura.Func.PatientOrientedMovement ¦ SRS.Allura.Func.RotateFDXDDetector | Enables the user to rotate the frontal stand for position at the side or head of the table for optimized patient imaging. Iso-center stand rotation around a vertical axis to accomplish improved access to the patient. Function does not apply to the Poly-G floor stand. In FlexArm (XY-fa) and biplane (Clea3 & L-arc3 only) configurations, if the user angulates or rotates the frontal beam, the detector and collimator angles shall be adapted so to keep the same orientation towards the patient. Enables the user to rotate the frontal FD20 detector for optimized imaging (manual and motorized). User shall be able to rotate the detector by 90° in order to switch between landscape and portrait usage of the detector. Beam width is adapted to maintain the selected detector field size. User shall be able to rotate the detector by any angle between 0° and 90° in order to align with specific anatomy. In FlexArm (XY-fa) and biplane (Clea3 & L-arc3 only) configurations, the collimator rotates synchronously to maintain the selected detector field-size. In all other configurations, the collimator is fixed and the beam-width is limited such that the irradiated area is completely visible.
UNS.SmartSuite.FlexibleViewing | Select from the available image sources which ones will be displayed on which part of the screen (viewports); and save this display; Change the display layout and rearrange the position and size of multiple viewports; and save this display; Swap/move keyboard and mouse focus over selected viewports; Enable/disable whether the display will be visible in the ER or not; Connect a dedicated keyboard to an external workstation modality (e.g. Heamo and/or EP). New, Have flexibility in scaling the size of multiple sources; Have flexibility in displaying different sources on different monitors; Show multiple sources with different resolutions on Philips flexible monitors (e.g. Flexvision, Flexspot). | SRS.Allura.Intgr.Flexspot.KeyboardMouse ¦ SRS.Allura.Intgr.Flexspot.AdditionalWorkspot ¦ SRS.Allura.Intgr.Flexspot ¦ SRS.Allura.Intgr.Flexvision ¦ SRS.Allura.Intgr.Flexvision.KeyboardMouse ¦ SRS.Allura.Intgr.VideoDisplayInterface | With FlexSpot, the system shall allow for remote control of auxiliary systems via keyboard and mouse. A second workspot (keyboard, mouse, single monitor) next to FlexSpot shall be offered, which can be a workspot in the control room or exam room. One video source can be selected to be shown. Video can originate from the system itself and from auxiliary systems (configurable set, e.g. Interventional Workspot, IVUS). The additional FlexSpot requires a license. Provides the user with the same features as for FlexVision in the control room. Exception is that FlexSpot doesn't support display of 2k2 images. In combination with VoIP based infrastructure, FlexSpot requires a license. A second monitor for the FlexSpot workspot shall be offered. This second monitor requires a license. FlexVision enables flexible viewing for displaying multiple video sources (live images but also video from other systems) on one large screen display in the exam room with which the user can determine the layout, the position and size of multiple video sources on one large screen display in the exam room. Layouts are predefined. A layout divides the large screen in viewports. The user shall be able to select a preset (a preset contains the layout and for each viewport the specific video source shown in the viewport) to be applied on the large screen. The user shall be able to create, delete, change, save and rename a preset. The user shall be able to select which source is displayed in which viewport. The user shall be able to hide/unhide the video source display in a viewport. The user shall be able to swap the video sources of two viewports in a preset. The user shall be able to enlarge a specific viewport and restore it to its original size. The user shall be able to display video sources in native resolution. Display of 2k2 images shall be supported. Video can originate from the system itself and from auxiliary systems (configurable set, e.g. Interventional Workspot, IVUS). In combination with VoIP based infra, FlexVision requires a license. The display of 2K2 images requires an additional license. With FlexVision, the system shall allow for remote control of auxiliary systems via on-screen keyboard and touchpad (on TSM) and / or wireless mouse. This requires a license. The system shall support the display of its own video sources and/or sources provided by auxiliary systems on its examination/control room monitors and large screen.
UNS.SmartSuite.FluoroIndicators | Switch on/off the volume of sounds for acquisition and fluoroscopy; Stop the 5 minute fluoro time buzzer; Indicate visually if actual X-ray is given or initiated e.g. “X-ray on” indicator light). | SRS.Allura.Func.IndicateRadiationOn ¦ SRS.Allura.Func.FluoroTimeBuzzer ¦ SRS.Allura.Func.RoomLightAndOperationLamp | During X-ray generation this is continuously indicated to the user. X-ray-on light in and outside the exam room; indications on the UI in exam room and control room. Continuous audible signal (fluoro/exposure buzzer); can be configured off. Warns the user after an elapsed accumulated fluoro time of 5 minutes: This safety mechanism can be customized on/off, dependent on the local needs and safety regulations. The user can deactivate the buzzer. Configurable at installation, the room-light and/or operation lamp can be switched or dimmed by a dedicated foot pedal, or synchronized to acquisition.
UNS.SmartSuite.FreeEndPositions | Move the FlexArm system from/to working position without interference with equipment/persons located at all sides of the table (e.g. anaesthesiologists). New, N.A. | SRS.Allura.Func.RotateBeam ¦ SRS.Allura.Func.ImageBeamRotation | The user shall be able to rotate the beam around the longitudinal-table-axis (motorized). Intuitive control and display orientation rules (see SRS.Allura.Func.DisplayOrientation) also apply if rotation causes a reversed beam direction. In biplane systems, the frontal and lateral beam can be rotated simultaneously. During motorized movement the rotation position shall be displayed. In FlexArm (XY-fa) and biplane (Clea3 & L-arc3 only) configurations, the frontal beam shall be rotatable i.e. collimator and detector rotate simultaneously and stay continuously aligned (IBR).
UNS.SmartSuite.FromStandbyToWorkingPosition | For FlexArm configuration return to the last X-ray Position from a Standby Position for any procedure while maintaining previous registration on Philips clinical applications (e.g. Interventional Workspot, EchoNav and Coronary tools. Example 3D roadmap, VesselNavigator overlay. Assuming patient and table remain in same position). New, Have the largest lateral standby distance on the market while maintaining 3D registration. | SRS.Allura.Func.XperCT ¦ SRS.Allura.Func.ParkorStandbyBeam ¦ SRS.Allura.Func.DRAFor3D ¦ SRS.Allura.Func.ClassicDRA ¦ SRS.Allura.Func.DualPhaseXperCT | This is a special case of DRA or DAR for 3D enabling reconstruction of soft-tissue. This function is only supported with an FD20 detector. The system shall offer dedicated EPX procedures and related settings for Neuro (DRA and DAR) and Abdomen (DRA only) applications. The system shall perform acquisition with up to 60 images/second to obtain acceptable scan time. The system shall be able to apply slit collimation to reduce patient dose. Enables the user to move the frontal and/or lateral stand out of the way to a standby or park position to improve access to the patient if the beam is temporarily not used. The different ways to move the stand to a park / standby position depend on the particular configuration (examples: lateral movement for Flexarm standby, longitudinal movement in ceiling stands, swing stand base in a Poly-G floor stand). DRA for 3D is a special case of DRA performed at predefined positions, to enable 3D reconstruction. Settings, including start- and end-positions, shall be defined in EPX. The system shall support DRA for 3D at least at the following positions: for FlexArm (XY-fa) configuration at head, doctor and nurse side for other FD20 and FD15 configurations at head and one side position (doctor or nurse) for FD12 only at head position SID is maximized to allow for 3D scans of obese patients. The user shall define the end- and start-position of the DRA scan, implicitly executing a test-scan to detect obstructions. The system shall be able to perform the DRA scan at the side or head of the table (C-arm roll and propellor respectively). DRA acquisition is only possible for the frontal channel. The user shall be able to acquire a second run to enable subtracted reviewing. Settings, including start- and end-positions, shall be defined in EPX, but if allowed in EPX, the user shall be able to choose another end- and/or start-position. DRA requires a license. This is a special case of XperCT where two consecutive DRA scans are being made by executing the scans in different phases of contrast propagation, which is known as Dual phase XperCT. The two scans are performed in opposite direction to reduce the overall procedure time to minimize the amount of contrast medium. For neuro applications, to allow the use of intravenous contrast injections in Dual phase XPerCT scans, the system shall provide a subtracted fluoroscopy flavor with enhanced contrast. It shall enable to detect the arrival of contrast agent in order to start the scan. This application shall be executable from the Control Room. Flexible Dynamic Peripheral Angiography (FDPA)
UNS.SmartSuite.FullBodyCoverage | Image (2D) the entire body with FlexArm from foot to head from both nurse and doctor sides of the table, without the need to move the patient table combination*. *not applicable when using a leg support board New, N.A. | SRS.Allura.Func.MoveBeamXY | The user shall be able to move the beam longitudinal and/or transversal (manual and motorized). To select a clinical area for acquisition as an alternative to positioning the tabletop. For standby parking (typically transversal only) during a surgical procedure. For normal parking when the stand is not used.
UNS.SmartSuite.HCA | Use a system that complies with the security requirements from HCA. New, N.A. | SRS.Security.HN.RemovableMediaBreached ¦ SRS.Security.RM.AttackViaMedicalDevice ¦ SRS.Security.MD.DataBreach ¦ SRS.Security.MD.AttackViaPhysical DeviceModification ¦ SRS.Security.MD.HospitalNetwork ¦ SRS.Security.PHD.PatientDataLoss ¦ SRS.Security.MD.RemoteAccessPoint ¦ SRS.Security.PHD.DataBreach ¦ SRS.Security.PHD.PhysicalAccess ¦ SRS.Security.PHD.MedicalDevice ¦ SRS.Security.PHD.PatientDataCorrupted ¦ SRS.Security.MD.PhysicalAccess ¦ SRS.Security.MD.ReputationDamage ¦ SRS.Security.MD.MalwareOnRemovable media ¦ SRS.Security.PHD.HospitalNetwork ¦ SRS.Security.PHD.PhysicalDismounting ¦ SRS.Security.HN.HospitalNetworkBreached ¦ SRS.Security.PHD.RemoteAccessPoint ¦ SRS.Security.MD.UnavailabilityOfMedicalDevice ¦ SRS.Security.HN.AttackViaMedicalDevice | The Azurion shall reduce the impact of: Breached Devices, Disclosed Patient Health data, Unavailability of data, in case Removable Media is breached through the Azurion. Standards, Legal Requirements International Standards The Removable Media shall be protected against attacks via Azurion. The Azurion shall reduce the impact of Data Breach in case Azurion is breached. The Azurion shall be protected against attacks via the hospital network. The Azurion shall reduce the impact of Reputation damage (manufacturer/hospital) in case Azurion is breached. The Azurion shall be protected against attacks via a remote access point. The Azurion shall reduce the impact of Reputation damage (manufacturer/hospital) in case Azurion is breached. The Patient Health Data shall be protected against attacks via physical access. The Patient Health Data shall be protected against attacks via the Azurion. The Azurion shall reduce the impact of Patient data corrupted in case of Patient Health Data is attacked. The Azurion shall be protected against attacks via physical access. The Azurion shall reduce the impact of Reputation damage (manufacturer/hospital) in case Azurion is breached. The Azurion shall be protected against attacks via malware on removable media. The Patient Health Data shall be protected against attacks via hospital network. The Patient Health Data shall be protected against attacks via physical dismounting. The Azurion shall reduce the impact of: Unavailability of Hospital Network, Breached hospital infrastructure, Disclosed customer data, in case Hospital Network is attacked through the Azurion. The Patient Health Data shall be protected against attacks via remote access point. The Azurion shall reduce the impact of Unavailability of Azurion in case Azurion is breached. The Hospital Network shall be protected against attacks via Azurion.
UNS.SmartSuite.Handover | Handover evidence indicating that the system is operational according to manufacture specification after any service activity; Get proof that the customer accepts the system in the image-guided lab. | SRS.Allura.Serv.AcceptanceTestReport ¦ SRS.Security.MD.RemoteAccessPoint ¦ SRS.Allura.Serv.SystemInformation.Export ¦ SRS.Allura.Serv.SystemID ¦ SRS.Allura.Serv.SiteLog | The system shall be able to generate an acceptance test report that covers all necessary tests for a formal handover to the customer. Diagnostics, Faultfinding Faultfinding is used in corrective maintenance to localize a problem and to determine follow-up actions. The Azurion shall be protected against attacks via a remote access point. The FSE shall be able to export the system information, targeting: Formatted report, containing information to uniquely identify the system. Local copy according standard Common Data File format. Remote Service Functionality Remote Service General Requirements The system shall be uniquely identified. A service report shall contain the system ID and hospital name. Non-Functional Requirements Upgrading Related The FSE shall be able to use an on-system logbook to register the systems history and comments by the FSE.
UNS.SmartSuite.HospitalITConformance | To bring the hospital infrastructure and associated network protocols up-to-date and connect the system to the updated configuration. New, N.A. | SRS.Allura.Func.SystemCustomization | Allows the user to customize system settings, that are available for the system's configuration, such as: Date, Time and Regional Settings User Administration and Administration settings List of annotation and performing physician Procedure cards and RIS code mapping Settings for viewing and quantitative analysis Media and DICOM export settings APC positions Linking DICOM export settings, APC positions and procedure help text to EPX procedures Remote support settings Defining APC pathways Saved settings are persistent over system restarts. System customization settings can be imported and exported. Administration Functions This section focuses on administrative functions to deal with patient data and related information. General Administration Concepts A local database includes all information that is entered manually and generated in acquisition. It can contain multiple patients/studies/procedure steps comprising patient data, images, etc. Local Administration Functions
UNS.SmartSuite.IQ | Get images that can be used by the physician to perform procedures. | SRS.Allura.Func.X-rayImageProcessing ¦ SRS.Allura.Func.AcquisitionMethods | Image processing is applied to enhance the diagnostic value and the user's appreciation of the images. Processing is EPX controlled to allow optimized images for each type of procedure and each user. Fluoro processing depends on the selected fluoro-flavour. The processing algorithm ClarityIQ is optionally available. The option OncoBlue enables ClarityIQ image processing restricted to abdominal examinations only. ClarityIQ and OncoBlue both require a separate license. Common Functionality For Fluoro And Exposure Image generation (also called acquisition) can be done in a number of ways: Fluoro: intended to provide a real-time X-ray image. Exposure: intended for real-time diagnostic imaging and documentation purposes. Compared to fluoro a higher signal to noise ratio (S/N) is necessary for diagnostic imaging and the images are always stored, enabling reviewing.
UNS.SmartSuite.ISPWorkflowIntegration | Exchange data between Azurion and IntelliSpace Portal. New, Automatically launch relevant images on IntelliSpacePortal* ; Prefetch of data upon selection of the patient acquisition on Azurion. *Available from ISP R8 onwards. | SRS.Allura.Intgr.Interfaces.DataExchange | The system shall provide communication interfaces to exchange data with compatible auxiliary systems Patient and examination related data Acquisition related data Dose data ECG data Geometry related data
UNS.SmartSuite.ImageEnhancement | Adjust image to enhance image review, e.g.: Zoom/pan the displayed image; Change image contrast, brightness and/or amount of edge enhancement; Invert image polarity on display (e.g. invert black vessels to white). New, N.A. | SRS.Allura.Func.VideoInvert ¦ SRS.Allura.Func.PanZoom ¦ SRS.Allura.Qual.ClinicalResponseTimes.Review ¦ SRS.Allura.Func.SetImageEnhancement ¦ SRS.Allura.Func.OverrideManualAdjustments | Changes the polarity of the image on display, e.g. black bones/vessels to white, and vice-versa. Image Presentation Enables the user to zoom the image on display and pan the zoom-center. The user can determine the zoom factor. The user can reset panning towards the image center. Annotation & Overlay Text Table: Categories/ items Max. notes: (all values in seconds ) Image stepping 0.2 Allowed 0.4 sec for run-subtract. Run stepping, non-subtract 0.3 Run stepping, subtract 0.6 Run overview 4*4, subtract & non-subtract 2 Run overview 4*4, run-subtract 4 File overview 4*4, non-subtract 2 File overview 4*4, subtract 4 Reviewing requirements are not applicable to mosaic overview modes, unless explicitly indicated otherwise. The user can interactively change image contrast, brightness and/or amount of edge-enhancement to adapt the subjective image quality to personal preference. This function resets settings (if altered during postprocessing) to “as during acquisition”. The user can select the scope of this function to all postprocessing settings, or a defined subset.
UNS.SmartSuite.ImageOffCenter | Position the FlexArm gantry centre laterally over at least 52 cm with respect to the table longitudinal centre line without moving staff or equipment which is positioned next to the table or the patient table combination; Create workspace, at least 118 cm, between the FlexArm stand centre and table longitudinal centre line for example to stay between stand and table, maintaining sterility of the user. New, Perform radial access procedures on both sides of the table. | SRS.Allura.Func.MoveBeamXY | The user shall be able to move the beam longitudinal and/or transversal (manual and motorized). To select a clinical area for acquisition as an alternative to positioning the tabletop. For standby parking (typically transversal only) during a surgical procedure. For normal parking when the stand is not used.
UNS.SmartSuite.ImageOverlay | Adjust post-processing parameters of the image; at image level, at run level and for fluoroscopy; Add portion of background information to processed image (e.g. Contrast Brightness and Edge enhancement (CBE), Landmark, ViewTraced, video invert etc.) Add/remove/adjust/configure/view/save/export the following data/ information to images: Relevant physiology data (e.g. ECG signals); Image acquisition related information (e.g. detector size, run time stamp); Annotations (predefined as well as added manually). Customize the default information shown on the live and review image. New, N.A. | SRS.Allura.Func.NormalSubtraction ¦ SRS.Allura.Func.PredefinedAnnotation ¦ SRS.Allura.Func.ChangeFluoroProcessing ¦ SRS.Allura.Func.StoreToMedia ¦ SRS.Allura.Func.SetImageEnhancement ¦ SRS.Allura.Func.SelectProcessingScope ¦ SRS.Allura.Func.ManualAnnotation ¦ SRS.Allura.Func.SelectPhysioDisplay ¦ SRS.Allura.Func.OverlayText ¦ SRS.Allura.Func.Tracing ¦ SRS.Allura.Func.DICOM.Export ¦ SRS.Allura.Func.Roadmap.DevicePhase.Processing ¦ SRS.Allura.Qual.ClinicalResponseTimes.Review ¦ SRS.Allura.Func.VideoInvert ¦ SRS.Allura.Func.CO2Tracing ¦ SRS.Allura.Func.OverrideManualAdjustments | Minimal movement in acquisition is assumed, to obtain enhanced view with single mask subtraction. Enables that the user can step through the run to see the progressive filling of arteries or veins. The user can switch subtraction on/off, and/or select another mask from the run. Manually, registration errors due to patient movements can be corrected ("Pixel-shift" function). A background portion can be added manually for orientation purposes ("Landmarking" function). Subtracted imaging requires a license. During manual annotation, a list can be displayed containing predefined and user defined text strings, which can be inserted within the current annotation block. The user shall be able to customize this list: The predefined annotations can be deleted and modified. The user defined annotations can be added, deleted and modified. Enables the user to change processing during LIH or Auto-cycle (e.g. contrast/brightness/edge-enhancement). The changes remain active until the same EPX application, another EPX application, patient type, performing physician or procedure step for acquisition is selected, or roadmap is activated. Exposure Related Functions This function enables to store image data on media (CD/DVD/USB). The user selects DICOM format or PC format. With DICOM format, the user can select to store the DICOM viewer with the exported DICOM files. With PC format, the user determines the file name for creating a PNG (Portable Network Graphic) file or an MPEG4 (Moving Picture Experts Group) file. The content can be deidentified and the user can define a fictive patient name for the deidentified content. The user can interactively change image contrast, brightness and/or amount of edge-enhancement to adapt the subjective image quality to personal preference. The user shall be able to set the scope of the interactive post-processing operations within a run: the scope may be set to either the whole run or to the active image for annotation; the scope may be set to either the whole run, to a selection of images within the run, or to the active image for image enhancement, video invert, pan, zoom, and landmarking; the scope may be set to either the whole run, to a selection of images within the run, the active image, active image and all preceding images in the run, active image and all following images in the run for pixel-shift; for all other operations the scope is always the whole run. Image Enhancement This enables the user to add text and/or shapes on a certain position into the image on display. Annotations can be repositioned, modified, hidden, copied and deleted. Annotations can be added to x-ray images and as well on secondary capture images. This function enables to switch on/off physio data as graphics overlay. The correct time relation between images and physio on display must be preserved. In case of multiple physio sequences, the user can select one recorded channel for display. Dynamic Viewing Overlay text comprises image/run-related information and is shown by default (e.g. number, geo-angles). Non-essential overlay text can be switched-off by the user to unclutter the display. Interactively, images can be added to provide a progressive build-up of the overview. In order to exclude some images from the trace operation, skip and undo functions are provided. The resulting trace image can be stored for later use (e.g. as SmartMask). Subtracted imaging requires a license. Enables to export (part-of) a study to a remote DICOM workstation or to a DICOM archive. Auto-export at the end of run and/or finish study, according EPX defined destination (multiple) and contents (e.g. images, downscaling). The user shall be able to get an overview of which runs will be auto-exported upon finishing the study and to which destination. For manual export, the user can select the content and the destination. For manual export, the content can be de-identified and the user can define a fictive patient name for the de-identified content. The selected destination determines the image processing format (raw or processed) which will be used. A biplane run is exported as two separate DICOM objects that refer to each other. On Last-Image-Hold the user can change settings interactively, also for the next fluoro run: Regular processing functions, e.g. pixel-shift, can be applied to obtain a better match. The intensity of the VesselMask and Device image can be altered independently (called ‘Fading’). Table: Categories/ items Max. notes: (all values in seconds ) Image stepping 0.2 Allowed 0.4 sec for run-subtract. Run stepping, non-subtract 0.3 Run stepping, subtract 0.6 Run overview 4*4, subtract & non-subtract 2 Run overview 4*4, run-subtract 4 File overview 4*4, non-subtract 2 File overview 4*4, subtract 4 Reviewing requirements are not applicable to mosaic overview modes, unless explicitly indicated otherwise. Changes the polarity of the image on display, e.g. black bones/vessels to white, and vice-versa. Image Presentation If CO2 contrast medium is used, vessels appear to be white, so in this case it is required to trace maximum instead of minimum pixel values. This function is compatible with all other trace options. Default selection is EPX defined, but can be overruled by the user. CO2 tracing requires a license. Bolus Chase Reconstruction This function resets settings (if altered during postprocessing) to “as during acquisition”. The user can select the scope of this function to all postprocessing settings, or a defined subset.
UNS.SmartSuite.ImageSubtraction | Make vessels visible without disturbing background; at image level, at run level and during fluoroscopy; Toggle between subtracted and un-subtracted display. New, N.A. | SRS.Allura.Qual.ClinicalResponseTimes.Acquisition ¦ SRS.Allura.Func.Roadmap.DevicePhase.VesselSubtract ¦ SRS.Allura.Func.NormalSubtraction ¦ SRS.Allura.Func.Roadmap.DevicePhase.LiveSubtract ¦ SRS.Allura.Func.Roadmap.DevicePhase.DeviceMask ¦ SRS.Allura.Func.ExposureSubtract ¦ SRS.Allura.Func.Roadmap.DevicePhase.Subtraction ¦ SRS.Allura.Func.AcquisitionMaskAveraging ¦ SRS.Allura.Func.RunSubtract | Table: Categories/ items Max. notes: (all values in seconds) Switch viewing to fluoroscopy 1.8 from pedal press till the first image shown Switch viewing to exposure 1.8 from pedal press till the first image shown, For non-DSA cine Switch viewing to exposure 2.5 from pedal press till the first image shown, For acquisition types with test-shot lock-in and frame-speeds greater or equal to 4 fr/s Switch fluoro to viewing 1.8 from pedal release till display of viewing image on screen Switch exposure to viewing 1.8 from pedal release till display of viewing image on screen Switch between exposure/fluoro 1.8 For acquisition types without test-shot lock-in and frame-speeds greater or equal to 4 fr/s Switch between exposure/fluoro 2.5 For acquisition types with test-shot lock-in and frame-speeds greater or equal to 4 fr/s Digital Subtraction Angiography (DSA): delay after test-image 1.5 Change live fluoro settings 1.5 Fluo flavour, field-size (incl FD-mode). Change live exposure settings 0.5 Frame-rates for multiphase acquisitions Delay between Xray-OFF and Indicators-OFF 1.0 Fluoro pipeline perception latency 0.200 Essential for eye-hand coordination. In this mode, the images shall be displayed subtracted with the DeviceMask resulting into the so-called Device image, and the Device image shall be displayed superimposed by the (faded) VesselMask. Optionally (EPX determined) the Device image shall be displayed superimposed by the LandmarkMask (containing the bones structure, derived from the VesselMask). Minimal movement in acquisition is assumed, to obtain enhanced view with single mask subtraction. Enables that the user can step through the run to see the progressive filling of arteries or veins. The user can switch subtraction on/off, and/or select another mask from the run. Manually, registration errors due to patient movements can be corrected ("Pixel-shift" function). A background portion can be added manually for orientation purposes ("Landmarking" function). Subtracted imaging requires a license. In this mode, the images shall be displayed subtracted with the DeviceMask resulting into the so-called Device image. Optionally (EPX determined) the Device image shall be displayed superimposed by the LandmarkMask (containing the bones structure, derived from the DeviceMask). During acquisition (in the LiveSubtract and VesselSubtract modes), the so-called DeviceMask shall be calculated from a number of received images (EPX determined). The DeviceMask shall be retained for subsequent acquisitions, unless one of the following cases occurs: The user requests a remask of the DeviceMask. Roadmap is deactivated. The user selects another SmartMask. The user selects a fluoro flavour. Displays the run subtracted after the mask is acquired. The mask is automatically selected (EPX procedure defined). The user can override subtraction on/off for the complete exposure run (before acquisition has started). Optionally automatic pixel-shift is applied to the live image (EPX controlled) to compensate for undesired movements. Subtracted imaging requires a license. During fluoro in the Device phase (after X-ray stabilization), subtraction can be toggled on/off without losing the related masks. Exposure subtraction can be enhanced by averaging several images to obtain an improved mask. This feature provides improved image quality with reduced noise. Settings are EPX determined and cannot be changed by the user. Subtracted view is obtained for an acquisition run exhibiting intended movement (e.g. DRA). This requires a license. Subtraction is based on an acquired mask run; the matching run is automatically selected. View-Trace An image sequence is composed into a single image, giving an overview of the filled vessel-tree structure. The composite image can be viewed in subtracted mode.
UNS.SmartSuite.ImageTranfer | Transfer image data to multiple destinations in multiple formats; Review the status (e.g. scheduled, in-progress, completed), of image data transfer; Receive confirmation (e.g. successful, non-successful) of image data transfer; Close examination and receive notification of closure; Re-open and re-close a closed examination. New, N.A. | SRS.Allura.Func.DICOM.Export/ImportJobManagement ¦ SRS.Allura.Func.StoreToMedia ¦ SRS.Allura.Func.ReopenFinishedStudy ¦ SRS.Allura.Func.DICOM.Export ¦ SRS.Allura.Func.FinishStudy ¦ SRS.Allura.SLS.IHE | This function enables to observe a submitted import or export job. The user can abort, delete and retry the job. This function enables to store image data on media (CD/DVD/USB). The user selects DICOM format or PC format. With DICOM format, the user can select to store the DICOM viewer with the exported DICOM files. With PC format, the user determines the file name for creating a PNG (Portable Network Graphic) file or an MPEG4 (Moving Picture Experts Group) file. The content can be deidentified and the user can define a fictive patient name for the deidentified content. The user is able to re-open a finished study/procedure step and select it for acquisition. The study/procedure step is re-initialized in such way that a complete workflow scenario may again be applied on the study/procedure step. Actions are similar compared to selection from the scheduled list, except: When a finished study/procedure step is selected, new images acquired are appended to the study/procedure step data. The procedure card, clinical-file structure, dose report data and other relevant data already attached to this study/procedure step are used as a starting point for any additional X-ray acquisitions. Finishing A Study Enables to export (part-of) a study to a remote DICOM workstation or to a DICOM archive. Auto-export at the end of run and/or finish study, according EPX defined destination (multiple) and contents (e.g. images, downscaling). The user shall be able to get an overview of which runs will be auto-exported upon finishing the study and to which destination. For manual export, the user can select the content and the destination. For manual export, the content can be de-identified and the user can define a fictive patient name for the de-identified content. The selected destination determines the image processing format (raw or processed) which will be used. A biplane run is exported as two separate DICOM objects that refer to each other. With this function the user indicates that this study is finished on the system. Per procedure step within the study, EPX defined automation tasks are started (e.g. export). After pending jobs (e.g. archiving) are successfully finished, the procedure step is finished and (if MPPS is configured) reported according DICOM MPPS. Unless set to protected, the occupied storage space is released for automatic deletion. Examination Related Presetting EPX Concept The system implements all of the transactions specified in the Integrating the Healthcare Enterprise (IHE) Technical Framework to support the IHE Integration Profiles, Actors and Options listed below: Table: Integration Profiles Implemented Actors Implemented Options Implemented Cardiac Catheterization Workflow (CATH) Acquisition Modality Broad Worklist Query Scheduled Workflow (SWF) Acquisition Modality Patient Based Worklist Query, Broad Worklist Query Patient Information Reconciliation (PIR) Acquisition Modality none Consistent Presentation of Images (CPI) Acquisition Modality none Radiation Exposure Monitoring (REM) Acquisition Modality none Consistent Time (CT) Time Client none Audit Trail and Node Authentication (ATNA) Secure Node none
UNS.SmartSuite.ImagingControl | Control the imaging settings at table side; Select/adjust/control the detector field size, FOV, position of shutters/wedges, Source Image Distance (SID), X-ray channel, fluoro-buzzer, fluoro-grab, fluoro-flavour, Roadmap Pro on/off. New, N.A. | SRS.Allura.Func.StoreFluoro ¦ SRS.Allura.Func.RotateFDXDDetector ¦ SRS.Allura.Func.PositioningTableTop ¦ SRS.Allura.Func.AngulateBeam ¦ SRS.Allura.Func.APC3DReference ¦ SRS.Allura.Func.TableCradle ¦ SRS.Allura.Func.Move/ResetWedges ¦ SRS.Allura.Func.SwingStandBase ¦ SRS.Allura.Func.EmergencyStop ¦ SRS.Allura.Func.ResetGeometry ¦ SRS.Allura.Conf.UI-RelatedItems ¦ SRS.Allura.Func.BiplaneAcquisition ¦ SRS.Allura.Func.APCPredefined ¦ SRS.Allura.Func.RotateBeam ¦ SRS.Allura.Func.Z-RotateStandBase(L-ArmRotation) ¦ SRS.Allura.Func.Move/ResetShutters ¦ SRS.Allura.Func.FluoroFlavours ¦ SRS.Allura.Func.MoveBeam ¦ SRS.Allura.Func.APCReference ¦ SRS.Allura.Func.TableTilting ¦ SRS.Allura.Func.APCTable ¦ SRS.Allura.Func.SelectTableSide ¦ SRS.Allura.Func.SelectDetectorFieldsize ¦ SRS.Allura.Func.TableLock ¦ SRS.Allura.Func.MoveBeamXY ¦ SRS.Allura.Func.Roadmap.SmartMask ¦ SRS.Allura.Func.ChangeSIDByDetector ¦ SRS.Allura.Func.ChangingTableHeight ¦ SRS.Allura.Func.APCStore/Recall ¦ SRS.Allura.Qual.ClinicalResponseTimes.Movement ¦ SRS.Allura.Func.FluoroTimeBuzzer ¦ SRS.Allura.Func.Roadmap | Fluoro images are cached with cyclic overwrite for an EPX defined time-span. During fluoroscopy, the user can grab (i.e. store) single fluoro images or an image sequence. Multiple grabs during a single fluoro run are stored for each grab in its own run. After fluoroscopy, auto-cycle displays the cached sequence. After fluoroscopy, the user can store this complete cached run as a normal run. The cache is cleared at the start of a next fluoro run. Enables the user to rotate the frontal FD20 detector for optimized imaging (manual and motorized). User shall be able to rotate the detector by 90° in order to switch between landscape and portrait usage of the detector. Beam width is adapted to maintain the selected detector field size. User shall be able to rotate the detector by any angle between 0° and 90° in order to align with specific anatomy. In FlexArm (XY-fa) and biplane (Clea3 & L-arc3 only) configurations, the collimator rotates synchronously to maintain the selected detector field-size. In all other configurations, the collimator is fixed and the beam-width is limited such that the irradiated area is completely visible. This applies to motorized or manual positioning of the tabletop in longitudinal and/or lateral direction. Table-type dependent manual positioning (so called 'float') is balanced (passive or motor-assist) to prevent the patient from sliding down (e.g. in case of tilt). Changes the angle between beam and longitudinal-table-axis (motorized). Angulation can be performed simultaneously with beam rotation. During motorized movement the angle position is displayed. The user shall be able to recall the frontal projection (rotation/angulation) that corresponds to a view on the Interventional Workspot (3DRA, XperCT, etc). Indicated by the Interventional Workspot, the SID shall go to minimum or to maximum (maximum to e.g. prevent hitting the needle) during the recall movement. APC 3D reference requires a license. The table rotates on the longitudinal axis (motorized). The cradle rotation center is ISO-centric within the limits set by the particular table type (Maquet/TruSystem OR Tables are excluded) The cradle movement automatically stops in horizontal position. Enables wedges to move in/out, or rotate. Reset puts the wedges outside the imaging area. The default position of wedges is EPX application/procedure determined. Enables the user to rotate the Poly-G floor stand around a vertical axis far outside the iso-center. Enables the user to image objects (e.g. patient's arm) that are not centered on the table. Enables the user to rotate the stand away from the table for parking purposes. Automatic Position Control (APC) Focus is on applying pre-defined or user-defined positions in controlling table and/or stand movements. The user shall be able to stop all motorized movements. Within operators reach, red-colored emergency stop buttons are provided. Movements to free the patient from the system remain possible. Normal operation is resumed after a system restart. The user shall be able to put the stand and table in a predefined position, in order to free the patient, or to provide a maximum of accessibility to the patient. Note: In case of TruSystem table or Maquet Magnus table with universal table top, only the stand is put in a predefined position. Table: sub tag item monoplane biplane .xs X-ray switches -1 footswitch 3f + -2 footswitch 6f + -3 wireless footswitch 3f o -4 wireless footswitch 6f o -5 additional footswitch CR 3f o -6 additional footswitch CR 6f o -7 additional footswitch ER 3f o -8 additional footswitch ER 6f o -9 X-ray handswitch + + Table: sub tag item Monoplane Monoplane Monoplane XY (fm/fa) XY (fm/fa) XY (fm/fa) biplane biplane biplane sub tag item AD7 XT AD7 XNT Maq AD7 XT AD7 XNT Maq Tru AD7 XT AD7 XNT Maq ui UI-modules -1 ControlModule NT +1) -2 ControlModule Tilt +1) -3 ControlModule Flex +1) +1) -4 ControlModule ORT o2) o2) +1,3) -5 ControlModule ORT Flex o2) o2) +1,3) -6 ControlModule Geo NT7) +1) -7 ControlModule Geo Tilt7) +1) -8 ControlModule Geo ORT7) o2) o2) + -9 ControlModule Geo NT IBR +1) -10 ControlModule Geo Tilt IBR +1) -11 ControlModule Imaging o4) o4) o4) o4) o4) o4) + + + -12 speed controller o5) o5) o5) o5) o5) o5) -13 pan handle o o o o o o -14 TouchScreenModule +6) +6) +6) +6) +6) +6) +6) +6) +6) -15 review module + + + + + + + + + -16 viewpad (cardio) # # # # # # # # # -17 viewpad (mixed) # # # # # # # # # -18 ControlModule Geo ORT IBR o2) o2) + Notes: 1) Additional ControlModule optional for CR 2) Optional as additional ControlModule on pedestal for ER 3) On pedestal 4) ControlModule imaging optionally available for ER (not in combination with 2)) 5) Required for FDPA, not an option for TruSystem 6) Additional TSM optional for CR and (on pedestal) for ER 7) Not for NPI, applicable for modules already present in the installed base Table: sub item Table type Table type Table type notes tag AD 7XT AD 7XNT Maquet TruSystem .mi miscellaneous -1 TSM swingarm + + + -2 Pedestal o o + trolley to attach UI modules; for AD7X(N)T including Pedestal Wall Conn. Box; for ORT including Surgery Wall Conn. Box -3 Two-way Intercom o o o -4 Keyboard + Mouse in CR + + + in CR; additional second keyboard + mouse in case of additional FlexSpot -5 Mouse in ER o o o Requires a license -7 MultiVision Switch o o o see SRS.Allura.Intgr.Multivision -8 KVM-Switch o o o see SRS.Allura.Intgr.KVM-Switching, requires DVI based video infrastructure -10 MCS Auxiliary box + + + for viewpad-receiver and X-on light on 3rd party booms In a biplane system, the user can acquire a frontal and lateral projection simultaneously, saving procedure time and contrast medium. This is manually pre-selected before the start of a run (for exposure), or based on the activated pedal(s) (for fluoro). The system shall provide pre-defined sets each comprising a sequence of projections (defined by rotation/angulation, SID and Field of View). The sets shall be EPX defined and a set can have max. 20 projections. The user shall be able to customize the sets. Specific sets can be customized for different users. The user shall be able to recall the projections of a set in sequential order, but alternatively the user shall be able to recall a projection directly. APC predefined requires a license. The user shall be able to rotate the beam around the longitudinal-table-axis (motorized). Intuitive control and display orientation rules (see SRS.Allura.Func.DisplayOrientation) also apply if rotation causes a reversed beam direction. In biplane systems, the frontal and lateral beam can be rotated simultaneously. During motorized movement the rotation position shall be displayed. Enables the user to rotate the frontal stand for position at the side or head of the table for optimized patient imaging. Iso-center stand rotation around a vertical axis to accomplish improved access to the patient. Function does not apply to the Poly-G floor stand. Enables shutters to move in/out. Opposite shutters move in pairs; horizontal and vertical pairs can be moved simultaneously. Opening stops at the edge of the current detector field (so becomes again visible at a larger field). Closing must be immediately visible (i.e. fast moving to edge of detector field, if initially larger). In case a user-defined position is overruled (by switching towards a smaller field-size), enlarging the field-size will automatically widen the shutters until the original user-defined position is reached. Reset puts the shutters at the edge of the detector field-size. The default position of shutters is EPX application/procedure determined. The user can select between 3 EPX defined flavours, determining e.g. fluoro-frame-rate and X-ray dose. The flavour-determined frame-rate can be 30, 25, 20, 15, 12.5, 10, 7.5, 6.25, 5, 3.75, 3.125, 2.5, 1.875, 1.25, 1.0, 0.625, 0.5 img/sec. In biplane, frontal frame-rate equals lateral frame-rate (available range same as for monoplane). A fourth flavour can be selected by the system in situations where low-load is necessary (e.g. for graceful degradation). For legal reasons, a low-dose fluoro flavour is available with a ratio normal/low of at least a factor 2. Configurable, a high-dose flavour with a maximum of 20 R/min is offered. A flavour can also be selected during fluoro. The user shall be able to move the frontal and/or lateral beam longitudinal to select a clinical area for acquisition (manual and motorized). This function only applies to ceiling suspended configurations. The user shall be able to recall the geometry positions (i.e. position of frontal and lateral stand, table, shutters and wedges) plus the receptor field size which corresponds to an image that is displayed in a live or reference viewport. In a biplane system, the user shall be able to recall a position for both frontal and lateral channels independently or simultaneously. Note: Table position recall is not supported for TruSystem tables in combination with FlexMove. On AD7X tables recall of longitudinal and lateral table position requires appropriate motors which are optional APC reference requires a license. This enables the user to tilt the tabletop to lift or lower the patient's head relative to his feet (motorized). The tilt rotation center is isocentric (except for TruSystem table) within the limits set by the particular table type. Customizable, the stand can follow the tilt movement in a synchronous way (except for TruSystem table). The tilting movement automatically stops in horizontal position. Isocentric tilt on the Magnus table requires a license. After the user has brought a point-of-interest into the system’s iso-center, the user shall be able to store the corresponding table position (i.e. table longitudinal, lateral, height, tilt and cradle). The system shall provide a default point-of-interest (EPX defined). When a new acquisition study is selected the 'stored table position' is reset such that the default point-of-interest is in the iso-center. The user shall be able to recall the stored table position or only the table height. The system shall indicate to the user if the table is in the stored position. Note: Not applicable for Hillstrom TruSystem table with universal table top and Maquet Magnus table with universal table top. On TruSystem tables with the SQ14-Xtra table top only height can be recalled. On AD7X tables recall of longitudinal and lateral table position requires appropriate motors, which are optional. APC table requires a license. The user shall be able in table-oriented mode to maintain intuitive movement control from four table sides. The user shall be able to select the table side on every movement-control UI module individually. Combined Patient- And Beam Positioning This section focuses on movements under direct user control, which are not stand or table specific. The detector field-size is the detector area that is being acquired and displayed. The user shall be able to select a field-size; for fluoro also during acquisition. The available field-sizes and related image quality characteristics depend on the type of detector. The default shall be restored when a new procedure step is selected for acquisition, and EPX-dependent at EPX procedure selection. The user shall be able to lock/unlock: The table pivot movements. The table lateral movements. All table movements. For the TruSystem table in combination with FlexMove, the lateral table lock/unlock is not supported via the system UI. Beam Positioning This section focuses on moving and positioning the stand under direct user control. The user shall be able to move the beam longitudinal and/or transversal (manual and motorized). To select a clinical area for acquisition as an alternative to positioning the tabletop. For standby parking (typically transversal only) during a surgical procedure. For normal parking when the stand is not used. Alternative to the roadmap Vessel phase, the user can directly select an image from any stored run and use it as the VesselMask (‘SmartMask’) for the Device phase. Before selecting, the smart mask can be enhanced using pixel-shift. Selecting a ‘SmartMask’ requires a license. Changes the Source Image Distance (SID) by moving the detector along the beam (manual and motorized). The beam width is adapted automatically to maintain the selected detector field-size. Moves the table up/down (motorized). The movement range depends on the table type and on the table tilt/cradle angle. The system shall allow the user to store and recall multiple geometry positions (i.e. position of stand, table, shutters and wedges) including the receptor field size. When a new acquisition study is selected, all stored positions shall be removed. The system shall indicate to the user if the system has not reached a selected position. Note: Table position recall is not supported for TruSystem tables in combination with FlexMove. On AD7X tables recall of longitudinal and lateral table position requires appropriate motors which are optional Table: Categories/ items Max. notes: (all values in seconds) Move shutters & wedges 0.5 from UI up to & including image display. APC recall (start geometry movement) 1.0 1.5 for Maquet & TruSystem: 2.5 if the table is in hibernation mode Start all other movements 1.0 1.5 for Maquet & TruSystem: 2.5 if the table is in hibernation mode Warns the user after an elapsed accumulated fluoro time of 5 minutes: This safety mechanism can be customized on/off, dependent on the local needs and safety regulations. The user can deactivate the buzzer. The user shall be able to activate or deactivate roadmap. If roadmap gets activated, the last used roadmap mode since EPX application selection is used, otherwise the default roadmap mode for this EPX application is used. Dependent on the roadmap mode, the system starts roadmap in either the Vessel phase or the Device phase. The active roadmap phase is indicated to the user. Selection of a SmartMask is an implicit activation of roadmap. The Device phase is automatically entered. Subtracted imaging requires a license.
UNS.SmartSuite.InjectorCoupling | Couple certain automated injector systems; Perform timed administration (injection) of contrast medium (precise timing depends on patient’s circulation, injection protocol, and imaging protocol); Set X-ray delay; Perform manual injection; Perform automatic injection; Couple/decouple automatic injection to/from acquisition at EPX level; Maintain coupling after acquisition; Perform acquisition with manual injection while coupling is on; Adjust contrast injector parameters at table side. New, N.A. | SRS.Allura.Intgr.Injectors ¦ SRS.Allura.Conf.RemainingComponents ¦ SRS.Allura.Func.ContrastInjection ¦ SRS.Allura.Func.SetInjectorCoupling | The system shall provide interfaces to mount, power and control an injector. Departmental Image Storage Providers and Image Storage Users Table: sub tag item availability notes .ps Partner Systems See for details corresponding requirement in Room-Specific Auxiliary Systems section -1 DoseAware Xtend o .mi Miscellaneous -1 Ceiling Suspended radiation shield o -2 Examination light o -3 X-ray spacer o limits minimum Source Skin Distance (Health and Human Services (HHS) requirement). Mandatory for 21 CFR 1020.30. -4 Drip stand o -5 Medical DVD recorder o Requires DVI based video infrastructure -6 1 phase UPS o see SRS.Allura.Func.PowerFailureBehaviour -9 UPS control c see SRS.Allura.Intgr.ExternalThreePhaseUPS -10 VideoSlave o see SRS.Allura.Intgr.VideoSlave -11 Open video interface o see SRS.Allura.Intgr.OpenVideoInterface.LargeScreen, -12 DVD burner o see SRS.Allura.Func.StoreToMedia -13 HD output o see SRS.Allura.Intgr.OpenVideoInterface.LargeScreenDownscaled; in DVI based video infrastructure it requires FlexVision open video interface -14 Video scan converter o for 50/60 Hz SLR TV output, requires DVI based video infrastructure -15 Wall-connection boxes o see SRS.Allura.Intgr.FHDVideoSources, requires DVI based video infrastructure -16 Auxiliary System Input Connection Box o see SRS.Allura.Intgr.FHDVideoSources, requires IP based video infrastructure -17 4K Auxiliary System Input Connection Box o see SRS.Allura.Intgr.4KVideoInputOutput, not available in combination with DVI based video infrastructure Core Functionality Core functionality is defined as functionality which is not dependent on the presence of any license key. The following functions are in the core: Clinical core functionality is a subset of all functionality as defined in chapter FUNCTIONAL REQUIREMENTS. The subset is determined by subtracting all functionality that is dependent on any license key. All functionality needed to install a system and to perform basic checks to initially verify correct system working condition. User messages (see SRS.Allura.Func.DisplayUserMessages). Tooltip (balloon) help. Software-Keyed Clinical Options A Clinical Option gives additional clinical functions/features/characteristics and is enabled through a software license key. Two types of licenses exist: counted license: used to implement licenses that restrict the number of items that can be used, e.g. number of video outputs. uncounted license: used to implement licenses that restrict the usage of a certain feature (on/off) For fluoro, contrast medium is to be administered manually, and no synchronization or coupling to a motorized injector is provided. Enables to couple motorized injection to acquisition. This requires a license. The user can select between coupled and uncoupled. At a new procedure step selection for acquisition, and customizable after every run, the injector is set to uncoupled. Timing of injection is based on a fixed "injector response time" and EPX defined image after which injection should start and “Contrast Arrival Time” (user overridable). In case the start of acquisition is delayed, a “Count-down bar” is shown representing remaining time. Type of injector coupling is customizable: One-knob: X-ray hand- or footswitch control both injection and X-ray. Two-knob: X-ray hand- or footswitch controls X-ray, but injection is enabled by injector handswitch. Acquisition Of Physiology Signals
UNS.SmartSuite.InspectCleanLubricate | Inspect, clean, lubricate, and adjust accordingly to the critical to safety and quality aspects of the system in the image-guided lab. | SRS.Allura.Serv.Maintenance.Planned | The system shall enable the performance of planned maintenance activities to ensure its functionality during its lifetime.
UNS.SmartSuite.IntegratedChecklistandProtocols | Make a scan of their checklist(s) (time-out document, compromising text and images) and display it on the screen; Add Checklists and Protocols (comprising text and images of e.g. injector protocols, kind of disposables etc.) to a Procedure Card; Recall Checklists and Protocols from a Procedure Card before or during the procedure. New, N.A. | SRS.Allura.Func.ProcedureCard ¦ SRS.Allura.UI.Help ¦ SRS.Allura.UI.Help.Customize | The user shall be offered a procedure based workflow by having procedure specific settings for a procedure step. A procedure card determines the procedure specific settings, being: The EPX selections the user can choose from during the procedure and a default EPX. The default preset for FlexVision. The default preset for FlexSpot. The default preset for Second FlexSpot. The patient orientation and surgical view Procedure guidance. Specific settings for compatible room specific auxiliary systems Predefined procedure cards shall be delivered with the system. The user can customize procedure cards and organize/group them. The user can set which procedure card shall be used when a procedure step is selected for acquisition. This can also be set automatically based on information from the WLM system (RIS code). The user shall be able to access integrated help and electronic Instructions for Use. This includes tooltips, task and procedure guidance and instruction material. The user shall be able to search for a specific topic within the Instructions for Use. The user shall be able to delete and modify a help header and link a new help header to an imported hospital protocol/checklist. This requires a license.
UNS.SmartSuite.IntegratedDoseAware | Automatically receive real-time feedback of radiation exposure to user and staff. New, N.A. | SRS.Allura.Intgr.DoseAwareXtend | The system shall optionally provide integration with the DoseAware Family product, which means: In combination with DoseAware Xtend, video can be shown on the Azurion display. Azurion provides X-ray procedure related information to DoseAware Xtend. A reference dosimeter can be attached to the frontal stand. In combination with DoseAware the system shall provide electrical and mechanical integration of the DoseAware basestation on the MCS frame.
UNS.SmartSuite.IntegratedHelp | Electronically search through the user manual for help on specific topics/tasks/problems (e.g. digital help function) integrated into the system. New, N.A. | SRS.Allura.UI.Help | The user shall be able to access integrated help and electronic Instructions for Use. This includes tooltips, task and procedure guidance and instruction material. The user shall be able to search for a specific topic within the Instructions for Use.
UNS.SmartSuite.IntegratedIFU | Have access to the latest/up-to-date Instructions for Use electronically on the system. New, N.A. | SRS.Allura.Serv.RemoteSWInstall ¦ SRS.Allura.SLS.ListOfApplicableStandards ¦ SRS.Allura.UI.Help | It shall be possible to distribute software packages over the RSN and remotely install these on the system. The local user is requested to start installation, perform necessary tests, and validate/accept. Without associated risks, user interaction is not needed. Remote Desktop Remote desktop enables a remote service center to assist a local user. The standards as mentioned in chapter APPLICABLE STANDARDS are applicable. Environmental Care The user shall be able to access integrated help and electronic Instructions for Use. This includes tooltips, task and procedure guidance and instruction material. The user shall be able to search for a specific topic within the Instructions for Use.
UNS.SmartSuite.IntegratedTraining | Access training materials on/onto the system (e.g. help materials, integrated applications/ simulations). New, N.A. | SRS.Allura.UI.Help | The user shall be able to access integrated help and electronic Instructions for Use. This includes tooltips, task and procedure guidance and instruction material. The user shall be able to search for a specific topic within the Instructions for Use.
UNS.SmartSuite.Intercom | Communicate verbally between ER and CR (e.g. by using an Intercom). New, N.A. | SRS.Allura.Func.Intercom | A two way intercom enables to communicate between control and exam room.
UNS.SmartSuite.IntuitiveCollisionDetectionFeedback | Take appropriate measures in order to avoid and/or resolve collision based on the information and guidance received; Receive visual feedback (e.g. icons) showing what is causing a collision and where. New, N.A. | SRS.Allura.Func.ExceptionHandling | A number of exceptional situations can be distinguished: A collision shall be indicated, also if other user messages are active. After a software error recovery, the system shall guarantee that previously stored clinical information is still available. If the user requests a function that cannot be executed in the current system state, for a reason which is not obvious to the user, the system shall display an appropriate user message. If the user requests a function and the system cannot execute the function because the system has run out of resources (e.g. storage full), the system shall display an appropriate user message. Graceful Degradation
UNS.SmartSuite.LIHTablePanning | Reposition the patient without using fluoroscopy (Zero Dose Positioning) including an indication of the new position on both the lateral as the frontal display, e.g. by using a graphical image area overlay of the resulting image area when the user moves the table, shutter or wedge or changes the detector field size. New, N.A. | SRS.Allura.Func.PositioningIndicationWithoutRadiation | This enables positioning before a run is started, in order to save patient radiation dose. During (manual or automatic) re-positioning of stand and table the anticipated radiation area shall be displayed (EPX controlled) as an overlay on an image for an EPX defined time. Note: Not supported for Hillrom TruSystem table or Maquet Magnus table with universal table top Position indication without radiation requires a license. Patient Positioning This section focuses on moving and positioning the table under direct user control.
UNS.SmartSuite.Languages | Read from a user interface in his/her native or at least in familiar language; Select one user interface language from a range of user interface languages. New, N.A. | SRS.Allura.UI.Clinical-UILanguage | The system can be configured according the following local languages: Table: English French German Spanish Italian Dutch Danish Norwegian Swedish Simplified Chinese* Traditional Chinese* Japanese* *Note: No manual (i.e. keyboard) input for Chinese and Japanese language Support for Japanese and Chinese requires a license.
UNS.SmartSuite.LastImageHoldOrAutoCycle | View/review after fluoroscopy and acquisition the last acquired images; Select that the last acquisition run is displayed in a cycle. New, N.A. | SRS.Allura.Func.Cycle ¦ SRS.Allura.Func.StoreFluoro ¦ SRS.Allura.Func.LastImageHoldOrAutoCycle | This function enables to display a sequence of images in cine mode. In file-cycle all exposure and fluoro runs in the active file of the current study are cycled sequentially. In run-cycle the complete active run is cycled. User can set speed. Intra-Image Post-Processing Focuses on intra-image processing functions for customized enhanced display. General Concepts Fluoro images are cached with cyclic overwrite for an EPX defined time-span. During fluoroscopy, the user can grab (i.e. store) single fluoro images or an image sequence. Multiple grabs during a single fluoro run are stored for each grab in its own run. After fluoroscopy, auto-cycle displays the cached sequence. After fluoroscopy, the user can store this complete cached run as a normal run. The cache is cleared at the start of a next fluoro run. Customizable, after acquisition the last exposed image is displayed or the run is cycled. Auto-cycle stops automatically after a customizable time-out (static image remains on display).
UNS.SmartSuite.LightsAndOperationLamp | Switch on/off or dim the room-light and/or operation lamp, optionally synchronized to acquisition. | SRS.Allura.Func.RoomLightAndOperationLamp | Configurable at installation, the room-light and/or operation lamp can be switched or dimmed by a dedicated foot pedal, or synchronized to acquisition.
UNS.SmartSuite.LinkWorkstation | Have a direct (real time link) connection with the Interventional WorkSpot and Coronary Tools and EchoNavigator and access to relevant Interventional Tools (e.g. StentBoost); Have adequate image quality for (3D) reconstructions/information at their Philips workstation(s). New, N.A. | SRS.Allura.Intgr.Interfaces.RealTimeImageLink ¦ SRS.Allura.Func.XperCT ¦ SRS.Allura.Func.DRAFor3D ¦ SRS.Allura.Func.ClassicDRA ¦ SRS.Allura.Intgr.Interfaces.ControlFromAuxiliarySystems ¦ SRS.Allura.Func.DualPhaseXperCT ¦ SRS.Allura.Func.Dual-AxisAcquisition ¦ SRS.Allura.Intgr.Interfaces.ControlAuxiliarySystems | The system shall provide communication interfaces to provide x-ray images of the frontal channel in real time to compatible auxiliary equipment during x-ray acquisition. This requires a license. This is a special case of DRA or DAR for 3D enabling reconstruction of soft-tissue. This function is only supported with an FD20 detector. The system shall offer dedicated EPX procedures and related settings for Neuro (DRA and DAR) and Abdomen (DRA only) applications. The system shall perform acquisition with up to 60 images/second to obtain acceptable scan time. The system shall be able to apply slit collimation to reduce patient dose. DRA for 3D is a special case of DRA performed at predefined positions, to enable 3D reconstruction. Settings, including start- and end-positions, shall be defined in EPX. The system shall support DRA for 3D at least at the following positions: for FlexArm (XY-fa) configuration at head, doctor and nurse side for other FD20 and FD15 configurations at head and one side position (doctor or nurse) for FD12 only at head position SID is maximized to allow for 3D scans of obese patients. The user shall define the end- and start-position of the DRA scan, implicitly executing a test-scan to detect obstructions. The system shall be able to perform the DRA scan at the side or head of the table (C-arm roll and propellor respectively). DRA acquisition is only possible for the frontal channel. The user shall be able to acquire a second run to enable subtracted reviewing. Settings, including start- and end-positions, shall be defined in EPX, but if allowed in EPX, the user shall be able to choose another end- and/or start-position. DRA requires a license. The system shall provide communication interfaces to allow compatible auxiliary systems to select EPX procedures which may or may not be specific for the auxiliary system (i.e. they are selectable only by the auxiliary system) set APC target positions lock/unlock the table set specific acquisition parameters (e.g. injector coupling, patient orientation, biopsy mode, etc.) push a secondary capture image to a reference viewport This is a special case of XperCT where two consecutive DRA scans are being made by executing the scans in different phases of contrast propagation, which is known as Dual phase XperCT. The two scans are performed in opposite direction to reduce the overall procedure time to minimize the amount of contrast medium. For neuro applications, to allow the use of intravenous contrast injections in Dual phase XPerCT scans, the system shall provide a subtracted fluoroscopy flavor with enhanced contrast. It shall enable to detect the arrival of contrast agent in order to start the scan. This application shall be executable from the Control Room. Flexible Dynamic Peripheral Angiography (FDPA) A run is acquired while the X-ray beam is rotated and angulated according a curved movement profile. This function is only available for the frontal channel. Scans are performed automatically according pre-defined acquisition/geometry settings; cardiac swing or helical. Cardiac Swing requires a license. Digital File & Data Management Focus is on the handling of files with images and/or related data, as perceived by the user. File & Data Operations The system shall provide communication interfaces to control compatible auxiliary systems from the systems user interface with wireless mouse on the auxiliary system’s graphical UI displayed on FlexVision with mouse and keyboard on the auxiliary system’s graphical UI displayed on Flexspot from TSM by a generic button based UI tailored by the auxiliary system or by a graphical UI streamed from the auxiliary system including (multi-)touch control. This requires a license control shall be established automatically when (mobile) compatible auxiliary system is connected
UNS.SmartSuite.LiveImageView | Display real-time images from fluoroscopy as well as exposure; View live and previously acquired images simultaneously; Record the image stream on a DVD during fluoroscopy/exposure; For monoplane: Save images for reference up to three different reference screens; For biplane: Save images for reference up to two different reference screens per channel. In monoplane mode save images up to 3 reference screens for the active channel. New, Have the ability to view multiple previously acquired images (i.e. reference screens) simultaneously with live; Show an external source in its native resolution on a 3rd party monitor; Show multiple sources in their native resolution; Use 4K/UHD input connectivity to display 3rd party sources; Deliver artefact-free images distributed with near-zero latency via optical fiber. | SRS.Allura.Intgr.4KVideoInputOutput ¦ SRS.Allura.Qual.ClinicalResponseTimes.Acquisition ¦ SRS.Allura.Intgr.FHDVideoSources ¦ SRS.Allura.Func.StoreReferenceRun/Image ¦ SRS.Allura.Func.DVDRecordControl ¦ SRS.Allura.Intgr.NativeIQ ¦ SRS.Allura.Func.ActivateExposure ¦ SRS.Allura.Func.SingleShotFootpedal ¦ SRS.Allura.Func.ActivateFluoro | It shall be possible to connect to the system (connection available on MCS or wall) QHD, UHD or 4K video sources (max 2) and output it: as downscaled FHD video source to the system monitor or external video system optionally directly (no switching possibility for a clinical user) to a 4K 3rd party monitor(s) (mounted on a 3rd party monitor boom or installed in the control room) Connecting one 4K video source and displaying its output on the 4K 3rd party monitor shall not impact the total number of FHD video sources of the auxiliary systems that can be connected to the system by more than 10. Connecting and outputting 4K external inputs are both licensed. These are counted licenses for the number of 4K auxiliary system input connection boxes and 4K outputs installed. The licenses are not enforced. Table: Categories/ items Max. notes: (all values in seconds) Switch viewing to fluoroscopy 1.8 from pedal press till the first image shown Switch viewing to exposure 1.8 from pedal press till the first image shown, For non-DSA cine Switch viewing to exposure 2.5 from pedal press till the first image shown, For acquisition types with test-shot lock-in and frame-speeds greater or equal to 4 fr/s Switch fluoro to viewing 1.8 from pedal release till display of viewing image on screen Switch exposure to viewing 1.8 from pedal release till display of viewing image on screen Switch between exposure/fluoro 1.8 For acquisition types without test-shot lock-in and frame-speeds greater or equal to 4 fr/s Switch between exposure/fluoro 2.5 For acquisition types with test-shot lock-in and frame-speeds greater or equal to 4 fr/s Digital Subtraction Angiography (DSA): delay after test-image 1.5 Change live fluoro settings 1.5 Fluo flavour, field-size (incl FD-mode). Change live exposure settings 0.5 Frame-rates for multiphase acquisitions Delay between Xray-OFF and Indicators-OFF 1.0 Fluoro pipeline perception latency 0.200 Essential for eye-hand coordination. In combination with IP based video integration infrastructure, FlexVision exam room display configuration and FlexSpot + 2nd FlexSpot, or MultiView + Additional MultiView (i.e. no other CR/ER displays), the system shall support connecting up to 20 FHD auxiliary systems and up to 4 integrated workstations. In combination with DVI based video infrastructure the system shall support up to 16 FHD video sources, of which up to 11 video sources (total number from the integrated workstations and auxiliary systems). The system shall offer the flexibility to connect auxiliary systems in the control room, the examination room (on MSCs, walls or equipment racks) and the technical room. For VoIP configurations, the support of FHD video sources is licensed. This is a counted license for the number of video sources. The license does not actually limit the number of FHD video sources that can be connected. The user shall be able to add the active image or run to a reference viewport. This option requires a license. Enables automatic start/stop of a medical DVD recorder, synchronous to generation of X-ray (fluoro and/or exposure). EPX controlled, the medical DVD recorder records a copy of the displayed images. This requires a license. When the video signal is not displayed on the system monitor the system shall be able to offer a transparent image path such that it can be used by a 3rd party system. Transparent image path between the system input and output is a point to point 3rd party video signal distribution that doesn’t make any change to the signal being distributed. Auxiliary Systems Data Exchange & Control This causes X-ray radiation, images are acquired, displayed and stored. The user is warned when the free space is insufficient to make the next exposure run. The run is stopped (in case of max. 1k2) after 1000 images, or when no more image storage free space is available. This enables the user to acquire a single image in vascular/lock-in mode directly via a separate footpedal. As long as the user activates fluoro, X-ray is generated and images are acquired and displayed. Exposure will temporarily interrupt fluoro; after the exposure switch is released, fluoro is automatically resumed (if still activated). After 10 minutes of uninterrupted fluoro, fluoro is automatically stopped.
UNS.SmartSuite.LocalStorage | Store/ retrieve fluoroscopy and exposure images with DICOM format on/from external destinations (e.g. PACS, USB, CD, DVD); Store fluoroscopy and exposure images with DICOM and non-DICOM format on local storage media/devices (e.g. USB, CD, DVD). New, N.A. | SRS.Allura.Conf.ImageStorageOptions ¦ SRS.Allura.Func.RetrieveFromMedia ¦ SRS.Allura.Func.StoreToMedia ¦ SRS.Allura.Intgr.Media ¦ SRS.Allura.Intgr.DVDRecorders | Table: sub tag 5122 1K2 (2K)2 .xs Capacity Keys 5 -1 Single + 100.000 25.000 6250 -2 Double6 o 200.000 50.000 12500 -3 Quadruple6 o 400.000 100.000 25000 .gc Frame-rate Keys 1 -1 Basic + 12 img/s 12 img/s 6 img/s -2 Cardiac6 o 30 img/s 3 30 img/s 3,4 6 img/s -8 Pediatric6 o 60 img/s 2 30 img/s 3,4 6 img/s Notes: These figures do not apply to fluoro or to scenarios and represent maximum frame-rates expressed in images/second (img/s) within limits set by detector and configuration. In biplane this is frame rate per channel, and in addition to 60 img/s also 50 img/s is supported. Cardiac and pediatric also supports 25 img/s instead of 30 img/s. With processing algorithm ClarityIQ (see SRS.Allura.Func.X-rayImageProcessing) only <= 25 img/s @ 1K2 is supported. The image sizes give an indication how many of these sized images can be stored, the system is not limited nor bound to these sizes as they vary depending on detector size and configuration. This option requires a license INTEGRATION & INTEROPERABILITY REQUIREMENTS This chapter focuses on integration with other systems/equipment, in and outside the exam room. UPS Integration This function enables to retrieve image data according DICOM from media and store it in the local database. SERVICE REQUIREMENTS This chapter focuses on requirements related to service activities (termed "Field-Service"). General Requirements Field-Service System Access This function enables to store image data on media (CD/DVD/USB). The user selects DICOM format or PC format. With DICOM format, the user can select to store the DICOM viewer with the exported DICOM files. With PC format, the user determines the file name for creating a PNG (Portable Network Graphic) file or an MPEG4 (Moving Picture Experts Group) file. The content can be deidentified and the user can define a fictive patient name for the deidentified content. In control room, support for storage on CD/DVD and USB according general PC format and DICOM. Also importing from CD/DVD and USB according DICOM is supported. UPGRADABILITY REQUIREMENTS Upgradability existing X-Ray systems This release will be available for initial deliveries (always in combination with IP based video distribution infrastructure) and as software only field upgrade for Azurion system releases < R3.0. Upgrade will not impact external interfaces of the upgraded system. FUNCTIONAL REQUIREMENTS In this chapter the functions are grouped according to their role in the normal workflow. Global Functional Behavior Workflow Flexibility Compatibility with a medical DVD recorder for real-time storage of acquisition images. DP is the video output standard (in combination with IP based video infrastructure) SLR Video (PAL, NTSC) is the video output standard (in combination with DVI based video infrastructure)
UNS.SmartSuite.ManoeuvreBeam | Manoeuvre the system out of current working area; Manoeuvre the system around other equipment; Manoeuvre the system to current working area, either manually, via controls or automated. New, N.A. | SRS.Allura.Func.MoveBeam ¦ SRS.Allura.Func.SelectTableSide ¦ SRS.Allura.Func.Z-RotateStandBase(L-ArmRotation) ¦ SRS.Allura.Func.SwingStandBase | The user shall be able to move the frontal and/or lateral beam longitudinal to select a clinical area for acquisition (manual and motorized). This function only applies to ceiling suspended configurations. The user shall be able in table-oriented mode to maintain intuitive movement control from four table sides. The user shall be able to select the table side on every movement-control UI module individually. Combined Patient- And Beam Positioning This section focuses on movements under direct user control, which are not stand or table specific. Enables the user to rotate the frontal stand for position at the side or head of the table for optimized patient imaging. Iso-center stand rotation around a vertical axis to accomplish improved access to the patient. Function does not apply to the Poly-G floor stand. Enables the user to rotate the Poly-G floor stand around a vertical axis far outside the iso-center. Enables the user to image objects (e.g. patient's arm) that are not centered on the table. Enables the user to rotate the stand away from the table for parking purposes. Automatic Position Control (APC) Focus is on applying pre-defined or user-defined positions in controlling table and/or stand movements.
UNS.SmartSuite.Marker | Put multiple marks on the live viewport, from TSM, ER and CR; Markers can be adjusted or erased; Markers stay visible during fluoroscopy and exposure; Markers adapt to changes of magnification (zoom, SID or detector format); Markers will stay when copy to ref; Markers will not be stored on image. New, N.A. | SRS.Allura.Func.Marker | The user shall be able to draw on TSM or live viewport graphical overlay lines in order to mark certain anatomy on the clinical image. This requires a license. The marker overlay shall become and stay visible on TSM and live viewport during x-ray acquisition until deleted by the user or until another patient is selected. Quantitative Analysis Functions General Concepts
UNS.SmartSuite.Measurements | Perform a measurement on an image (angle, ratio, length); Perform manual and automatic calibration; Perform a distance measurement on TSM Pro. New, N.A. | SRS.Allura.Func.Measurement | The user shall be able to add measurements on a certain position into the image on display or on the TSM. This requires a license. The following functions are supported: length of a line, angle between 2 lines, and ratio between the length of 2 lines. Measurements can be repositioned, modified, hidden, copied and deleted. Measurement functionality shall support auto calibration and manual calibration methods to relate distances on display, expressed in pixels, to real-object distances (see SRS.Allura.Func.QACalibration) Accuracy of a length measurement shall be within ± 5% in automatic calibration when the measured object is in the iso-center and if the length of the object to be measured covers at least 50 pixels on the monitor Accuracy of an angle measurement shall be within ± 2.0 degrees. Multi-Image Post-Processing This section focuses on obtaining additional diagnostic value out of combining multiple images. Image Subtraction
UNS.SmartSuite.MechanicalInstallation | Mechanically install the system using the appropriate instruction guide in the image-guided lab. | SRS.Allura.Serv.FSInstructions ¦ SRS.Allura.UI.Help ¦ SRS.Allura.Serv.FSTools | Instructions are available on how to unpack, install, de-install, decommission, and packing and disposal of the system and components. The user shall be able to access integrated help and electronic Instructions for Use. This includes tooltips, task and procedure guidance and instruction material. The user shall be able to search for a specific topic within the Instructions for Use. Tools not in the standard service toolkit are included in the system delivery or made available separately.
UNS.SmartSuite.MechnicalDeinstallation | Mechanically deinstall the system in the image-guided lab based on the appropriate instructions; Mechanically deinstall parts of the system for corrective maintenance & Upgrade/FCO. | SRS.Allura.Serv.SingleLabConnection ¦ SRS.Allura.Serv.FSInstructions ¦ SRS.Allura.UI.Help | The system shall provide one entry point to the lab to service compatible room specific auxiliary systems, i.e. logging retrieval from and remote desktop of other configured systems shall be supported. The connected auxiliary systems shall be accessible even if the X-ray system is switched off. The remote & local FSE shall be able to troubleshoot connectivity between the entry point and other connected systems; and between the entry point and RSN. Miscellaneous Instructions are available on how to unpack, install, de-install, decommission, and packing and disposal of the system and components. The user shall be able to access integrated help and electronic Instructions for Use. This includes tooltips, task and procedure guidance and instruction material. The user shall be able to search for a specific topic within the Instructions for Use.
UNS.SmartSuite.MonitorConfiguration | Adjust the system to his particular viewing needs; Use the system in hybrid labs with minimal interference with Laminar Airflow; Be able in interventional labs to position the viewing support at 3 sides of the table, doctor, nurse and foot end; Adjust the screens on viewing distance and viewing height; Choose a wide range of screen configurations to match the intended procedure mix of the lab. Screen configurations range from separate loose screens to screens combined on single booms to large screens or combinations of those. New, N.A. | SRS.Allura.Conf.MCS-RelatedItems ¦ SRS.Allura.UI.PositionMCSFrame | Table: sub item notes tag Rail susp 3rd party boom 3rd party boom XL Springarm Springarm XL Up to two rail suspended MCS and in total up to four MCS of different type and configuration can be combined in one system .frm MCS frame -1 Frame + + + + -2 VxH= 1x2 % % % 2 x 27” -3 VxH= 1x2 + 1 % 2 x 19" + 1 x 27" -4 VxH= 2x2 % % % 4 x 27” -5 VxH= 2x2 + 2 % % 6 x 27” -7 Large Screen % % % 58” XL -8 Large Screen + 2 % % % 58” XL + 2 x 27” .dsp MCS displays max. number is configuration dependent; In combination with IP based video infrastructure the total number of all monitors installed on MCSs/booms in ER and FHD monitors installed outside MCSs/booms in CR and ER never exceeds 20. The maximum doesn’t include FlexVision, Slave FlexVision, (Second) FlexSpot, Basic CR, MultiView and Additional MultiView monitors -1 27” diagnostic color % % o % o 27" color monitor, min. 400 cd/m2, resolution 1920x1080, aspect ratio 16:9 -2 32” diagnostic color1) o 32" color monitor, min. 400 cd/m2, resolution 1920x1080, aspect ratio 16:9 -4 19” multi-purpose o 19" color monitor (only for VxH= 2x1 + 1) min. 250 cd/m2, resolution 1280x1024, aspect ratio 5:4 -5 58” XL % % % 58” color monitor, min. 400 cd/m2, resolution 3840x2160, aspect ratio 16:9 -9 Dummy fill o o o .swm Switchable Monitors -1 27” or 32” diagnostic color1) o o o Up to 16 switchable monitors, can also be offered without MCS Note: 1) 32" monitors are not part of NPI. The user shall be able to position the MCS frame(s) to have a clear view on the monitors from the workspot.
UNS.SmartSuite.MoveToPark | In case the FlexArm stand is not needed inter procedural, have no hinder of the system and have it parked/placed outside the working area, for example for cleaning, during the night, or for during a procedures that do not require imaging Park the FlexArm in central rail configuration away against the wall at foot side or head side of the room New, N.A. | SRS.Allura.Func.ParkorStandbyBeam | Enables the user to move the frontal and/or lateral stand out of the way to a standby or park position to improve access to the patient if the beam is temporarily not used. The different ways to move the stand to a park / standby position depend on the particular configuration (examples: lateral movement for Flexarm standby, longitudinal movement in ceiling stands, swing stand base in a Poly-G floor stand).
UNS.SmartSuite.MoveToStandby | Move the FlexArm system to and from a Standby Position by table side controls. A standby position is an optimal* position of the system in the room, when imaging is temporary not needed during a procedure. * Optimal is defined as where the gantry is out of the way during the procedure New, N.A. | SRS.Allura.Func.ParkorStandbyBeam | Enables the user to move the frontal and/or lateral stand out of the way to a standby or park position to improve access to the patient if the beam is temporarily not used. The different ways to move the stand to a park / standby position depend on the particular configuration (examples: lateral movement for Flexarm standby, longitudinal movement in ceiling stands, swing stand base in a Poly-G floor stand).
UNS.SmartSuite.NewIdentityUserInterface | Adjust visual display settings depending on the type of procedure (only relevant features should be displayed); Have focus on the clinical task and especially on the images (e.g. the legibility of visual display through better icon size, color, contrast, backlight, etc.). New, N.A. | SRS.Allura.UI.Design.Application ¦ SRS.Allura.Func.PresettingFunctions ¦ SRS.Allura.UI.UIModule ¦ SRS.Allura.Func.ProcedureCard ¦ SRS.Allura.UI.Design.FlexibleViewing ¦ SRS.Allura.UI.KeyboardMouse | Each system application shall have the same user interface design for use in the control room; likewise each system application shall have the same user interface design for use in the exam room to make it easy for the user to transition from one application to the next. The control room acquisition and review monitors shall have the same user interface design to host the system applications. Presetting is determined by the information in the study/procedure step that is current for acquisition. The user can modify the information. The procedure card to be used can be modified. The performing physician can be selected from a pre-defined physician list. The patient type can be entered/modified manually. The EPX application/procedure group/procedure can be selected from a list (which contains only items that can be performed within the system's current configuration). EPX application, procedure group and procedure can be selected in control room, as well as in exam room. EPX application, procedure group and procedure can be selected during x-ray. The selected settings will then be applied for the next exposure or fluoroscopy run. The review and control modules shall support backlighting to allow easy visibility of buttons in e.g. dark examination rooms and give user guidance. The UI modules at or near the table shall allow control under sterile conditions (e.g. under a sterile cover). The user shall be able to detach the UI modules from the pedestal resp. table and attach them to the table resp. pedestal. The user shall be offered a procedure based workflow by having procedure specific settings for a procedure step. A procedure card determines the procedure specific settings, being: The EPX selections the user can choose from during the procedure and a default EPX. The default preset for FlexVision. The default preset for FlexSpot. The default preset for Second FlexSpot. The patient orientation and surgical view Procedure guidance. Specific settings for compatible room specific auxiliary systems Predefined procedure cards shall be delivered with the system. The user can customize procedure cards and organize/group them. The user can set which procedure card shall be used when a procedure step is selected for acquisition. This can also be set automatically based on information from the WLM system (RIS code). FlexVision and FlexSpot shall have the same user interface design for flexible viewing. The keyboard and mouse button behavior shall be harmonized throughout the different system applications in the control room.
UNS.SmartSuite.ORTable | Control all geometry movements from a single control panel; Attach and detach the TSO and TSM from the table rail during the clinical procedure; Choose different table tops; Store and recall table position and set the table to stored positions; To use the system with reverse table top and 180 deg. pivot. | SRS.Allura.Intgr.Magnus ¦ SRS.Allura.Conf.UI-RelatedItems ¦ SRS.Allura.Conf.Table-RelatedItems ¦ SRS.Allura.Intgr.TruSystem ¦ SRS.Allura.UI.UIModule | Compatibility with Magnus OR table of Maquet. Move table via system’s UI modules and Magnus UI modules to position table top, tilt, cradle and set height. Move table for reset geometry, APC iso-center and FDPA. Stop table movements in case of emergency stop. Detect collisions and allow override table movements. Synchronize table lock setting, patient orientation setting, table positions, status and capabilities. Table: sub tag item monoplane biplane .xs X-ray switches -1 footswitch 3f + -2 footswitch 6f + -3 wireless footswitch 3f o -4 wireless footswitch 6f o -5 additional footswitch CR 3f o -6 additional footswitch CR 6f o -7 additional footswitch ER 3f o -8 additional footswitch ER 6f o -9 X-ray handswitch + + Table: sub tag item Monoplane Monoplane Monoplane XY (fm/fa) XY (fm/fa) XY (fm/fa) biplane biplane biplane sub tag item AD7 XT AD7 XNT Maq AD7 XT AD7 XNT Maq Tru AD7 XT AD7 XNT Maq ui UI-modules -1 ControlModule NT +1) -2 ControlModule Tilt +1) -3 ControlModule Flex +1) +1) -4 ControlModule ORT o2) o2) +1,3) -5 ControlModule ORT Flex o2) o2) +1,3) -6 ControlModule Geo NT7) +1) -7 ControlModule Geo Tilt7) +1) -8 ControlModule Geo ORT7) o2) o2) + -9 ControlModule Geo NT IBR +1) -10 ControlModule Geo Tilt IBR +1) -11 ControlModule Imaging o4) o4) o4) o4) o4) o4) + + + -12 speed controller o5) o5) o5) o5) o5) o5) -13 pan handle o o o o o o -14 TouchScreenModule +6) +6) +6) +6) +6) +6) +6) +6) +6) -15 review module + + + + + + + + + -16 viewpad (cardio) # # # # # # # # # -17 viewpad (mixed) # # # # # # # # # -18 ControlModule Geo ORT IBR o2) o2) + Notes: 1) Additional ControlModule optional for CR 2) Optional as additional ControlModule on pedestal for ER 3) On pedestal 4) ControlModule imaging optionally available for ER (not in combination with 2)) 5) Required for FDPA, not an option for TruSystem 6) Additional TSM optional for CR and (on pedestal) for ER 7) Not for NPI, applicable for modules already present in the installed base Table: sub item Table type Table type Table type notes tag AD 7XT AD 7XNT Maquet TruSystem .mi miscellaneous -1 TSM swingarm + + + -2 Pedestal o o + trolley to attach UI modules; for AD7X(N)T including Pedestal Wall Conn. Box; for ORT including Surgery Wall Conn. Box -3 Two-way Intercom o o o -4 Keyboard + Mouse in CR + + + in CR; additional second keyboard + mouse in case of additional FlexSpot -5 Mouse in ER o o o Requires a license -7 MultiVision Switch o o o see SRS.Allura.Intgr.Multivision -8 KVM-Switch o o o see SRS.Allura.Intgr.KVM-Switching, requires DVI based video infrastructure -10 MCS Auxiliary box + + + for viewpad-receiver and X-on light on 3rd party booms Table: sub AD7XT AD7XNT Maquet tag item TruSystem notes . ptf Features -1 neuro tabletop # # -2 cardio tabletop # # -3 universal tabletop % -4 carbonfiber & neuro tabletops % variants A0…A7 for Maquet variants Carbon Floatline, Carbon Spine, SQ14-Xtra and universal for TruSystem -5 tabletop brake option o o VA kit: when power off, friction brake is applied to prevent movement of tabletop -6 table pivot option o o + see SRS.Allura.Func.PivotRotationTableBase -7 table swivel option o o only for Clea-F, see SRS.Allura.Func.SwivelMovementTableBase -8 motorized -float o o + motorized longitudinal and lateral tabletop movements -9 table tilt option o + see SRS.Allura.Func.TableTilting -10 table cradle option o + see SRS.Allura.Func.TableCradle -13 floor spacer o o ‘retention plate’, as mitigation against spillage in OR .pta AD7XT / 7XNT accessories -1 table-mounted radiation shield o o -2 peripheral filters o o -3 arm supports o o -4 head support o o -5 patient straps + + Mandatory in combination with AD7XT and AD7NXT -6 rail accessory clamp o o -7 tabletop accessory clamp o o -8 mattress o o -9 table covers (non-sterile) o o in hybrid OR, table-base, for cleaning purposes -10 additional OP-rail o o extended positioning for UI modules -11 auxiliary OP-rail kit o o positioning for third party equipment -13 neuro wedge o o for patient posturing -14 table rail cable guide + + -15 ratchet compressor o o compression belts with Ratchet-winding mechanism for symmetrical compression -16 handgrip and clamp set o o table clamps and handgrips -17 neuro head holder c c -18 Leg support board o o For use in AngioCT Compatibility with TruSystem 7500 OR table of Hillrom In combination with the FlexArm: Move table via system's UI modules and TruSystem UI modules to position table top, tilt, cradle and set height. Move table for reset geometry, APC iso-center. Stop table movements in case of emergency stop. Detect collisions and allow override table movements. Synchronize table lock settings, patient orientation setting, table positions, status and capabilities. Excluded are: FDPA, iso-centric tilt, and Syncra tilt. In combination with the FlexMove there is limited integration: Table can be locked by the system The system's emergency stop also stops all active table movements The review and control modules shall support backlighting to allow easy visibility of buttons in e.g. dark examination rooms and give user guidance. The UI modules at or near the table shall allow control under sterile conditions (e.g. under a sterile cover). The user shall be able to detach the UI modules from the pedestal resp. table and attach them to the table resp. pedestal.
UNS.SmartSuite.OperateSystem | Operate the system such that the movements of the system are effective and in control, both manually and via the operator controls. | SRS.Allura.Func.MovementDisable/Enable ¦ SRS.Allura.SLS.movement-integrity ¦ SRS.Allura.SLS.unexpected-manual-behavior | A disable movement function is provided to avoid accidental activation of movement related functions of table and available stands by laymen (e.g. cleaners). Programmed movements will not be available when movement is disabled. The disable function cannot be activated during programmed movements. Default state after power-up is enabled. Out of control movements (e.g. spontaneous movements, movements with high speed than requested) shall be prevented in all conditions of the system (during start-up, normal use, power failure, power off). Measures shall be taken to prevent unintended and/or uncontrolled manual movement behavior.
UNS.SmartSuite.OptimalWorkingPosition | Move the FlexArm system to an optimal working position continuous in a Rz rotation range of +135 to -135 degrees while maintaining the selected projection position to create space for e.g. physician, anaesthesia, equipment or visibility on the display. *optimal = a position in the range of 270 degrees minimizing interference with patient accessibility, equipment and viewing angles of the monitors. New, N.A. | SRS.Allura.Func.Z-RotateStandBase(L-ArmRotation) ¦ SRS.Allura.Func.RotateBeam ¦ SRS.Allura.Func.RotateFDXDDetector | Enables the user to rotate the frontal stand for position at the side or head of the table for optimized patient imaging. Iso-center stand rotation around a vertical axis to accomplish improved access to the patient. Function does not apply to the Poly-G floor stand. The user shall be able to rotate the beam around the longitudinal-table-axis (motorized). Intuitive control and display orientation rules (see SRS.Allura.Func.DisplayOrientation) also apply if rotation causes a reversed beam direction. In biplane systems, the frontal and lateral beam can be rotated simultaneously. During motorized movement the rotation position shall be displayed. Enables the user to rotate the frontal FD20 detector for optimized imaging (manual and motorized). User shall be able to rotate the detector by 90° in order to switch between landscape and portrait usage of the detector. Beam width is adapted to maintain the selected detector field size. User shall be able to rotate the detector by any angle between 0° and 90° in order to align with specific anatomy. In FlexArm (XY-fa) and biplane (Clea3 & L-arc3 only) configurations, the collimator rotates synchronously to maintain the selected detector field-size. In all other configurations, the collimator is fixed and the beam-width is limited such that the irradiated area is completely visible.
UNS.SmartSuite.PackingandDisposal | Dispose of the system accordingly to the standards/guidelines. | SRS.Allura.Serv.FSInstructions ¦ SRS.Allura.UI.Help | Instructions are available on how to unpack, install, de-install, decommission, and packing and disposal of the system and components. The user shall be able to access integrated help and electronic Instructions for Use. This includes tooltips, task and procedure guidance and instruction material. The user shall be able to search for a specific topic within the Instructions for Use.
UNS.SmartSuite.ParallelWorking | Simultaneously run image acquisitions and work on other tasks in the Control Room on the same patient during X-ray generation; Review and process previously and intra-procedurally acquired images Perform QA (without interrupting an examination) Film (i.e. select/flag images) and export images; Patient administration. Simultaneously run image acquisitions and work on other tasks in the Control Room on different patients during X-ray generation; Schedule an examination on another patient in advance to the current examination; Schedule an examination on another patient during (and without interrupting) the current examination; Review, post process, measure, print and transfer images of one patient in parallel to performing an examination on another patient; Perform QA during (without interrupting an examination); Film (i.e. select/flag images) and export images; Display data of the current patient in the ER, coming from an auxiliary system (e.g. PACS) to enable reviewing previous studies of this patient; Use the display(s) in the ER to look at prior images of the current patient or to perform a QA or Bolus Chase Reconstruction (BCR) for the current patient; Always see the Live X-ray while simultaneously performing patient administration/scheduling/QA task; Work simultaneously on different touch screen modules in different applications in Exam Room and Control Room. New, Tailored configuration of the workspot with the possibility of adding additional monitors (up to 6) with a multiswitch capability. | SRS.Allura.Intgr.Multivision ¦ SRS.Allura.Func.Concurrency ¦ SRS.Allura.Func.PriorityConstraints ¦ SRS.Allura.Intgr.SwitchableMonitors ¦ SRS.Allura.Intgr.Flexvision | MultiVision enables to display video on the exam room display originating from multiple video sources. The user can select the current video source for display. Video can originate from the system itself and/or from auxiliary systems (e.g. Interventional Workspot, IVUS). Not available in combination with Switchable Monitors or FlexVision. In a DVI configuration, the control room review monitor is selectable as video source and max. 4 other video sources are configurable. MultiVision requires a license This addresses the possibility to use the same or different functions by multiple users at the same time. Administrative functions can be performed in the control room concurrent with activities in the exam room with exception of Bolus Chase Reconstruction and 2D-Quantitative Analysis. Local or remote service monitoring can be performed while the system is in use. Distribution/networking, printing and archiving can be performed with progress in the background. Based on switch-to licenses, auxiliary systems can be operated concurrently. The control room user can work on the same or a different patient and/or study from the exam room. (With FlexSpot) Two control room users can perform different functions on the same patient/study. Priority Handling In case of conflicting or ambiguous user requests the following constraints shall apply: Overruling priority (high to low): 1) exposure, 2) fluoro, 3) other functions (except movement requests and concurrent working in control room) that are applied to the current acquisition study. A function shall be executed as long as requested from at least one UI, with exception of motorized movement requests: parallel movement requests shall be served as long as they are not conflicting. External (surgery) tables are equipped with their own user interface modules. For Maquet Magnus, parallel table movement requests shall result in stopping of the motorized table movements. A movement request via a TruSystem UI will stop a table movement requested via the system UI, without starting the newly requested movement. A table request via the system UI is ignored as long as a table movement request via the TruSystem is active. Exception Handling The system shall provide an option to flexibly switch during the procedure all available video sources to monitors in the examination room (within or outside MCSs/booms) and control room. Video sources can be displayed on at most 16 monitors. Switchable Monitors requires a license. In combination with DVI based video infrastructure, if FlexVision is present at most 8 switchable monitors are supported. In VoIP based infrastructure the license is a counted license for the number of Switchable Monitors. FlexVision enables flexible viewing for displaying multiple video sources (live images but also video from other systems) on one large screen display in the exam room with which the user can determine the layout, the position and size of multiple video sources on one large screen display in the exam room. Layouts are predefined. A layout divides the large screen in viewports. The user shall be able to select a preset (a preset contains the layout and for each viewport the specific video source shown in the viewport) to be applied on the large screen. The user shall be able to create, delete, change, save and rename a preset. The user shall be able to select which source is displayed in which viewport. The user shall be able to hide/unhide the video source display in a viewport. The user shall be able to swap the video sources of two viewports in a preset. The user shall be able to enlarge a specific viewport and restore it to its original size. The user shall be able to display video sources in native resolution. Display of 2k2 images shall be supported. Video can originate from the system itself and from auxiliary systems (configurable set, e.g. Interventional Workspot, IVUS). In combination with VoIP based infra, FlexVision requires a license. The display of 2K2 images requires an additional license.
UNS.SmartSuite.ParkBeam | Manually and motorized (user controlled) park/position the system (if temporarily not used) out of the way to improve access to the patient. New, N.A. | SRS.Allura.Func.ParkorStandbyBeam | Enables the user to move the frontal and/or lateral stand out of the way to a standby or park position to improve access to the patient if the beam is temporarily not used. The different ways to move the stand to a park / standby position depend on the particular configuration (examples: lateral movement for Flexarm standby, longitudinal movement in ceiling stands, swing stand base in a Poly-G floor stand).
UNS.SmartSuite.PatientAndBeam | Operate the system from multiple user positions (e.g. Table Side controlled, pedestal side controlled and remotely controlled from the Control Room); Rotate the frontal and lateral beam simultaneously in biplane mode; Get user guidance/feedback about geometry positions, e.g. height, SID, cradle/tilt, shutter and wedge position. New, The frontal Arc to rotate independently of the detector and collimator, guaranteeing patient-oriented images in every angulation (Azurion 7 B20/15 LN). | SRS.Allura.Func.MoveBeam ¦ SRS.Allura.Func.APCTable ¦ SRS.Allura.Func.RotateBeam ¦ SRS.Allura.Func.SwingStandBase ¦ SRS.Allura.Func.ChangeSIDByDetector ¦ SRS.Allura.Func.XperCT ¦ SRS.Allura.Func.RotateFDXDDetector ¦ SRS.Allura.Func.APCStore/Recall ¦ SRS.Allura.Func.DisplayActualPosition ¦ SRS.Allura.Qual.ClinicalResponseTimes.Movement ¦ SRS.Allura.Func.MoveBeamXY ¦ SRS.Allura.Func.AngulateBeam ¦ SRS.Allura.Func.Z-RotateStandBase(L-ArmRotation) | The user shall be able to move the frontal and/or lateral beam longitudinal to select a clinical area for acquisition (manual and motorized). This function only applies to ceiling suspended configurations. After the user has brought a point-of-interest into the system’s iso-center, the user shall be able to store the corresponding table position (i.e. table longitudinal, lateral, height, tilt and cradle). The system shall provide a default point-of-interest (EPX defined). When a new acquisition study is selected the 'stored table position' is reset such that the default point-of-interest is in the iso-center. The user shall be able to recall the stored table position or only the table height. The system shall indicate to the user if the table is in the stored position. Note: Not applicable for Hillstrom TruSystem table with universal table top and Maquet Magnus table with universal table top. On TruSystem tables with the SQ14-Xtra table top only height can be recalled. On AD7X tables recall of longitudinal and lateral table position requires appropriate motors, which are optional. APC table requires a license. The user shall be able to rotate the beam around the longitudinal-table-axis (motorized). Intuitive control and display orientation rules (see SRS.Allura.Func.DisplayOrientation) also apply if rotation causes a reversed beam direction. In biplane systems, the frontal and lateral beam can be rotated simultaneously. During motorized movement the rotation position shall be displayed. Enables the user to rotate the Poly-G floor stand around a vertical axis far outside the iso-center. Enables the user to image objects (e.g. patient's arm) that are not centered on the table. Enables the user to rotate the stand away from the table for parking purposes. Automatic Position Control (APC) Focus is on applying pre-defined or user-defined positions in controlling table and/or stand movements. Changes the Source Image Distance (SID) by moving the detector along the beam (manual and motorized). The beam width is adapted automatically to maintain the selected detector field-size. This is a special case of DRA or DAR for 3D enabling reconstruction of soft-tissue. This function is only supported with an FD20 detector. The system shall offer dedicated EPX procedures and related settings for Neuro (DRA and DAR) and Abdomen (DRA only) applications. The system shall perform acquisition with up to 60 images/second to obtain acceptable scan time. The system shall be able to apply slit collimation to reduce patient dose. Enables the user to rotate the frontal FD20 detector for optimized imaging (manual and motorized). User shall be able to rotate the detector by 90° in order to switch between landscape and portrait usage of the detector. Beam width is adapted to maintain the selected detector field size. User shall be able to rotate the detector by any angle between 0° and 90° in order to align with specific anatomy. In FlexArm (XY-fa) and biplane (Clea3 & L-arc3 only) configurations, the collimator rotates synchronously to maintain the selected detector field-size. In all other configurations, the collimator is fixed and the beam-width is limited such that the irradiated area is completely visible. The system shall allow the user to store and recall multiple geometry positions (i.e. position of stand, table, shutters and wedges) including the receptor field size. When a new acquisition study is selected, all stored positions shall be removed. The system shall indicate to the user if the system has not reached a selected position. Note: Table position recall is not supported for TruSystem tables in combination with FlexMove. On AD7X tables recall of longitudinal and lateral table position requires appropriate motors which are optional The actual position of table and stand is displayed to the user. Note: Not supported for the TruSystem table, when in combination with the FlexMove. Table: Categories/ items Max. notes: (all values in seconds) Move shutters & wedges 0.5 from UI up to & including image display. APC recall (start geometry movement) 1.0 1.5 for Maquet & TruSystem: 2.5 if the table is in hibernation mode Start all other movements 1.0 1.5 for Maquet & TruSystem: 2.5 if the table is in hibernation mode The user shall be able to move the beam longitudinal and/or transversal (manual and motorized). To select a clinical area for acquisition as an alternative to positioning the tabletop. For standby parking (typically transversal only) during a surgical procedure. For normal parking when the stand is not used. Changes the angle between beam and longitudinal-table-axis (motorized). Angulation can be performed simultaneously with beam rotation. During motorized movement the angle position is displayed. Enables the user to rotate the frontal stand for position at the side or head of the table for optimized patient imaging. Iso-center stand rotation around a vertical axis to accomplish improved access to the patient. Function does not apply to the Poly-G floor stand.
UNS.SmartSuite.PatientComfort | Provide patients with a comfortable position on the table (e.g. ergonomic mattress). New, N.A. | SRS.Allura.Conf.Table-RelatedItems | Table: sub AD7XT AD7XNT Maquet tag item TruSystem notes . ptf Features -1 neuro tabletop # # -2 cardio tabletop # # -3 universal tabletop % -4 carbonfiber & neuro tabletops % variants A0…A7 for Maquet variants Carbon Floatline, Carbon Spine, SQ14-Xtra and universal for TruSystem -5 tabletop brake option o o VA kit: when power off, friction brake is applied to prevent movement of tabletop -6 table pivot option o o + see SRS.Allura.Func.PivotRotationTableBase -7 table swivel option o o only for Clea-F, see SRS.Allura.Func.SwivelMovementTableBase -8 motorized -float o o + motorized longitudinal and lateral tabletop movements -9 table tilt option o + see SRS.Allura.Func.TableTilting -10 table cradle option o + see SRS.Allura.Func.TableCradle -13 floor spacer o o ‘retention plate’, as mitigation against spillage in OR .pta AD7XT / 7XNT accessories -1 table-mounted radiation shield o o -2 peripheral filters o o -3 arm supports o o -4 head support o o -5 patient straps + + Mandatory in combination with AD7XT and AD7NXT -6 rail accessory clamp o o -7 tabletop accessory clamp o o -8 mattress o o -9 table covers (non-sterile) o o in hybrid OR, table-base, for cleaning purposes -10 additional OP-rail o o extended positioning for UI modules -11 auxiliary OP-rail kit o o positioning for third party equipment -13 neuro wedge o o for patient posturing -14 table rail cable guide + + -15 ratchet compressor o o compression belts with Ratchet-winding mechanism for symmetrical compression -16 handgrip and clamp set o o table clamps and handgrips -17 neuro head holder c c -18 Leg support board o o For use in AngioCT
UNS.SmartSuite.PatientData | Prevent mix-up of patient data/information (e.g. identify viewports with name of current patient, or a warning; resolve patient mix); Starting imaging without having entered any patient data (e.g. set-up one of the X-ray protocols as default for emergency cases). New, N.A. | SRS.Allura.Func.Merge ¦ SRS.Allura.Func.ProcedureCard ¦ SRS.Allura.Func.ImplicitSelectionDefaultStudy ¦ SRS.Allura.SLS.consistent-patient-and-image-data | The user shall be able to move a run(s) from one procedure step to a procedure step of another patient. Also, the user shall be able to import a study from media to another patient in the local database. Selecting Study For Acquisition Before the start of acquisition the related study for acquisition is defined. In case a WLM system is connected, a study can consist of multiple scheduled procedure steps. In that case, the user selects a procedure step for acquisition. The user shall be offered a procedure based workflow by having procedure specific settings for a procedure step. A procedure card determines the procedure specific settings, being: The EPX selections the user can choose from during the procedure and a default EPX. The default preset for FlexVision. The default preset for FlexSpot. The default preset for Second FlexSpot. The patient orientation and surgical view Procedure guidance. Specific settings for compatible room specific auxiliary systems Predefined procedure cards shall be delivered with the system. The user can customize procedure cards and organize/group them. The user can set which procedure card shall be used when a procedure step is selected for acquisition. This can also be set automatically based on information from the WLM system (RIS code). If a user didn't select a study/procedure step for acquisition and performs acquisition, the system creates a study/procedure step under which the acquired images are stored. All patient and image data shall be displayed and maintained in a consistent manner throughout the system. Manual changes in patient related data shall only be done after explicit confirmation of the operator.
UNS.SmartSuite.PatientDoseFeedback | Monitor Dose Area Product (DAP) and Air-Kerma (AK; skin-entrance-dose); Receive a warning if the accumulated AK is above a configurable limit. New, N.A. | SRS.Allura.Func.DisplayAndReportPatientDose | This applies to Dose Area Product (DAP) and skin-entrance-dose (air-kerma). Dose info is displayed during and after acquisition. Generated and displayed values comply with legal demands for required accuracy. The user is warned if accumulated air-kerma within the current body zone is above a customizable threshold value (default value 2 Gray). The cumulative (zone-)dose figures are comprised in the study.
UNS.SmartSuite.PatientOrientation | To display the patient’s anatomy (or clinical image) default in such a way that it adheres to the clinical practice (head on top side; left/right as if facing the patient) as commonly used for the various examinations; Prevent diagnostic risks in mistaking left/right patient anatomy; To deviate from image display default e.g. adjusting the display orientation at table side, during the case. New, Indicate the correct patient positioning with respect to the system to ensure proper image display of the region of interest. | SRS.Allura.Func.PatientOrientation ¦ SRS.Allura.Conf.SupportedConfigurations ¦ SRS.Allura.Func.SelectSurgicalView ¦ SRS.Allura.SLS.consistent-image-orientation | The patient can be positioned on the table in one of 4 ways (nose up/down, legs down/up). The actual patient orientation (nose up/down, legs down/up) can be entered into the system. Default orientation is procedure card dependent. The orientation can be changed by the user; the override remains valid until a new procedure card is selected. This setting is also used for intuitive control of shutters and wedges. The following FD configurations shall be supported: Table: Configuration Characteristic discriminators Azurion FD12 Poly-G Ceiling/Floor stand with FD12 detector Azurion FD15 Clea-Floor stand with FD15 detector Azurion FD20 Clea-Ceiling/Floor stand with FD20 detector Azurion FD12/12 Poly-G Floor with FD12 & Larc-C with FD12 Azurion FD20/12 Clea-Floor with FD20 & Larc-N1) or Larc-C with FD12 Azurion FD20/15 Clea-Floor with FD20 & Larc-N or Larc-C with FD15 Azurion FD20 OR-table Clea-Ceiling stand with FD20 detector Azurion FD12/12 OR-table Poly-G Floor with FD12 & Larc-C with FD12 Azurion FD20/12 OR-table Clea-Floor with FD20 & Larc-C with FD12 Azurion FD20/15 OR-table Clea-Floor with FD20 & Larc-N or Larc-C with FD15 Note 1): Azurion FD20/12 Clea-Floor with FD20 & Larc-N with FD12 is not supported for NPI. This configuration is only supported for the installed base. The following configurations are also supported as remanufactured systems: Table: Configuration Characteristic discriminators Azurion FD12 Poly-G Ceiling stand with FD12 detector and AD7 table Azurion FD20 Clea Floor stand with FD20 detector and AD7 table (no FlexMove) Detailed Configuration requirements Applied symbols: 1) In subsequent tables, additional detailing of configurations uses following abbreviations and indications: Mpl = Monoplane system Ce = Monoplane Ceiling system, suspended by longitudinal rail XY-fa = Monoplane Ceiling system, suspended by double-pivot arm (FlexArm) XY-fm = Monoplane Ceiling system, suspended by X-Y rails (FlexMove) Fl = Floor mounted Monoplane system Bpl = Biplane system (FD12/12 and FD20/12) NB = Neuro Biplane (FD20/15) Mag = Magnus Maquet Tru = Hillrom TruSystem 7500 2) Symbols are used in the tables have following meaning + = included in a system, not an option ++ = twice included in a system, not an option (in case of a Biplane system) # = each system comprises exactly one of the items marked with # in one cell % = each system comprises at least one of the items marked with % in one cell. o = option; is an extension based on extra hardware and/or based on a (software) license key c = compatibility based upon interfacing with the system 3) Use of sub-tags To deal with limited table space, sub-tags are used; to be appended to the specified table tag. Example: Table tag = SRS.Allura.Conf.XperFDSeriesConfigurations; sub-tag = .std-1 Resulting complete tag = SRS.Allura.Conf.XperFDSeriesConfigurations.std-1 Enables the user to select between surgical view and diagnostic view. The surgical view is only available on a configuration with an FD20 and FD15. The reversed acquired images must be clearly marked (e.g. inserted icon) to prevent diagnostic mistakes later on. The patient orientation in the X-ray image shall be consistent with the patient to table position and other information as provided by the user.
UNS.SmartSuite.PatientOrientedMovements | To control the FlexArm projections and movements of the gantry (rotate, angulate, translate) by moving the X-ray beam (detector + tube) relative to the longitudinal axis of the table / patient; Navigate to the region of interest through automated FlexArm movements: Image beam rotation, roll, prop, dual Rz rotation, and rail movement in such a way that the user has the perception to steer the X-ray beam. New, N.A. | SRS.Allura.Func.PatientOrientedMovement | In FlexArm (XY-fa) and biplane (Clea3 & L-arc3 only) configurations, if the user angulates or rotates the frontal beam, the detector and collimator angles shall be adapted so to keep the same orientation towards the patient.
UNS.SmartSuite.PatientPositioning | Position the patient on the table; Lock the table to prevent unintended table movement; Adjust the height of the table for ergonomic purposes; Perform interventions via radial, subclavian and groin access; from head, arm and groin position; user standing at left or right side; Indicate the proper patient positioning with respect to the system to ensure proper image display; New, Adjust the table position to optimize the clinical setup. | SRS.Allura.Func.SelectTableSide ¦ SRS.Allura.Func.TableLock ¦ SRS.Allura.Func.ChangingTableHeight ¦ SRS.Allura.Func.PositioningTableTop ¦ SRS.Allura.Func.DisplayActualPosition ¦ SRS.Allura.Func.DisplayOrientation | The user shall be able in table-oriented mode to maintain intuitive movement control from four table sides. The user shall be able to select the table side on every movement-control UI module individually. Combined Patient- And Beam Positioning This section focuses on movements under direct user control, which are not stand or table specific. The user shall be able to lock/unlock: The table pivot movements. The table lateral movements. All table movements. For the TruSystem table in combination with FlexMove, the lateral table lock/unlock is not supported via the system UI. Beam Positioning This section focuses on moving and positioning the stand under direct user control. Moves the table up/down (motorized). The movement range depends on the table type and on the table tilt/cradle angle. This applies to motorized or manual positioning of the tabletop in longitudinal and/or lateral direction. Table-type dependent manual positioning (so called 'float') is balanced (passive or motor-assist) to prevent the patient from sliding down (e.g. in case of tilt). The actual position of table and stand is displayed to the user. Note: Not supported for the TruSystem table, when in combination with the FlexMove. The patient shall be displayed ‘head-up’, facing towards the user, regardless of patient orientation on the table (nose/legs up/down), and stand/detector orientation. This ‘diagnostic view’ adheres to standard clinical practice and prevents diagnostic risks in mistaking left/right patient anatomy. For the lateral channel in a biplane system the same rule applies if the detector is rotated above the table, the image shall not be flipped if the lateral stand is rotated towards a lateral position. There are 2 exceptions: if ‘surgical view’ is chosen, the image will be left/right reversed in case a patient is positioned nose down (rationale: in order to have intuitive eye-hand coordination). in neuro applications (EPX controlled) the lateral-channel display will be left/right reversed (rationale: based on neuro preference; the lateral view has no risk of misinterpreting left/right anatomy).
UNS.SmartSuite.Pedestal | Stand in between the pedestal and the table during the procedure; Move the pedestal (in relation to footswitch); Operate the system with minimum footprint of the pedestal; Operate the system with minimum loose cables in the working area; Operate the system with minimal cables on the floor (e.g. between wall and pedestal) during imaging procedures and open surgery (to prevent tripping hazard, and blocking transport of equipment); Operate the system with minimum interference of wires of the control units; Easily and effectively clean the pedestal and related wires; Move the Control Module, Imaging Module and TSM from the pedestal to the table and vice versa. New, N.A. | SRS.Allura.UI.Workspot ¦ SRS.Allura.Conf.UI-RelatedItems ¦ SRS.Allura.SLS.patient-contamination | Movement and X-ray generating functions shall be controllable at or near the table (e.g on a pedestal). Table: sub tag item monoplane biplane .xs X-ray switches -1 footswitch 3f + -2 footswitch 6f + -3 wireless footswitch 3f o -4 wireless footswitch 6f o -5 additional footswitch CR 3f o -6 additional footswitch CR 6f o -7 additional footswitch ER 3f o -8 additional footswitch ER 6f o -9 X-ray handswitch + + Table: sub tag item Monoplane Monoplane Monoplane XY (fm/fa) XY (fm/fa) XY (fm/fa) biplane biplane biplane sub tag item AD7 XT AD7 XNT Maq AD7 XT AD7 XNT Maq Tru AD7 XT AD7 XNT Maq ui UI-modules -1 ControlModule NT +1) -2 ControlModule Tilt +1) -3 ControlModule Flex +1) +1) -4 ControlModule ORT o2) o2) +1,3) -5 ControlModule ORT Flex o2) o2) +1,3) -6 ControlModule Geo NT7) +1) -7 ControlModule Geo Tilt7) +1) -8 ControlModule Geo ORT7) o2) o2) + -9 ControlModule Geo NT IBR +1) -10 ControlModule Geo Tilt IBR +1) -11 ControlModule Imaging o4) o4) o4) o4) o4) o4) + + + -12 speed controller o5) o5) o5) o5) o5) o5) -13 pan handle o o o o o o -14 TouchScreenModule +6) +6) +6) +6) +6) +6) +6) +6) +6) -15 review module + + + + + + + + + -16 viewpad (cardio) # # # # # # # # # -17 viewpad (mixed) # # # # # # # # # -18 ControlModule Geo ORT IBR o2) o2) + Notes: 1) Additional ControlModule optional for CR 2) Optional as additional ControlModule on pedestal for ER 3) On pedestal 4) ControlModule imaging optionally available for ER (not in combination with 2)) 5) Required for FDPA, not an option for TruSystem 6) Additional TSM optional for CR and (on pedestal) for ER 7) Not for NPI, applicable for modules already present in the installed base Table: sub item Table type Table type Table type notes tag AD 7XT AD 7XNT Maquet TruSystem .mi miscellaneous -1 TSM swingarm + + + -2 Pedestal o o + trolley to attach UI modules; for AD7X(N)T including Pedestal Wall Conn. Box; for ORT including Surgery Wall Conn. Box -3 Two-way Intercom o o o -4 Keyboard + Mouse in CR + + + in CR; additional second keyboard + mouse in case of additional FlexSpot -5 Mouse in ER o o o Requires a license -7 MultiVision Switch o o o see SRS.Allura.Intgr.Multivision -8 KVM-Switch o o o see SRS.Allura.Intgr.KVM-Switching, requires DVI based video infrastructure -10 MCS Auxiliary box + + + for viewpad-receiver and X-on light on 3rd party booms The system shall be designed for: Cleaning and disinfection of system parts in the patient environment. Creating an aseptic environment.
UNS.SmartSuite.PhysioAcquisition | Obtain X-ray image information in relation to the heart cycle e.g. by acquisition, display and storage of physiological signals (limited to electrical signals, i.e. ECG) together with the X-ray images; Manually switch on/off the overlay of the physiological signals on the X-ray images. New, N.A. | SRS.Allura.Func.PhysioAcquisition ¦ SRS.Allura.Intgr.PhysioEquipment(ECG) ¦ SRS.Allura.Func.SelectPhysioDisplay | The system can acquire physiological signals together with the X-ray images. This requires a license. Number of channels (up to 4) is configurable. Only compatible with electrical signals (e.g. ECG; no compatibility with pressure, etc.). EPX determined storage (on/off) of all inputs; recording only in parallel with exposure. Image Generation Scenarios This section focuses on acquisition of images during a controlled geometry movement. Dynamic Rotational Angiography In a DRA scan the X-ray beam rotates around the patient's area of interest placed within the iso center. Compatibility with ECG signals from haemo equipment (acquisition of physiological signals, see SRS.Allura.Func.PhysioAcquisition). This function enables to switch on/off physio data as graphics overlay. The correct time relation between images and physio on display must be preserved. In case of multiple physio sequences, the user can select one recorded channel for display. Dynamic Viewing
UNS.SmartSuite.PointerOnViewingContainer | Control a visible pointer on a viewport to indicate to others an area of interest of an image (e.g. finger on TSM in ER, or mouse in CR); See the pointer on all instances of that viewport. New, N.A. | SRS.Allura.UI.Image.Pointer | The user shall be able to show and move a pointer over an area of interest in an image to support collaboration between exam room and/or control room. This requires a license.
UNS.SmartSuite.PowerFailure |  | SRS.Allura.Func.PowerFailureBehaviour ¦ SRS.Allura.Intgr.ExternalThreePhaseUPS ¦ SRS.Allura.Intgr.BackupPowerActive ¦ SRS.Allura.Conf.Table-RelatedItems ¦ SRS.Allura.Conf.RemainingComponents ¦ SRS.Allura.Intgr.AuxiliarySystems | When the mains power fails: Safe operation shall remain guaranteed (e.g. block non-balanced movements, unblock manual movements to enable to free patient). Stored patient and system data shall be preserved (except for the last run if no Uninterruptible Power Supply (1 phase UPS) is installed). If the 1 phase UPS is installed, the system shall be shut down gracefully instead of an abrupt power off. After power is restored, normal operation shall be resumed after pressing the Power on button. Room Services The system shall support powering from a dedicated external 3-phase UPS in case of a power outage of the hospital mains. The system shall provide an interface to switch to reduced power in case of a power outage of the hospital mains. Configurable at installation, high-power consumption functions (like exposure) are switched off in case the hospital runs on backup power or the system is powered from an external three phase UPS. Low-load fluoro and functions to free a patient remain available. Mechanical Integration Table: sub AD7XT AD7XNT Maquet tag item TruSystem notes . ptf Features -1 neuro tabletop # # -2 cardio tabletop # # -3 universal tabletop % -4 carbonfiber & neuro tabletops % variants A0…A7 for Maquet variants Carbon Floatline, Carbon Spine, SQ14-Xtra and universal for TruSystem -5 tabletop brake option o o VA kit: when power off, friction brake is applied to prevent movement of tabletop -6 table pivot option o o + see SRS.Allura.Func.PivotRotationTableBase -7 table swivel option o o only for Clea-F, see SRS.Allura.Func.SwivelMovementTableBase -8 motorized -float o o + motorized longitudinal and lateral tabletop movements -9 table tilt option o + see SRS.Allura.Func.TableTilting -10 table cradle option o + see SRS.Allura.Func.TableCradle -13 floor spacer o o ‘retention plate’, as mitigation against spillage in OR .pta AD7XT / 7XNT accessories -1 table-mounted radiation shield o o -2 peripheral filters o o -3 arm supports o o -4 head support o o -5 patient straps + + Mandatory in combination with AD7XT and AD7NXT -6 rail accessory clamp o o -7 tabletop accessory clamp o o -8 mattress o o -9 table covers (non-sterile) o o in hybrid OR, table-base, for cleaning purposes -10 additional OP-rail o o extended positioning for UI modules -11 auxiliary OP-rail kit o o positioning for third party equipment -13 neuro wedge o o for patient posturing -14 table rail cable guide + + -15 ratchet compressor o o compression belts with Ratchet-winding mechanism for symmetrical compression -16 handgrip and clamp set o o table clamps and handgrips -17 neuro head holder c c -18 Leg support board o o For use in AngioCT Table: sub tag item availability notes .ps Partner Systems See for details corresponding requirement in Room-Specific Auxiliary Systems section -1 DoseAware Xtend o .mi Miscellaneous -1 Ceiling Suspended radiation shield o -2 Examination light o -3 X-ray spacer o limits minimum Source Skin Distance (Health and Human Services (HHS) requirement). Mandatory for 21 CFR 1020.30. -4 Drip stand o -5 Medical DVD recorder o Requires DVI based video infrastructure -6 1 phase UPS o see SRS.Allura.Func.PowerFailureBehaviour -9 UPS control c see SRS.Allura.Intgr.ExternalThreePhaseUPS -10 VideoSlave o see SRS.Allura.Intgr.VideoSlave -11 Open video interface o see SRS.Allura.Intgr.OpenVideoInterface.LargeScreen, -12 DVD burner o see SRS.Allura.Func.StoreToMedia -13 HD output o see SRS.Allura.Intgr.OpenVideoInterface.LargeScreenDownscaled; in DVI based video infrastructure it requires FlexVision open video interface -14 Video scan converter o for 50/60 Hz SLR TV output, requires DVI based video infrastructure -15 Wall-connection boxes o see SRS.Allura.Intgr.FHDVideoSources, requires DVI based video infrastructure -16 Auxiliary System Input Connection Box o see SRS.Allura.Intgr.FHDVideoSources, requires IP based video infrastructure -17 4K Auxiliary System Input Connection Box o see SRS.Allura.Intgr.4KVideoInputOutput, not available in combination with DVI based video infrastructure Core Functionality Core functionality is defined as functionality which is not dependent on the presence of any license key. The following functions are in the core: Clinical core functionality is a subset of all functionality as defined in chapter FUNCTIONAL REQUIREMENTS. The subset is determined by subtracting all functionality that is dependent on any license key. All functionality needed to install a system and to perform basic checks to initially verify correct system working condition. User messages (see SRS.Allura.Func.DisplayUserMessages). Tooltip (balloon) help. Software-Keyed Clinical Options A Clinical Option gives additional clinical functions/features/characteristics and is enabled through a software license key. Two types of licenses exist: counted license: used to implement licenses that restrict the number of items that can be used, e.g. number of video outputs. uncounted license: used to implement licenses that restrict the usage of a certain feature (on/off) With FlexSpot it shall be possible to install dedicated Philips auxiliary systems in the cabinet in the technical room. FlexSpot shall provide infrastructure to gain access from control room to the media devices available on these PCs; on DVI based video configurations the maximum number of PCs from dedicated Philips auxiliary systems installed in the cabinet in the technical room is three. on VoIP based video configuration the maximum number of PCs from dedicated Philips auxiliary systems installed in the cabinet in the technical room is four. The maximum number of RTO clients in the system (TR/CR/ER) is five for VoIP based configurations, and six for DVI based configurations. Without FlexSpot it shall be possible to install a single PC from dedicated Philips auxiliary systems in the cabinet in the technical room. However no further infrastructure except power and network access is provided.
UNS.SmartSuite.Pre-procedurePlanning | Use/reuse pre-procedure planning from other systems (e.g. recall projection angles planned on HeartNavigator) for more complex cases. New, N.A. | SRS.Allura.Func.DICOM.Import ¦ SRS.Allura.SLS.IHE ¦ SRS.Allura.Qual.ClinicalResponseTimes.Arch | Enables the user to manually import images from a remote destination (workstation or archive). The user selects a remote destination, determines the image data and imports it; or on a remote destination, the user determines the image data and pushes it to the system. An original biplane run is reconstructed from separate (frontal and lateral) DICOM objects. The system implements all of the transactions specified in the Integrating the Healthcare Enterprise (IHE) Technical Framework to support the IHE Integration Profiles, Actors and Options listed below: Table: Integration Profiles Implemented Actors Implemented Options Implemented Cardiac Catheterization Workflow (CATH) Acquisition Modality Broad Worklist Query Scheduled Workflow (SWF) Acquisition Modality Patient Based Worklist Query, Broad Worklist Query Patient Information Reconciliation (PIR) Acquisition Modality none Consistent Presentation of Images (CPI) Acquisition Modality none Radiation Exposure Monitoring (REM) Acquisition Modality none Consistent Time (CT) Time Client none Audit Trail and Node Authentication (ATNA) Secure Node none performance values apply to transfer over a network with a speed of at least 1Gbps. Table: Categories/ items Max. notes: (all values in seconds) Export 10 runs * 100 images max. 1K2 200 multiframe, 8 bit, with Jpeg Import 40 images max. 1K2 900 singleframe 12bit, with & without Jpeg Export 40 images max. 1K2 100 singleframe 12bit, with & without Jpeg Printing 1 sheet with 12 images 60 Printer resolution satisfying source resol. Field-Service Response Times
UNS.SmartSuite.Preparation | Obtain information that is needed to ensure the rightful preparation for a service activity before visiting the system. | SRS.Allura.Serv.FunctionalTests ¦ SRS.Allura.Serv.RemoteConnectivityStatus ¦ SRS.Allura.Serv.HardwareTests ¦ SRS.Allura.Serv.NetworkDiagnosis | The system shall provide tests to check the functional integrity of the system or part of the system. The tests shall provide information to enable the FSE to find a fault by reasoning, based on supplied documentation. Appropriate (Edoc) documentation shall be provided to guide the FSE in performing successive tests. The system shall inform the clinical user about the remote connection status The system shall provide a remote connection test for clinical user and hospital IT. For VPN connected systems it shall be possible to up- and download remote configuration settings from/to the system. Applied to diagnose the correct functioning of the Field Replaceable Units (FRU’s) and related interfaces. Power-On SelfTests (‘POST’) are performed automatically at system power-up. Build-In SelfTests (‘BIST’) enable the FSE to examine a particular FRU into greater detail. The system shall provide means to detect and diagnose network connectivity issues for the (system internal) control network for the real time image link for the connection to external DICOM nodes on the hospital network for remote service connection
UNS.SmartSuite.Printing | Select images and create a film sheet from these images; Print the film sheet; Preview the film sheet prior to printing; Select printer and modify printer settings; Submit print job and manage print queues; Create and print film sheet at table side prior/during/after an examination; Print via DICOM and non DICOM printer. New, N.A. | SRS.Allura.Qual.ClinicalResponseTimes.Arch ¦ SRS.Allura.Intgr.LocalPaperPrinters ¦ SRS.Allura.Intgr.NetworkFilmPrinters ¦ SRS.Allura.Func.PrintSheetCreation ¦ SRS.Allura.Func.PrintJobControl | performance values apply to transfer over a network with a speed of at least 1Gbps. Table: Categories/ items Max. notes: (all values in seconds) Export 10 runs * 100 images max. 1K2 200 multiframe, 8 bit, with Jpeg Import 40 images max. 1K2 900 singleframe 12bit, with & without Jpeg Export 40 images max. 1K2 100 singleframe 12bit, with & without Jpeg Printing 1 sheet with 12 images 60 Printer resolution satisfying source resol. Field-Service Response Times Compatibility with general off-the-shelf local and network 2D printers. Injector Integration The system can be connected to Network Film Printers. Compatibility according DICOM (Basic Grayscale Print Meta class). Supported DICOM features: Gray Scale Display Standard, Presentation LUT (Look Up Table). Support for most commonly used printers to guarantee printing image quality. Compatibility With Media This function enables manual selection and composition of multiple images on a film sheet. This requires a license. The system presents an already filled sheet. Default print settings can be customized. The user can change the layout and/or add/delete/re-position images. A preview function is available to observe the result before actual printing. This function enables the user to start and control the actual printing process. This requires a license. The user can modify customization-based defaults regarding printer settings and destination. The content can be deidentified and the user can define a fictive patient name for the deidentified content. The user can submit the job, view job status, abort, delete and retry the job. Dose Reporting
UNS.SmartSuite.ProcedureCards | Select the system set-up, edit, customize, and create their own specific ProcedureCards including quick-save during procedure and the ability to save under different groups (e.g. per clinical discipline, hospital department or per user); Select the system set-up (based on the above parameters) from the ProcedureCard with a single click; Select a ProcedureCard from at least 20 default ProcedureCards; at least 3 per clinical area; Automatically select the configured ProcedureCards, based on the RIS/HIS/CIS code of the scheduled procedure; Select a ProcedureCard within a maximum of 3 clicks once the user selects the patient and there is no RIS/HIS/CIS code coupling; Import, modify/adapt, configure and export (default) procedural information (ProcedureCards) including at least 6 parameters: Short description of procedure; Set of X-ray protocols which should be directly available related to the clinical procedure (incl. geometry position, APC position and injector coupling); Patient orientation; A pre-set screen layout which will be used as default for viewing in ER (FlexVision); A pre-set screen layout which will be used as default for viewing in CR (FlexSpot); On-screen display of hospital specific checklists & protocols s in exam room and control room; Injection Protocol (only in case off SmartCT). New, N.A. | SRS.Allura.Func.PredefinedAnnotation ¦ SRS.Allura.UI.Help.Customize ¦ SRS.Allura.Func.SystemCustomization ¦ SRS.Allura.Func.PatientOrientation ¦ SRS.Allura.Func.ProcedureCard ¦ SRS.Allura.Func.Workflow.Flexibility | During manual annotation, a list can be displayed containing predefined and user defined text strings, which can be inserted within the current annotation block. The user shall be able to customize this list: The predefined annotations can be deleted and modified. The user defined annotations can be added, deleted and modified. The user shall be able to delete and modify a help header and link a new help header to an imported hospital protocol/checklist. This requires a license. Allows the user to customize system settings, that are available for the system's configuration, such as: Date, Time and Regional Settings User Administration and Administration settings List of annotation and performing physician Procedure cards and RIS code mapping Settings for viewing and quantitative analysis Media and DICOM export settings APC positions Linking DICOM export settings, APC positions and procedure help text to EPX procedures Remote support settings Defining APC pathways Saved settings are persistent over system restarts. System customization settings can be imported and exported. Administration Functions This section focuses on administrative functions to deal with patient data and related information. General Administration Concepts A local database includes all information that is entered manually and generated in acquisition. It can contain multiple patients/studies/procedure steps comprising patient data, images, etc. Local Administration Functions The patient can be positioned on the table in one of 4 ways (nose up/down, legs down/up). The actual patient orientation (nose up/down, legs down/up) can be entered into the system. Default orientation is procedure card dependent. The orientation can be changed by the user; the override remains valid until a new procedure card is selected. This setting is also used for intuitive control of shutters and wedges. The user shall be offered a procedure based workflow by having procedure specific settings for a procedure step. A procedure card determines the procedure specific settings, being: The EPX selections the user can choose from during the procedure and a default EPX. The default preset for FlexVision. The default preset for FlexSpot. The default preset for Second FlexSpot. The patient orientation and surgical view Procedure guidance. Specific settings for compatible room specific auxiliary systems Predefined procedure cards shall be delivered with the system. The user can customize procedure cards and organize/group them. The user can set which procedure card shall be used when a procedure step is selected for acquisition. This can also be set automatically based on information from the WLM system (RIS code). The user shall be able to have a workflow according local preferences. Different studies/procedure steps can overlap in time, i.e. The user shall be able to suspend the current acquisition study/procedure step, select another study/procedure step for acquisition and later select the suspended study/procedure step for acquisition again. The user shall be able to select a study for viewing besides an already selected study for acquisition. Sequence of workflow steps is not fixed, i.e. the user shall be able to switch between administration, image generation, viewing, post-processing, printing, reporting, archiving and distribution for the same study. Customizable automated workflow, i.e. Automatic DICOM export. Automatic dose report creation. Automatic DICOM dose structured report. Automatic finishing of the current acquisition study when another study is selected for acquisition. Automatic setting a study to protected against deletion upon finishing of the study. Automatic selecting of procedure card based on RIS code. Concurrency
UNS.SmartSuite.ProvideInformation | Provide the hospital administration with patient/ procedure information during/at the end of the examination (e.g. export onto the hospital information system (e.g. RIS/CIS); Access patient data after closing the examination Export information to the hospital system about what (events and setting) actually has been performed by the system during/at the end of the procedure (e.g. MPPS manager). New, N.A. | SRS.Allura.Intgr.WLMSystems ¦ SRS.Allura.Func.SelectStudyFromScheduledList ¦ SRS.Allura.Func.ShowStudyInfo ¦ SRS.Allura.Func.FinishStudy ¦ SRS.Allura.SLS.IHE | The system can be connected to the hospitals Worklist management system (e.g. RIS/CIS). This requires a license. Compatibility according DICOM conformance statement. Support for DICOM-SCU Worklist and MPPS (Modality Performed Procedure Step) protocol. This enables the user to select a study/procedure step for acquisition from the scheduled study list. The patient name is displayed clearly in the exam room. The system is prepared for acquisition; information about any blockage is shown. The system acquisition parameters are set according to the current EPX application/procedure. For WLM based procedure steps, progress is reported according DICOM MPPS. The user shall be able to view the information of a selected study, comprising at least: Study info: overview of runs, applied EPXs, cumulative dose, fluoro time. Run/image info: exposure parameters, projections. The user shall be able to filter the runs within the selected study. With this function the user indicates that this study is finished on the system. Per procedure step within the study, EPX defined automation tasks are started (e.g. export). After pending jobs (e.g. archiving) are successfully finished, the procedure step is finished and (if MPPS is configured) reported according DICOM MPPS. Unless set to protected, the occupied storage space is released for automatic deletion. Examination Related Presetting EPX Concept The system implements all of the transactions specified in the Integrating the Healthcare Enterprise (IHE) Technical Framework to support the IHE Integration Profiles, Actors and Options listed below: Table: Integration Profiles Implemented Actors Implemented Options Implemented Cardiac Catheterization Workflow (CATH) Acquisition Modality Broad Worklist Query Scheduled Workflow (SWF) Acquisition Modality Patient Based Worklist Query, Broad Worklist Query Patient Information Reconciliation (PIR) Acquisition Modality none Consistent Presentation of Images (CPI) Acquisition Modality none Radiation Exposure Monitoring (REM) Acquisition Modality none Consistent Time (CT) Time Client none Audit Trail and Node Authentication (ATNA) Secure Node none
UNS.SmartSuite.QA | Start Quantitative Analysis; Outline and quantify of contrast-medium filled vessels (2DQCA and 2DQVA) and ventricles (2DLVA and 2DRVA for monoplane and biplane), respectively, in angiographic X-ray images. New, N.A. | SRS.Allura.Func.VesselAnalysis ¦ SRS.Allura.Func.QA ¦ SRS.Allura.Func.VentricularAnalysis | The system shall provide vessel analysis for coronary (QCA) and vascular arteries (QVA). The Quantitative Analysis shall allow the user to indicate a vessel segment in an un-subtracted or (for QVA only) in a subtracted X-ray image. The application shall then automatically detect the contrast-medium filled lumen in that segment and outline its borders as graphical overlays. The user shall be able to manually correct the contour. The application shall also display the profile of the vessel-diameter for the vessel segment and compute characteristics of the stenosis. QCA and QVA both require a separate license. The system shall provide Coronary Analysis, Vascular Analysis, Left Ventricular Analysis and Right Ventricular Analysis. Coronary Analysis (QCA) and Vascular Analysis (QVA) are monoplane applications i.e. analysis is performed on a single image of one channel only. Left and Right Ventricular Analysis (LVA, RVA) are available as mono- and biplane application i.e. in biplane configurations ventricular analysis can be performed on images of one or two channels. The user selects an image or run for analysis, using available image navigation functions. The user selects the desired analysis function, and performs the related applicable workflow. QA-calibration can be applied to express quantitative outcomes on an absolute scale (e.g. in mm). Results are shown on the review monitor but can also be copied to one of the reference viewports A QA report is generated and stored as a Secondary Capture image in the system database. This report can be displayed, exported or printed. Calibration The system shall provide ventricular analysis for the left (LVA) and right ventricle (RVA) After identification by the user of the aortic valve plane and the apex, the LVA application shall automatically trace and display the outline of a contrast-medium filled left ventricle. Biplane LVA shall similarly establish automatic contours for both the frontal and the lateral plane. RVA and LVA application shall allow the user to manually draw contours around the ventricular lumina in end-diastole and in end-systole, and in biplane series for the frontal and the lateral plane, respectively. RVA and LVA then computes ventricular characteristics like volumes, ejection fraction and performs wall motion analysis. LVA and RVA both require a separate license.
UNS.SmartSuite.QualityControl | Perform system quality control checks (in order to assure safety, regulatory compliance and validated system performance (as standardized by NEMA). New, N.A. | SRS.Allura.SLS.ListOfApplicableStandards | The standards as mentioned in chapter APPLICABLE STANDARDS are applicable. Environmental Care
UNS.SmartSuite.QuickSaveDuringProcedure | Save image(s); Save a changed screen layout and save under a different name; Store temporarily the table position with a single click ; Store temporarily the position of the system (beam carrier(s) and table), including the detector orientation, FOV, shutters and wedge positions with a single click. New, N.A. | SRS.Allura.Func.APCReference ¦ SRS.Allura.Func.APCTable ¦ SRS.Allura.Func.APC3DReference ¦ SRS.Allura.Func.APCStore/Recall ¦ SRS.Allura.Intgr.Flexvision ¦ SRS.Allura.Func.APCCommon ¦ SRS.Allura.Intgr.Flexvision.Snapshot | The user shall be able to recall the geometry positions (i.e. position of frontal and lateral stand, table, shutters and wedges) plus the receptor field size which corresponds to an image that is displayed in a live or reference viewport. In a biplane system, the user shall be able to recall a position for both frontal and lateral channels independently or simultaneously. Note: Table position recall is not supported for TruSystem tables in combination with FlexMove. On AD7X tables recall of longitudinal and lateral table position requires appropriate motors which are optional APC reference requires a license. After the user has brought a point-of-interest into the system’s iso-center, the user shall be able to store the corresponding table position (i.e. table longitudinal, lateral, height, tilt and cradle). The system shall provide a default point-of-interest (EPX defined). When a new acquisition study is selected the 'stored table position' is reset such that the default point-of-interest is in the iso-center. The user shall be able to recall the stored table position or only the table height. The system shall indicate to the user if the table is in the stored position. Note: Not applicable for Hillstrom TruSystem table with universal table top and Maquet Magnus table with universal table top. On TruSystem tables with the SQ14-Xtra table top only height can be recalled. On AD7X tables recall of longitudinal and lateral table position requires appropriate motors, which are optional. APC table requires a license. The user shall be able to recall the frontal projection (rotation/angulation) that corresponds to a view on the Interventional Workspot (3DRA, XperCT, etc). Indicated by the Interventional Workspot, the SID shall go to minimum or to maximum (maximum to e.g. prevent hitting the needle) during the recall movement. APC 3D reference requires a license. The system shall allow the user to store and recall multiple geometry positions (i.e. position of stand, table, shutters and wedges) including the receptor field size. When a new acquisition study is selected, all stored positions shall be removed. The system shall indicate to the user if the system has not reached a selected position. Note: Table position recall is not supported for TruSystem tables in combination with FlexMove. On AD7X tables recall of longitudinal and lateral table position requires appropriate motors which are optional FlexVision enables flexible viewing for displaying multiple video sources (live images but also video from other systems) on one large screen display in the exam room with which the user can determine the layout, the position and size of multiple video sources on one large screen display in the exam room. Layouts are predefined. A layout divides the large screen in viewports. The user shall be able to select a preset (a preset contains the layout and for each viewport the specific video source shown in the viewport) to be applied on the large screen. The user shall be able to create, delete, change, save and rename a preset. The user shall be able to select which source is displayed in which viewport. The user shall be able to hide/unhide the video source display in a viewport. The user shall be able to swap the video sources of two viewports in a preset. The user shall be able to enlarge a specific viewport and restore it to its original size. The user shall be able to display video sources in native resolution. Display of 2k2 images shall be supported. Video can originate from the system itself and from auxiliary systems (configurable set, e.g. Interventional Workspot, IVUS). In combination with VoIP based infra, FlexVision requires a license. The display of 2K2 images requires an additional license. During APC movements the SID shall be maximized in order to reduce the chance of a collision. In case of a collision APC movement is stopped. With FlexVision, a single viewport or all viewports at once can be grabbed and stored as image under the current acquisition study.
UNS.SmartSuite.Raillength | Adapt the system to the size of the room by the option of different rail lengths; Be able to move the FlexArm system to the far end of the room to free up space around the table. New, N.A. | SRS.Allura.Conf.Stand-RelatedItems | Table: sub Poly-G Poly-G Poly-G Clea Clea Clea Clea L-arc tag Item Ce XY fm Fl Ce XY fm XY fa Fl Bpl notes .std -1 Extended rail o o o o o o FlexArm: 4.3m; standard: 6.0m or 7.8m. -2 L-arm long # # # # -3 L-arm short # # # # -4 Flexarm long # -5 Flexarm short # -6 Ceiling rail covers o See SRS.Allura.Phys.LaminarAirflowCompatibility -7 Optical imaging system o o o see SRS.Allura.Intgr.OpticalImagingSystem -8 Coverbags c c c c c c c c Arc / Tube / Detector / TSM
UNS.SmartSuite.ReducePatientMovement | Have means to compensate for the patient movements during a scan thus minimizing image artifacts. | SRS.Allura.Func.Roadmap.AutomaticPixelShift | During the roadmap Vessel and Device phases, optionally automatic pixel-shift is applied to the live image (EPX controlled) to compensate for undesired movements.
UNS.SmartSuite.ReviewModule | Perform controlled image review: Work with one hand on the review module and one hand on the mouse in order to create an efficient workflow for reviewing and image processing (e.g. ergonomic module design). Have improved visibility of icons in the dark (e.g. backlighting); New, N.A. | SRS.Allura.Func.MosaicOverview ¦ SRS.Allura.Conf.UI-RelatedItems ¦ SRS.Allura.UI.UIModule ¦ SRS.Allura.Func.Cycle ¦ SRS.Allura.Qual.ClinicalResponseTimes.Review ¦ SRS.Allura.Func.NavigateToImageData | Images are minified and displayed as an N*N mosaic in one or more pages. In Run-Overview the mosaic displays each image in the run. In File-Overview the mosaic displays a single image from each run in the file. In mosaic display mode navigation functions remain available. Table: sub tag item monoplane biplane .xs X-ray switches -1 footswitch 3f + -2 footswitch 6f + -3 wireless footswitch 3f o -4 wireless footswitch 6f o -5 additional footswitch CR 3f o -6 additional footswitch CR 6f o -7 additional footswitch ER 3f o -8 additional footswitch ER 6f o -9 X-ray handswitch + + Table: sub tag item Monoplane Monoplane Monoplane XY (fm/fa) XY (fm/fa) XY (fm/fa) biplane biplane biplane sub tag item AD7 XT AD7 XNT Maq AD7 XT AD7 XNT Maq Tru AD7 XT AD7 XNT Maq ui UI-modules -1 ControlModule NT +1) -2 ControlModule Tilt +1) -3 ControlModule Flex +1) +1) -4 ControlModule ORT o2) o2) +1,3) -5 ControlModule ORT Flex o2) o2) +1,3) -6 ControlModule Geo NT7) +1) -7 ControlModule Geo Tilt7) +1) -8 ControlModule Geo ORT7) o2) o2) + -9 ControlModule Geo NT IBR +1) -10 ControlModule Geo Tilt IBR +1) -11 ControlModule Imaging o4) o4) o4) o4) o4) o4) + + + -12 speed controller o5) o5) o5) o5) o5) o5) -13 pan handle o o o o o o -14 TouchScreenModule +6) +6) +6) +6) +6) +6) +6) +6) +6) -15 review module + + + + + + + + + -16 viewpad (cardio) # # # # # # # # # -17 viewpad (mixed) # # # # # # # # # -18 ControlModule Geo ORT IBR o2) o2) + Notes: 1) Additional ControlModule optional for CR 2) Optional as additional ControlModule on pedestal for ER 3) On pedestal 4) ControlModule imaging optionally available for ER (not in combination with 2)) 5) Required for FDPA, not an option for TruSystem 6) Additional TSM optional for CR and (on pedestal) for ER 7) Not for NPI, applicable for modules already present in the installed base Table: sub item Table type Table type Table type notes tag AD 7XT AD 7XNT Maquet TruSystem .mi miscellaneous -1 TSM swingarm + + + -2 Pedestal o o + trolley to attach UI modules; for AD7X(N)T including Pedestal Wall Conn. Box; for ORT including Surgery Wall Conn. Box -3 Two-way Intercom o o o -4 Keyboard + Mouse in CR + + + in CR; additional second keyboard + mouse in case of additional FlexSpot -5 Mouse in ER o o o Requires a license -7 MultiVision Switch o o o see SRS.Allura.Intgr.Multivision -8 KVM-Switch o o o see SRS.Allura.Intgr.KVM-Switching, requires DVI based video infrastructure -10 MCS Auxiliary box + + + for viewpad-receiver and X-on light on 3rd party booms The review and control modules shall support backlighting to allow easy visibility of buttons in e.g. dark examination rooms and give user guidance. The UI modules at or near the table shall allow control under sterile conditions (e.g. under a sterile cover). The user shall be able to detach the UI modules from the pedestal resp. table and attach them to the table resp. pedestal. This function enables to display a sequence of images in cine mode. In file-cycle all exposure and fluoro runs in the active file of the current study are cycled sequentially. In run-cycle the complete active run is cycled. User can set speed. Intra-Image Post-Processing Focuses on intra-image processing functions for customized enhanced display. General Concepts Table: Categories/ items Max. notes: (all values in seconds ) Image stepping 0.2 Allowed 0.4 sec for run-subtract. Run stepping, non-subtract 0.3 Run stepping, subtract 0.6 Run overview 4*4, subtract & non-subtract 2 Run overview 4*4, run-subtract 4 File overview 4*4, non-subtract 2 File overview 4*4, subtract 4 Reviewing requirements are not applicable to mosaic overview modes, unless explicitly indicated otherwise. Enables the user to zoom-in progressively on a particular file and/or a run, down to a specific image. After selection of a study for viewing or for acquisition, the first exposure or fluoro run is active for viewing. Exposure overrules the selection and makes the latest acquired run active for viewing.
UNS.SmartSuite.Roadmap | Superimpose vessel image during live fluoroscopy (on both the frontal as the lateral channel) to guide a device (catheter, guide wire or other interventional device) through the vascular structure during fluoroscopy in all patient positions; Superimpose vessel image in (un-)subtracted mode. New, N.A. | SRS.Allura.Func.Roadmap.AutomaticPixelShift ¦ SRS.Allura.Qual.ClinicalResponseTimes.Acquisition ¦ SRS.Allura.Func.Roadmap.DevicePhase ¦ SRS.Allura.Func.Roadmap.DevicePhase.VesselSubtract ¦ SRS.Allura.Func.Roadmap.SmartMask ¦ SRS.Allura.Func.Roadmap.Mode ¦ SRS.Allura.Func.Roadmap.DevicePhase.LiveSubtract ¦ SRS.Allura.Func.Roadmap.DevicePhase.VesselUnSubtract ¦ SRS.Allura.Func.Roadmap.DevicePhase.DeviceMask ¦ SRS.Allura.Func.Roadmap.DevicePhase.Misregistration ¦ SRS.Allura.Func.Roadmap.VesselPhase ¦ SRS.Allura.Func.Roadmap.DevicePhase.Processing ¦ SRS.Allura.Func.Roadmap.DevicePhase.Subtraction ¦ SRS.Allura.Func.Roadmap | During the roadmap Vessel and Device phases, optionally automatic pixel-shift is applied to the live image (EPX controlled) to compensate for undesired movements. Table: Categories/ items Max. notes: (all values in seconds) Switch viewing to fluoroscopy 1.8 from pedal press till the first image shown Switch viewing to exposure 1.8 from pedal press till the first image shown, For non-DSA cine Switch viewing to exposure 2.5 from pedal press till the first image shown, For acquisition types with test-shot lock-in and frame-speeds greater or equal to 4 fr/s Switch fluoro to viewing 1.8 from pedal release till display of viewing image on screen Switch exposure to viewing 1.8 from pedal release till display of viewing image on screen Switch between exposure/fluoro 1.8 For acquisition types without test-shot lock-in and frame-speeds greater or equal to 4 fr/s Switch between exposure/fluoro 2.5 For acquisition types with test-shot lock-in and frame-speeds greater or equal to 4 fr/s Digital Subtraction Angiography (DSA): delay after test-image 1.5 Change live fluoro settings 1.5 Fluo flavour, field-size (incl FD-mode). Change live exposure settings 0.5 Frame-rates for multiphase acquisitions Delay between Xray-OFF and Indicators-OFF 1.0 Fluoro pipeline perception latency 0.200 Essential for eye-hand coordination. In this roadmap phase, the fluoro images shall be displayed according the current roadmap mode settings (EPX), meaning in one of the following Device phase modes: LiveSubtract; VesselSubtract; or VesselUnSubtract. Exposure runs can be made in-between Device phase runs without losing the related masks. In this mode, the images shall be displayed subtracted with the DeviceMask resulting into the so-called Device image, and the Device image shall be displayed superimposed by the (faded) VesselMask. Optionally (EPX determined) the Device image shall be displayed superimposed by the LandmarkMask (containing the bones structure, derived from the VesselMask). Alternative to the roadmap Vessel phase, the user can directly select an image from any stored run and use it as the VesselMask (‘SmartMask’) for the Device phase. Before selecting, the smart mask can be enhanced using pixel-shift. Selecting a ‘SmartMask’ requires a license. The user shall be able to select a roadmap mode. Up to 4 roadmap modes shall be present (EPX) to set the exposure-control and image processing-control parameters for the Vessel phase and image processing-control parameters for the Device phase. The current selected fluoro flavour settings are used as the acquisition-control parameters for the Device phase. In this mode, the images shall be displayed subtracted with the DeviceMask resulting into the so-called Device image. Optionally (EPX determined) the Device image shall be displayed superimposed by the LandmarkMask (containing the bones structure, derived from the DeviceMask). In this mode, the images shall be displayed unsubtracted (no DeviceMask), but the images shall be displayed superimposed by the (faded) VesselMask. During acquisition (in the LiveSubtract and VesselSubtract modes), the so-called DeviceMask shall be calculated from a number of received images (EPX determined). The DeviceMask shall be retained for subsequent acquisitions, unless one of the following cases occurs: The user requests a remask of the DeviceMask. Roadmap is deactivated. The user selects another SmartMask. The user selects a fluoro flavour. To prevent mis-interpretation, subtracted display is switched-off automatically and a user message is displayed if the VesselMask or DeviceMask does not match live fluoro (regarding SID and geo projection). In this roadmap phase, a VesselMask from the acquired exposure images shall be created, which will serve as a background image during the roadmap Device phase. Contrast is manually injected during the run and subsequent images are ‘traced’: this is for each particular image spot (a pixel) the minimum occurring value is retained (TraceMin). Alternatively (for CO2 contrast: white vessels), the maximum occurring value is traced (TraceMax); (selection TraceMin/TraceMax is determined by EPX). The images are shown background-subtracted (after X-ray stabilization). On Last-Image-Hold, the resulting image is the VesselMask. After the run is made, the roadmap Device phase is automatically entered. On Last-Image-Hold the user can change settings interactively, also for the next fluoro run: Regular processing functions, e.g. pixel-shift, can be applied to obtain a better match. The intensity of the VesselMask and Device image can be altered independently (called ‘Fading’). During fluoro in the Device phase (after X-ray stabilization), subtraction can be toggled on/off without losing the related masks. The user shall be able to activate or deactivate roadmap. If roadmap gets activated, the last used roadmap mode since EPX application selection is used, otherwise the default roadmap mode for this EPX application is used. Dependent on the roadmap mode, the system starts roadmap in either the Vessel phase or the Device phase. The active roadmap phase is indicated to the user. Selection of a SmartMask is an implicit activation of roadmap. The Device phase is automatically entered. Subtracted imaging requires a license.
UNS.SmartSuite.RollSpeed | Perform a 3D roll scan fast enough for the patient to hold his/her breath. New, N.A. | SRS.Allura.Qual.ClinicalResponseTimes.3D-Acquisition | Table: Categories/ items Max. notes: (all values in seconds) Fastest 3D acquisition head side 5 -120° … +120° Fastest 3D acquisition in nurse/doctor side 6 -90° … + 90°
UNS.SmartSuite.RoomSetup | Position system easy and efficiently (manually/from TSO); Maintain the sterile field when moving the FD system, while using the system over time; Position system so as to minimally disrupt laminar airflow in OR settings; Position system so that the physician can see other staff during the procedure; In case of biplane, work with both channels or work with either the frontal or lateral channel Work with the frontal channel if the lateral channel is in parked position. Position image monitors at a comfortable viewing angle for physician when standing in head/side/arm/left/right position; Have a clear view on the monitor from work position. New, Choose a wide range of screen configurations to match the intended procedure mix of the lab. Screen configurations range from separate loose screens to screens combined on single booms to large screens or combinations of those. | SRS.Allura.Func.MoveBeam ¦ SRS.Allura.Func.SelectActiveChannel ¦ SRS.Allura.Func.SwingStandBase ¦ SRS.Allura.SLS.patient-contamination ¦ SRS.Allura.SLS.patient-contamin-airborne ¦ SRS.Allura.Conf.MCS-RelatedItems ¦ SRS.Allura.Phys.LaminarAirflowCompatibility ¦ SRS.Allura.UI.PositionMCSFrame ¦ SRS.Allura.Func.MoveBeamXY ¦ SRS.Allura.Func.Z-RotateStandBase(L-ArmRotation) | The user shall be able to move the frontal and/or lateral beam longitudinal to select a clinical area for acquisition (manual and motorized). This function only applies to ceiling suspended configurations. In a biplane system this enables the user to select the acquisition channel(s): frontal, lateral or biplane. Enables the user to rotate the Poly-G floor stand around a vertical axis far outside the iso-center. Enables the user to image objects (e.g. patient's arm) that are not centered on the table. Enables the user to rotate the stand away from the table for parking purposes. Automatic Position Control (APC) Focus is on applying pre-defined or user-defined positions in controlling table and/or stand movements. The system shall be designed for: Cleaning and disinfection of system parts in the patient environment. Creating an aseptic environment. The system shall be designed to minimize its contribution to airborne contaminants reaching the patient. Table: sub item notes tag Rail susp 3rd party boom 3rd party boom XL Springarm Springarm XL Up to two rail suspended MCS and in total up to four MCS of different type and configuration can be combined in one system .frm MCS frame -1 Frame + + + + -2 VxH= 1x2 % % % 2 x 27” -3 VxH= 1x2 + 1 % 2 x 19" + 1 x 27" -4 VxH= 2x2 % % % 4 x 27” -5 VxH= 2x2 + 2 % % 6 x 27” -7 Large Screen % % % 58” XL -8 Large Screen + 2 % % % 58” XL + 2 x 27” .dsp MCS displays max. number is configuration dependent; In combination with IP based video infrastructure the total number of all monitors installed on MCSs/booms in ER and FHD monitors installed outside MCSs/booms in CR and ER never exceeds 20. The maximum doesn’t include FlexVision, Slave FlexVision, (Second) FlexSpot, Basic CR, MultiView and Additional MultiView monitors -1 27” diagnostic color % % o % o 27" color monitor, min. 400 cd/m2, resolution 1920x1080, aspect ratio 16:9 -2 32” diagnostic color1) o 32" color monitor, min. 400 cd/m2, resolution 1920x1080, aspect ratio 16:9 -4 19” multi-purpose o 19" color monitor (only for VxH= 2x1 + 1) min. 250 cd/m2, resolution 1280x1024, aspect ratio 5:4 -5 58” XL % % % 58” color monitor, min. 400 cd/m2, resolution 3840x2160, aspect ratio 16:9 -9 Dummy fill o o o .swm Switchable Monitors -1 27” or 32” diagnostic color1) o o o Up to 16 switchable monitors, can also be offered without MCS Note: 1) 32" monitors are not part of NPI. FlexArm and FlexMove configurations (FD20 XY-fa and XY-fm) shall not impact compliance with hygiene standard (DIN-1946-4, norm Raum Klasse 1A) in operating rooms having a Laminar Airflow. For FlexArm configurations flow-optimized rail covers shall be optionally available. The user shall be able to position the MCS frame(s) to have a clear view on the monitors from the workspot. The user shall be able to move the beam longitudinal and/or transversal (manual and motorized). To select a clinical area for acquisition as an alternative to positioning the tabletop. For standby parking (typically transversal only) during a surgical procedure. For normal parking when the stand is not used. Enables the user to rotate the frontal stand for position at the side or head of the table for optimized patient imaging. Iso-center stand rotation around a vertical axis to accomplish improved access to the patient. Function does not apply to the Poly-G floor stand.
UNS.SmartSuite.Safe3DScan | Be sure that 3D scan can be done without any collisions. | SRS.Allura.Func.ClassicDRA | The user shall define the end- and start-position of the DRA scan, implicitly executing a test-scan to detect obstructions. The system shall be able to perform the DRA scan at the side or head of the table (C-arm roll and propellor respectively). DRA acquisition is only possible for the frontal channel. The user shall be able to acquire a second run to enable subtracted reviewing. Settings, including start- and end-positions, shall be defined in EPX, but if allowed in EPX, the user shall be able to choose another end- and/or start-position. DRA requires a license.
UNS.SmartSuite.SafeSystemUse | Use a system that is safe to use and complies with the applicable international standards. Prevent harm to patients or hospital by: Using a system that provides protection against unacceptable safety risks in hazardous situations (for example, protection against X-ray, high temperature); Receiving guidance and instructions on minimizing the risk of causing harm in hazardous situations. New, N.A. | SRS.Allura.SLS.movement-integrity ¦ SRS.Allura.SLS.table-movement-integrity ¦ SRS.Allura.SLS.acoustic-noise ¦ SRS.Allura.SLS.maintain-minimal-functionality ¦ SRS.Allura.SLS.patient-contamination ¦ SRS.Allura.Func.RemoteServiceActiveIndication ¦ SRS.Allura.SLS.fse-contamin ¦ SRS.Allura.SLS.electr-shock-prev-staff-patient ¦ SRS.Allura.SLS.intended-image-quality ¦ SRS.Allura.SLS.DoD ¦ SRS.Allura.SLS.structural-integrity ¦ SRS.Allura.SLS.collision-unintend-ui ¦ SRS.Allura.Phys.MainsPower ¦ SRS.Allura.SLS.movements-failure-detection-and-fault-tolerant ¦ SRS.Allura.SLS.safe-surfaces ¦ SRS.Allura.Phys.ClimaticEnvironmentalConditions ¦ SRS.Allura.SLS.replay-path-free ¦ SRS.Allura.SLS.patient-contamin-airborne ¦ SRS.Allura.SLS.collision-harm-reduction ¦ SRS.Allura.SLS.patient-move-collision ¦ SRS.Allura.Phys.ElectromagneticCompatibility(EMC) ¦ SRS.Allura.SLS.unexpected-manual-behavior ¦ SRS.Allura.Phys.MechanicalEnvironmentalConditions ¦ SRS.Allura.SLS.injector-safety ¦ SRS.Allura.SLS.radiation-patient-dose ¦ SRS.Allura.SLS.collision-avoidance ¦ SRS.Allura.SLS.prevent-using-incorrect-in-patient-measurements ¦ SRS.Allura.SLS.image-latency ¦ SRS.Allura.Func.EnableRemoteServiceAccess ¦ SRS.Allura.SLS.patient-move-fall ¦ SRS.Allura.Func.SwitchSystemPowerOn ¦ SRS.Allura.Serv.RemoteInterference ¦ SRS.Allura.SLS.careless-manual-operation ¦ SRS.Allura.SLS.fse-safety ¦ SRS.Allura.SLS.peripheral-equipment ¦ SRS.Allura.SLS.consistent-image-orientation ¦ SRS.Allura.SLS.consistent-patient-and-image-data ¦ SRS.Allura.SLS.electr-shock-prevention-fse ¦ SRS.Allura.SLS.prevent-lih-live-mixup ¦ SRS.Allura.SLS.thermal-safety ¦ SRS.Allura.SLS.unintend-ui-patient-falling ¦ SRS.Allura.SLS.ListOfApplicableStandards ¦ SRS.Allura.SLS.careless-operator-patient-fall ¦ SRS.Allura.SLS.xray-imaging-failure-detection-and-fault-tolerant ¦ SRS.Allura.SLS.correct-overlay ¦ SRS.Allura.SLS.radiation-user-dose ¦ SRS.Allura.SLS.tipping-equipment | Out of control movements (e.g. spontaneous movements, movements with high speed than requested) shall be prevented in all conditions of the system (during start-up, normal use, power failure, power off). The table shall maintain its structural integrity during patient transfer and examination in order to avoid the patient falling from the table. The system shall not produce noise that can lead to hearing impairment (i.e. more than 80 dB(A)). When the system becomes unavailable for the intended treatment or diagnosis, minimal functionality is maintained as long as possible to terminate the procedure and release the patient from the system. The system shall be designed for: Cleaning and disinfection of system parts in the patient environment. Creating an aseptic environment. When a remote service session is active, this is indicated to the user. Measures shall be taken to guide a field service engineer in working with a contaminated system. Measures shall be taken to prevent electric shock of persons via accessible live parts or tube flash over. Measures shall be taken to avoid insufficient image quality during clinical procedures. The system fulfils the requirements for certification according to the Department of Defense (DoD) Risk Management Framework. SYSTEM INTEGRATION AND VERIFICATION REQUIREMENTS This chapter describes requirements to perform the system integration and system verification activities (I&V). The system shall be designed for testability for the manual and automated system test activities. With automated is meant that no human interaction is required and/or no attendance is required. This approach is called in this document “AiT” meaning Automations in System Test. Structural integrity of parts and suspensions that constitute a hazard when they fall shall be safeguarded. Measures shall be taken to prevent unintended motorized movements related to user interface design and user interaction with the system. The system shall be compatible with TN (including multi-source), TN-C, TN-S, TN-C-S symmetric 3-phase mains power configurations in the voltage range 380-480V nominal with a fundamental frequency of 50 or 60Hz. Compatibility with other standard power configurations shall be achieved via external voltage conversion transformers or compatible power inverters. Note that the system is not compatible with TT-, IT- and asymmetric grounded power systems. These mains power configurations shall be converted to one of the above mains power configurations with an isolation transformer. Environmental Conditions Measures shall be taken to detect errors in the control of motorized movements. When errors are detected, that not directly result in a hazardous situation: The system shall allow movements (possibly with reduced performance). The user is informed. The system parts that are designed to come into direct contact with the patient shall be made of material that are known to be bio-compatible. Following are the expected climatic conditions to which systems classified for operation in "Indoor, non-regulated room" will be exposed. The system shall not be damaged and remain safe after exposure to these conditions, while in operation, transport or storage. Table: Conditions representative for: Description of requirement Environmental Conditions Operation Temperature (Performance) +10°C/+30°C Operation Temperature (Safety) +10°C/+35°C Operation Maximum Temperature for unit in rack Ambient T + 5-15 °C (performance; maximum 45 °C) Operation Temperature change rate ≤ 0.5 °C/min Operation Relative Humidity (Performance) 20% - 80% Operation Relative Humidity (Safety) 20% - 93% Operation Condensation No Transport & Storage Temperature Constant: -25°C (3 days) +70°C (4 days) Cycles 25°C/ 97% RH - 40°C/ 93% RH (6 cycles) -10°C - 45°C (1 cycle) Transport & Storage Relative Humidity 5% - 95% Transport & Storage Condensation Possible1) Operation, Transport & Storage Air Pressure 700 - 1100 hPa Note: 1) The packaging during transport and storage protects PC cabinets against moisture from condensation. Directly after exposure to transport and storage conditions, the exposed part or system does not need to be in normal operational mode, but must remain safe. It must return to normal operational mode after acclimatization, once the climatic conditions to which it is exposed are within the operational limits. When the system is executing a programmed movement (pre-programmed or defined by the user) measures shall be taken so that no collision occur. The system shall be designed to minimize its contribution to airborne contaminants reaching the patient. In cases where motorized movements (despite preventive measures) do lead to collisions of mechanical parts with a person, the collision forces shall be limited to prevent harm. Measures shall be taken to prevent that a movement of the patient results in a motorized or manual movement or results in a collision (i.e. entrapment) of the patient. According to Standard IEC 60601-1-2. Measures shall be taken to prevent unintended and/or uncontrolled manual movement behavior. There shall be no impact on system operation and safety after exposure to repetitive bumps and vibration, that are comparable to the expected mechanical conditions during transport. Only injectors that are compatible with the system and its environment are to be used. Contrast injection shall be single fault safe and 'dead-man' controlled by the user. Contrast injection without X-ray imaging shall be avoided. The system shall inform the user during and after the examination with patient dose information, to decide on continuation of the treatment. The system shall give a user warning above a (configurable) threshold. During motorized movements, the system shall help the operator to avoid collisions with a person. Measures shall be taken to prevent incorrect results of image based measurements and to warn the user when the measurement results are possibly incorrect. To enable procedures which require hand-eye coordination, the perceived image latency shall not exceed 230 ms. With this function the user explicitly enables and grants access for remote service activities. Dependent on the type of remote service activity the user will be restricted in using the system. Measures shall be provided to prevent that a patient repositioning on the table results in the patient falling from the table. The user shall be able to switch the system on which includes system self-tests and service diagnostics. Remote access shall be explicitly accepted by the local user unless no interference to the local user occurs. The local user shall be able to terminate remote access (e.g. to start an emergency procedure). Remote Monitoring & Alerting During manual movements the system shall help the operator to prevent collisions with a person. Measures shall be taken to guide a field service engineer in (de)installing and maintaining the system in a safe way. Privacy And Security Measures shall be taken to guarantee safe mechanical, electrical and functional compatibility of peripheral equipment and the system. The patient orientation in the X-ray image shall be consistent with the patient to table position and other information as provided by the user. All patient and image data shall be displayed and maintained in a consistent manner throughout the system. Manual changes in patient related data shall only be done after explicit confirmation of the operator. Measures shall be taken to train the field service engineer not to work with live parts, but switch the system off. The system shall prevent the user from mistaking a still or Last Image Hold image for a live image. The system shall not have hot accessible surfaces which can lead to thermal injuries (burns). Measures shall be taken to prevent unintended table movements related to user interface design and user interaction with the system. The standards as mentioned in chapter APPLICABLE STANDARDS are applicable. Environmental Care The system shall provide means to prevent that a user initiated table movement, may lead to the patient falling from the table, and it shall be made clear how to set-up the system for patient transfer. Measures shall be taken to detect errors in X-ray imaging functionality. When errors are detected that not directly result in loss of imaging functionality: The system shall allow imaging (possibly with reduced performance). The user is informed. Measures shall be taken to avoid mismatches between overlay images and the actual image. Measures shall be taken to minimize the exposure of the user to X-ray during operation of the system. Measures shall be taken to prevent movable and static parts of the system from tipping over.
UNS.SmartSuite.SecurityProtection | Protect the patient privacy as well as to prevent loss of information or inability to perform clinical procedures, as a result of security breaches; Provide protection against all known security threats; Receive guidance and instructions how to protect the system against security threats and how to minimize the vulnerability of the system; Create and manage user accounts. New, Have protection over patient and personal data across the network by using user authentication and encryption protocols; Comply with the industry security and privacy standards. | SRS.Security.HN.RemovableMediaBreached ¦ SRS.Allura.Func.RemoteServiceActiveIndication ¦ SRS.Allura.SLS.DoD ¦ SRS.Security.RM.AttackViaMedicalDevice ¦ SRS.Security.MD.DataBreach ¦ SRS.Security.MD.AttackViaPhysical DeviceModification ¦ SRS.Security.MD.HospitalNetwork ¦ SRS.Security.PHD.PatientDataLoss ¦ SRS.Security.MD.RemoteAccessPoint ¦ SRS.Security.PHD.DataBreach ¦ SRS.Security.PHD.PhysicalAccess ¦ SRS.Security.PHD.MedicalDevice ¦ SRS.Security.PHD.PatientDataCorrupted ¦ SRS.Allura.Serv.AccessRightsAndAuthorization ¦ SRS.Security.MD.PhysicalAccess ¦ SRS.Security.MD.ReputationDamage ¦ SRS.Security.MD.MalwareOnRemovable media ¦ SRS.Security.PHD.HospitalNetwork ¦ SRS.Allura.Func.EnableRemoteServiceAccess ¦ SRS.Security.PHD.PhysicalDismounting ¦ SRS.Security.HN.HospitalNetworkBreached ¦ SRS.Allura.Serv.RemoteInterference ¦ SRS.Security.PHD.RemoteAccessPoint ¦ SRS.Security.MD.UnavailabilityOfMedicalDevice ¦ SRS.Security.HN.AttackViaMedicalDevice | The Azurion shall reduce the impact of: Breached Devices, Disclosed Patient Health data, Unavailability of data, in case Removable Media is breached through the Azurion. Standards, Legal Requirements International Standards When a remote service session is active, this is indicated to the user. The system fulfils the requirements for certification according to the Department of Defense (DoD) Risk Management Framework. SYSTEM INTEGRATION AND VERIFICATION REQUIREMENTS This chapter describes requirements to perform the system integration and system verification activities (I&V). The system shall be designed for testability for the manual and automated system test activities. With automated is meant that no human interaction is required and/or no attendance is required. This approach is called in this document “AiT” meaning Automations in System Test. The Removable Media shall be protected against attacks via Azurion. The Azurion shall reduce the impact of Data Breach in case Azurion is breached. The Azurion shall be protected against attacks via the hospital network. The Azurion shall reduce the impact of Reputation damage (manufacturer/hospital) in case Azurion is breached. The Azurion shall be protected against attacks via a remote access point. The Azurion shall reduce the impact of Reputation damage (manufacturer/hospital) in case Azurion is breached. The Patient Health Data shall be protected against attacks via physical access. The Patient Health Data shall be protected against attacks via the Azurion. The Azurion shall reduce the impact of Patient data corrupted in case of Patient Health Data is attacked. The user's authorization level shall determine the available service functionality, e.g. without a key only changes with no associated risks are allowed. The Azurion shall be protected against attacks via physical access. The Azurion shall reduce the impact of Reputation damage (manufacturer/hospital) in case Azurion is breached. The Azurion shall be protected against attacks via malware on removable media. The Patient Health Data shall be protected against attacks via hospital network. With this function the user explicitly enables and grants access for remote service activities. Dependent on the type of remote service activity the user will be restricted in using the system. The Patient Health Data shall be protected against attacks via physical dismounting. The Azurion shall reduce the impact of: Unavailability of Hospital Network, Breached hospital infrastructure, Disclosed customer data, in case Hospital Network is attacked through the Azurion. Remote access shall be explicitly accepted by the local user unless no interference to the local user occurs. The local user shall be able to terminate remote access (e.g. to start an emergency procedure). Remote Monitoring & Alerting The Patient Health Data shall be protected against attacks via remote access point. The Azurion shall reduce the impact of Unavailability of Azurion in case Azurion is breached. The Hospital Network shall be protected against attacks via Azurion.
UNS.SmartSuite.SettingtoWork | Perform the main tasks* that are needed to make sure that the system is operational according to specifications in the image-guided lab. *power on for the first time, configuration, adjustments. | SRS.Allura.Qual.FSResponseTimes ¦ SRS.Allura.UI.ServiceUILanguage ¦ SRS.Allura.Serv.UserInterface ¦ SRS.Allura.SLS.movements-failure-detection-and-fault-tolerant ¦ SRS.Allura.Func.TubeLoadProtectionAndIndication ¦ SRS.Allura.Serv.Configure ¦ SRS.Allura.Func.SystemRestart ¦ SRS.Allura.Func.DisplayServiceGuidance ¦ SRS.Allura.SLS.maintain-minimal-functionality ¦ SRS.Allura.Qual.ClinicalResponseTimes.System ¦ SRS.Allura.Serv.Adjust ¦ SRS.Allura.Func.GracefulDegradation ¦ SRS.Allura.Func.ExceptionHandling ¦ SRS.Allura.Func.DisplayUserMessages ¦ SRS.Allura.SLS.xray-imaging-failure-detection-and-fault-tolerant ¦ SRS.Allura.Qual.IQFactory/ServiceMeasurements ¦ SRS.Allura.Serv.Customize | Table: Categories/ items Max. notes Field-Service Tool sec Local start-up service framework: - from start-up to log-on screen 15 - from log-on screen to display framework 10 after entering service key password - start-up with already logged on service user 20 Remote start-up service framework 100 Local shut-down service framework: - from close to service conclusion screen 6 after removing service key or after closing the close confirmation screen - from service conclusion screen to shut-down 3 EPX Management sec Start-up EPX Management Tool 30 includes loading of database settings in editor Store modified settings for verification 30 up to modified settings active for try-out Batch verification (one modified EPX-Application/Procedure) 30 +10 sec per additional modified EPX-Application/Procedure Store modified changes to system 120 up to usable clinical system Calibration min Calibrate FD20 90 Calibrate FD15 90 Calibrate FD12 90 Frontal Stand calibration 340 calibrate geometry without X-ray Lateral Stand calibration 225 calibrate geometry without X-ray Table calibration 75 calibrate geometry without X-ray Generator calibration 90 Image transfer sec Image max. 5122 - 12 bit 150 over 40 Kbps link Image Quality This section describes the basic Image Quality requirements. The purpose is to describe the system in terms of Image Quality requirements such, that if a system meets the requirements in this specification, the image quality of that system is guaranteed (and acceptable for the end user). The service UI for service engineers is in English. QUALITY AND PERFORMANCE REQUIREMENTS In this chapter the requirements are given for the quality and performance requirements. X-ray Performance Figures The service UI shall provide access to the whole system. The service UI shall be based on a standardized service framework and the EPX management tool. The system shall provide on-line help information. Measures shall be taken to detect errors in the control of motorized movements. When errors are detected, that not directly result in a hazardous situation: The system shall allow movements (possibly with reduced performance). The user is informed. The system shall safeguard X-ray-generation components and indirectly protect the user and the patient against potential harmful consequences from overloading the tube. Low-load fluoro shall remain available. Configuration functions enable the manual selection of basic settings related to: The system itself (e.g. licensing file, connected peripherals, display topology, remote service). The departmental/hospital environment (e.g. printers, archive, worklist management systems). The user shall be able to restart the system to recover faster from a system failure compared to subsequently switching off and on. Error events initiate a service guidance. The service guidance is logged via the logging capabilities provided by the system. Additionally the error related service guidance is reported to the user. When the system becomes unavailable for the intended treatment or diagnosis, minimal functionality is maintained as long as possible to terminate the procedure and release the patient from the system. Table: Categories/ items Max. notes: (all values in seconds) Regular System startup (power-up) 300 Time until functionality required for the intended use is available. System startup after total power down 300 (China only) 360 (Rest of the world) Measured after external power supply restored to the system. Warm restart 90 Cold restart 360 Full cycle from “off” button press to full functionality. Recovery geometry after emergency stop 105 Time until all geometry movements available. Build-in correction measures shall provide adjustment functionality to bring the system within specified limits. The service framework shall provide adjustment criteria and procedures to the FSE and hospital technician. In case of a system failure the available graceful degradation mode is entered, in which the system operates with reduced functionality and/or performance; the following rules apply: Via user/service guidance the graceful degradation situation is reported to the user. In a biplane system a defect in one channel must not prevent use of the other channel. The system will not try to automatically find alternate ways to realize the requested functionality. It is of the utmost importance that service functionality remains available for diagnosis. Demo System Behaviour A number of exceptional situations can be distinguished: A collision shall be indicated, also if other user messages are active. After a software error recovery, the system shall guarantee that previously stored clinical information is still available. If the user requests a function that cannot be executed in the current system state, for a reason which is not obvious to the user, the system shall display an appropriate user message. If the user requests a function and the system cannot execute the function because the system has run out of resources (e.g. storage full), the system shall display an appropriate user message. Graceful Degradation Comprises user guidance or error messages, which is to be displayed at the relevant workspot(s). Static messages are displayed as long as a cause exists (e.g. a system error). Triggered messages disappear automatically after a predefined time period. Confirmation by the user is not required. Measures shall be taken to detect errors in X-ray imaging functionality. When errors are detected that not directly result in loss of imaging functionality: The system shall allow imaging (possibly with reduced performance). The user is informed. Requirements for installation and manufacturing quality, related to guarantee reproduction of IQ. All system-wide level 1 measurements shall be executed and passed as part of manufacturing and/or installation and/or preventive and/or corrective maintenance.Additionally, lower-level measurement tests shall be available to be executed for diagnostic and regulatory purposes. Table: What is measured Level 1 Comments Dose per image detector Dose Test and Dose Calibration Noise estimate in ROI on FD Dose Test and Dose Calibration Stability noise value in ROI on FD Fast Dose Verification kV-Stabilized Energy and Signal flow 1 see SRS.Allura.Qual.IQKVStabilized Video stabilization level Energy and Signal flow 1 see SRS.Allura.Qual.IQWorkingPoint mAs-Stabilized Test Shot Energy and Signal flow 2 see SRS.Allura.Qual.IQmAsStabilized mAs-Stabilized exposure Energy and Signal flow 2 see SRS.Allura.Qual.IQmAsStabilized Calculated power to validate EDL function EDL verification Regulatory requirement described by FDA: 21CFR1020.32 (clause(d)). Contrast Monitor Performance see SRS.Allura.Qual.IQDisplaySignalTransfer Vignetting Monitor Performance see SRS.Allura.Qual.IQBrightnessUniformity Sharpness Monitor Performance see SRS.Allura.Qual.IQDisplaySignalTransfer Low contrast check Monitor Performance see SRS.Allura.Qual.IQDisplaySignalTransfer What is measured Lower Level DAP difference AEP meter and dose probe Skin dose difference AK displayed and dose probe DAP Measurement and AirKerma Verification Regulatory requirements described by IEC60601-2-43 (clauses:203.6.4.5 Dosimetric indications) and 21CFR1020.32 (clause (k)(6)) Test to check alignment of x-ray beam with detector field of view Field limitation Regulatory requirements described by IEC 60601-1-3 (clause 8.5.3), FDA 21CFR 1020.32 (clause (b)(2) etc), IEC 60601-2-43 (clause 203.8.5.3). Test for verifying x-ray beam vs detector entrance plane Xray Beam alignment see SRS.Allura.Qual.IQBeamAlignment & isolated part of Field Limitation test (without sagging and stand maximal application conditions). Horizontal focal spot Vertical focal spot Focal Spot size Focal spot size is conform IEC60336. Focal spot sizes are included in Performance Test Limits for fault finding. Small focus of x-ray tube Large focus of x-ray tube Limiting Resolution System Level SRS.Allura.Qual.IQLimitingResolutionSystemLevel Beam quality x-ray beam in mm Al Half Value Layer Regulatory requirements described by IEC 60601-1-3 (clause 7.1, half value layers in x-ray equipment). X-ray output x-ray tube X-ray Output SRS.Allura.Qual.IQX-rayOutput Check difference real kV and kV indication ∆kV for 6 measurements (5 manual exposure, 1 fluo technique) kV-Beam is kV-Desk Regulatory requirement described by IEC 61223-3-1 (clause 5.2.1, X-ray tube voltage). Qualitative image impression for Germany Monitor SMPTE Regulatory requirement described by German DIN6868-157:2014-11 What is measured Other Air Kerma verification Dose measurement for verification of the Entrance Doserate Limitation (EDL) mechanism Regulatory requirement described by FDA: 21CFR1020.32 (clause(d)). PHYSICAL REQUIREMENTS Mechanical No specific requirements apart from targets as derived from requirements stated in section 12.3 Standards, Legal Requirements. Customization of EPX apply to the optimization of settings to local hospital and/or user preferences: Customization of EPX settings shall be available through the EPX tool, integrated in the service UI. Permission to modify certain parameters shall depend on the authorization level of a particular FSE. It shall be possible to check all applicable EPX applications and procedures, or immediately after a modification is made to an EPX application and procedure. The FSE shall be able to store the modified values or revert to latest verified and applied settings. Verification & Report The scope is to verify and report that the system is within its intended specification (pass/fail).
UNS.SmartSuite.ShutDown | Shut-down the system. New, N.A. | SRS.Allura.Qual.ClinicalResponseTimes.System ¦ SRS.Allura.Func.SwitchSystemPowerOff | Table: Categories/ items Max. notes: (all values in seconds) Regular System startup (power-up) 300 Time until functionality required for the intended use is available. System startup after total power down 300 (China only) 360 (Rest of the world) Measured after external power supply restored to the system. Warm restart 90 Cold restart 360 Full cycle from “off” button press to full functionality. Recovery geometry after emergency stop 105 Time until all geometry movements available. The user shall be able to switch the system off; data is preserved to ensure acquired data does not get lost. Power outlets for external systems/equipment shall stay powered, but the user shall be able to switch off the power to these power outlets.
UNS.SmartSuite.SitePlanning | Obtain all information that is needed to prepare for the installation/upgrade. | SRS.Allura.Serv.SingleLabConnection | The system shall provide one entry point to the lab to service compatible room specific auxiliary systems, i.e. logging retrieval from and remote desktop of other configured systems shall be supported. The connected auxiliary systems shall be accessible even if the X-ray system is switched off. The remote & local FSE shall be able to troubleshoot connectivity between the entry point and other connected systems; and between the entry point and RSN. Miscellaneous
UNS.SmartSuite.SiteReadinessCheck | Obtain information that is needed to ensure that the prerequisites of system installation are met. | SRS.Allura.Serv.Maintenance.Planned | The system shall enable the performance of planned maintenance activities to ensure its functionality during its lifetime.
UNS.SmartSuite.SmallRoom | Fit FlexArm in smaller rooms (35m² and more); Calculate the ROAR by having Imaging Reach of the System data available. New, N.A. | SRS.Allura.Conf.MinimalRoomSize | The systems as mentioned in SRS.Allura.Conf.XperFDSeriesConfigurations in their minimum configuration shall fit in a room with the following minimum sizes: Table: System configuration (for all FD configurations) Minimum room size notes Fl 5.8 m length x 4.1 m width = 24 m2 Ce / NB 6.1 m length x 4.1 m width = 25 m2 Bpl 6.5 m length x 4.1 m width = 27 m2 XY-fa 7.1 m length x 4.9 m width = 35 m2 XY-fm 7.8 m length x 6.2 m width = 48 m2
UNS.SmartSuite.SmoothInteraction | Move the system in such a way that the acceleration and de-acceleration, the speed and the logic of automated movements is perceived by the user as being effective and in a controlled manner (not too fast, not too slow); Move the system with minimal* user interaction. *) minimal is with the least possible amount of user interaction. New, N.A. | SRS.Allura.SLS.movement-integrity ¦ SRS.Allura.Func.SelectTableSide ¦ SRS.Allura.SLS.unexpected-manual-behavior ¦ SRS.Allura.Func.PriorityConstraints | Out of control movements (e.g. spontaneous movements, movements with high speed than requested) shall be prevented in all conditions of the system (during start-up, normal use, power failure, power off). The user shall be able in table-oriented mode to maintain intuitive movement control from four table sides. The user shall be able to select the table side on every movement-control UI module individually. Combined Patient- And Beam Positioning This section focuses on movements under direct user control, which are not stand or table specific. Measures shall be taken to prevent unintended and/or uncontrolled manual movement behavior. In case of conflicting or ambiguous user requests the following constraints shall apply: Overruling priority (high to low): 1) exposure, 2) fluoro, 3) other functions (except movement requests and concurrent working in control room) that are applied to the current acquisition study. A function shall be executed as long as requested from at least one UI, with exception of motorized movement requests: parallel movement requests shall be served as long as they are not conflicting. External (surgery) tables are equipped with their own user interface modules. For Maquet Magnus, parallel table movement requests shall result in stopping of the motorized table movements. A movement request via a TruSystem UI will stop a table movement requested via the system UI, without starting the newly requested movement. A table request via the system UI is ignored as long as a table movement request via the TruSystem is active. Exception Handling
UNS.SmartSuite.SoftwareInstallation | Install software on the system that match the workflow prerequisites*. *hardware and software compatibility. | SRS.Allura.Serv.RemoteConnectivityStatus ¦ SRS.Allura.Serv.Upgrading ¦ SRS.Allura.Serv.InstallSoftware | The system shall inform the clinical user about the remote connection status The system shall provide a remote connection test for clinical user and hospital IT. For VPN connected systems it shall be possible to up- and download remote configuration settings from/to the system. The following data shall be retained during an upgrade or can be restored from a previously stored back-up: Configuration and calibration data EPX (except manually added image processing tastes) and customization data Maintenance Related The initial installation of software and a system upgrade shall adhere to the following. The user shall be able to perform the installation/upgrade from a single entry point. Software options shall be pre-installed and activated through a standard licensing mechanism (sw-license file). User interaction shall only be required in the first 10 minutes of starting the installation/upgrade (this excludes configuration of the system).
UNS.SmartSuite.SoundsRoom | Have a low level of noise in the room (no disturbing noise in the room, except when immediately attention required, like heart stops). | SRS.Allura.SLS.acoustic-noise ¦ SRS.Allura.Phys.Noise | The system shall not produce noise that can lead to hearing impairment (i.e. more than 80 dB(A)). Acoustical noise level for different areas are listed below: Max. 60 dB(A) in exam room, with the following exceptions: Max. 75 db(A) in exam room during ≥ 30 °/sec DRA scan; Max. 71 dB(A) in exam room during all other geometry movements. Max. 60 dB(A) in control room. Max. 65 dB(A) in technical room.
UNS.SmartSuite.StartExamination | Start the examination from the Control Room; Start the examination from the Exam Room from the Additional FlexSpot. New, N.A. | SRS.Allura.Func.SelectStudyFromScheduledList ¦ SRS.Allura.Qual.ClinicalResponseTimes.Admin ¦ SRS.Allura.Func.ReopenFinishedStudy ¦ SRS.Allura.Intgr.Flexspot.AdditionalWorkspot ¦ SRS.Allura.Func.SelectStudyFromSuspendedList | This enables the user to select a study/procedure step for acquisition from the scheduled study list. The patient name is displayed clearly in the exam room. The system is prepared for acquisition; information about any blockage is shown. The system acquisition parameters are set according to the current EPX application/procedure. For WLM based procedure steps, progress is reported according DICOM MPPS. Table: Categories/ items Max. notes: (all values in seconds) General response on UI requests e.g. switch exposure channel mono/biplane switch roadmap-on multiphase frame-rate subtraction on/off 1.0 one action; e.g. display a list, menu. Delete study 1.0 Select study for acquisition 5.0 Select study for viewing 2.8 Select EPX application for current study 4.0 Select EPX procedure for current study 1.2 in idle state, not valid when EPX is selected during an active fluoroscopy run The user is able to re-open a finished study/procedure step and select it for acquisition. The study/procedure step is re-initialized in such way that a complete workflow scenario may again be applied on the study/procedure step. Actions are similar compared to selection from the scheduled list, except: When a finished study/procedure step is selected, new images acquired are appended to the study/procedure step data. The procedure card, clinical-file structure, dose report data and other relevant data already attached to this study/procedure step are used as a starting point for any additional X-ray acquisitions. Finishing A Study A second workspot (keyboard, mouse, single monitor) next to FlexSpot shall be offered, which can be a workspot in the control room or exam room. One video source can be selected to be shown. Video can originate from the system itself and from auxiliary systems (configurable set, e.g. Interventional Workspot, IVUS). The additional FlexSpot requires a license. This function lets the user select a study/procedure step for acquisition from the suspended study list. Actions are similar compared to selection from the scheduled list, except: The last used settings are used to continue acquisition (e.g. dose data, clinical files). Newly acquired images are appended to the study/procedure step data.
UNS.SmartSuite.StartUp | Start-up the whole system; Enable/disable X-ray to prevent accidental exposures (e.g. during cleaning, patient transfer etc.). Turn on the displays only in FlexSpot configuration, FlexVision (Pro) and Switchable Monitors (i.e. to access the non-X-ray sources). New, N.A. | SRS.Allura.Phys.MainsPower ¦ SRS.Allura.Func.X-RayDisable/Enable ¦ SRS.Allura.Qual.ClinicalResponseTimes.System ¦ SRS.Allura.Func.SwitchSystemPowerOn.VideoOnly ¦ SRS.Allura.Func.SwitchSystemPowerOn | The system shall be compatible with TN (including multi-source), TN-C, TN-S, TN-C-S symmetric 3-phase mains power configurations in the voltage range 380-480V nominal with a fundamental frequency of 50 or 60Hz. Compatibility with other standard power configurations shall be achieved via external voltage conversion transformers or compatible power inverters. Note that the system is not compatible with TT-, IT- and asymmetric grounded power systems. These mains power configurations shall be converted to one of the above mains power configurations with an isolation transformer. Environmental Conditions A disable/enable X-ray function is provided to avoid accidental activation of X-ray related functions such as radiation, coupled contrast injection and programmed movements by laymen (e.g. cleaners). The disable function cannot be activated during fluoro and exposure active. Default state after power-up is configurable. Table: Categories/ items Max. notes: (all values in seconds) Regular System startup (power-up) 300 Time until functionality required for the intended use is available. System startup after total power down 300 (China only) 360 (Rest of the world) Measured after external power supply restored to the system. Warm restart 90 Cold restart 360 Full cycle from “off” button press to full functionality. Recovery geometry after emergency stop 105 Time until all geometry movements available. On Systems with FlexSpot(s) and/or FlexVision the user shall be able to switch on only the video part of the system, meaning the monitors and the control for FlexSpot(s), FlexVision, Switchable monitors (including Multivision) and (Additional, 3rd or 4th) Multiview including TSM. The user shall be able to switch the system on which includes system self-tests and service diagnostics.
UNS.SmartSuite.Sterility | Cover system parts with sterile covers in the immediate vicinity of patients. Have a draping solution for TSM with regard to: Sterility: positioning and control over sterile sections Visibility: transparent without reflections Responsiveness: response rate using (dirty) gloves comparable to TSM without draping. New, N.A. | SRS.Allura.UI.UIModule ¦ SRS.Allura.UI.TSM.Workspot ¦ SRS.Allura.Conf.Stand-RelatedItems | The review and control modules shall support backlighting to allow easy visibility of buttons in e.g. dark examination rooms and give user guidance. The UI modules at or near the table shall allow control under sterile conditions (e.g. under a sterile cover). The user shall be able to detach the UI modules from the pedestal resp. table and attach them to the table resp. pedestal. The TSM shall offer a workspot in examination and/or control room to control system internal applications but also applications running on other systems. The touch screen shall be operable under table side conditions i.e. wearing gloves, covered with a sterile drape and spilled with fluids. The user shall be able to use an application on a certain TSM, while another user shall be able to use any other connected application on another TSM. The internal state of an application that is shown on multiple TSMs shall be synchronized. Table: sub Poly-G Poly-G Poly-G Clea Clea Clea Clea L-arc tag Item Ce XY fm Fl Ce XY fm XY fa Fl Bpl notes .std -1 Extended rail o o o o o o FlexArm: 4.3m; standard: 6.0m or 7.8m. -2 L-arm long # # # # -3 L-arm short # # # # -4 Flexarm long # -5 Flexarm short # -6 Ceiling rail covers o See SRS.Allura.Phys.LaminarAirflowCompatibility -7 Optical imaging system o o o see SRS.Allura.Intgr.OpticalImagingSystem -8 Coverbags c c c c c c c c Arc / Tube / Detector / TSM
UNS.SmartSuite.SystemNoise | Have a low level of noise in the room of the system when in use (no disturbing system noise when moving the system, except when immediately attention required). New, N.A. | SRS.Allura.Phys.Noise | Acoustical noise level for different areas are listed below: Max. 60 dB(A) in exam room, with the following exceptions: Max. 75 db(A) in exam room during ≥ 30 °/sec DRA scan; Max. 71 dB(A) in exam room during all other geometry movements. Max. 60 dB(A) in control room. Max. 65 dB(A) in technical room.
UNS.SmartSuite.TSMNavigation | Switch between modalities displayed on the TSM Pro within a maximum of 2 user interactions. New, N.A. | SRS.Allura.UI.TSM.Workspot ¦ SRS.Allura.Intgr.Interfaces.ControlAuxiliarySystems | The TSM shall offer a workspot in examination and/or control room to control system internal applications but also applications running on other systems. The touch screen shall be operable under table side conditions i.e. wearing gloves, covered with a sterile drape and spilled with fluids. The user shall be able to use an application on a certain TSM, while another user shall be able to use any other connected application on another TSM. The internal state of an application that is shown on multiple TSMs shall be synchronized. The system shall provide communication interfaces to control compatible auxiliary systems from the systems user interface with wireless mouse on the auxiliary system’s graphical UI displayed on FlexVision with mouse and keyboard on the auxiliary system’s graphical UI displayed on Flexspot from TSM by a generic button based UI tailored by the auxiliary system or by a graphical UI streamed from the auxiliary system including (multi-)touch control. This requires a license control shall be established automatically when (mobile) compatible auxiliary system is connected
UNS.SmartSuite.TSMPlugandPlay | Automatically display 3rd party modality information on the TSM when the 3rd party is connected to the network. New, N.A. | SRS.Allura.Intgr.Interfaces.ControlAuxiliarySystems | The system shall provide communication interfaces to control compatible auxiliary systems from the systems user interface with wireless mouse on the auxiliary system’s graphical UI displayed on FlexVision with mouse and keyboard on the auxiliary system’s graphical UI displayed on Flexspot from TSM by a generic button based UI tailored by the auxiliary system or by a graphical UI streamed from the auxiliary system including (multi-)touch control. This requires a license control shall be established automatically when (mobile) compatible auxiliary system is connected
UNS.SmartSuite.TSMtouchpad | Use the TSM as touchpad/ keyboard for manipulating/ navigating the FlexVision (Pro) at table side, also under a sterile bag. New, Perform new interactions at tableside when compared to previous generations of Allura systems such as: See live clinical images in 2D; Navigate images in 2D; Use mouse pointer which is visible in all screens showing X-ray images, to improve communication in ER/CR; Control the shutters and wedges from TSM. | SRS.Allura.Intgr.Flexvision.KeyboardMouse | With FlexVision, the system shall allow for remote control of auxiliary systems via on-screen keyboard and touchpad (on TSM) and / or wireless mouse. This requires a license.
UNS.SmartSuite.Table | Position the table top in any orientation; Move the table top to position the region of interest in the centre of the X-ray beam; Move the table top height (up/down) to optimize the working height for the physician; Move the table (with pivot or swivel) to be able to reach the entire leg area of a patient during an examination; Image the patient from head to feet without moving the patient (full body coverage); Perform an iso-centric tilt bringing the patient's head or feet down e.g. to use earth gravity for contrast fluid positioning (not applicable to FlexArm with Hillrom table); Perform a cradle tilt to allow the surgeon easy access during surgery or for patient positioning purposes e.g. bring the patient into a special position required for CO2 usage, biopsies of kidney or to slide the patient off the table; Pivot the table to move the arm of the patient into the X-ray beam for examination of the patient's arm or to bring the patient on/off the table; Pivot the table 180 degrees; Avoid accidental activation of movement related functions of table and available stands. New, N.A. | SRS.Allura.Func.MovementDisable/Enable ¦ SRS.Allura.Func.TableCradle ¦ SRS.Allura.Func.TableTilting ¦ SRS.Allura.Func.ChangingTableHeight ¦ SRS.Allura.Func.SwivelMovementTableBase ¦ SRS.Allura.Func.PositioningTableTop ¦ SRS.Allura.Conf.Table-RelatedItems ¦ SRS.Allura.Func.PivotRotationTableBase | A disable movement function is provided to avoid accidental activation of movement related functions of table and available stands by laymen (e.g. cleaners). Programmed movements will not be available when movement is disabled. The disable function cannot be activated during programmed movements. Default state after power-up is enabled. The table rotates on the longitudinal axis (motorized). The cradle rotation center is ISO-centric within the limits set by the particular table type (Maquet/TruSystem OR Tables are excluded) The cradle movement automatically stops in horizontal position. This enables the user to tilt the tabletop to lift or lower the patient's head relative to his feet (motorized). The tilt rotation center is isocentric (except for TruSystem table) within the limits set by the particular table type. Customizable, the stand can follow the tilt movement in a synchronous way (except for TruSystem table). The tilting movement automatically stops in horizontal position. Isocentric tilt on the Magnus table requires a license. Moves the table up/down (motorized). The movement range depends on the table type and on the table tilt/cradle angle. Provides additional longitudinal span to bring the table closer to the stand (AD7-X(N)T only). The table base can be positioned in one of two places (movement is motorized). This applies to motorized or manual positioning of the tabletop in longitudinal and/or lateral direction. Table-type dependent manual positioning (so called 'float') is balanced (passive or motor-assist) to prevent the patient from sliding down (e.g. in case of tilt). Table: sub AD7XT AD7XNT Maquet tag item TruSystem notes . ptf Features -1 neuro tabletop # # -2 cardio tabletop # # -3 universal tabletop % -4 carbonfiber & neuro tabletops % variants A0…A7 for Maquet variants Carbon Floatline, Carbon Spine, SQ14-Xtra and universal for TruSystem -5 tabletop brake option o o VA kit: when power off, friction brake is applied to prevent movement of tabletop -6 table pivot option o o + see SRS.Allura.Func.PivotRotationTableBase -7 table swivel option o o only for Clea-F, see SRS.Allura.Func.SwivelMovementTableBase -8 motorized -float o o + motorized longitudinal and lateral tabletop movements -9 table tilt option o + see SRS.Allura.Func.TableTilting -10 table cradle option o + see SRS.Allura.Func.TableCradle -13 floor spacer o o ‘retention plate’, as mitigation against spillage in OR .pta AD7XT / 7XNT accessories -1 table-mounted radiation shield o o -2 peripheral filters o o -3 arm supports o o -4 head support o o -5 patient straps + + Mandatory in combination with AD7XT and AD7NXT -6 rail accessory clamp o o -7 tabletop accessory clamp o o -8 mattress o o -9 table covers (non-sterile) o o in hybrid OR, table-base, for cleaning purposes -10 additional OP-rail o o extended positioning for UI modules -11 auxiliary OP-rail kit o o positioning for third party equipment -13 neuro wedge o o for patient posturing -14 table rail cable guide + + -15 ratchet compressor o o compression belts with Ratchet-winding mechanism for symmetrical compression -16 handgrip and clamp set o o table clamps and handgrips -17 neuro head holder c c -18 Leg support board o o For use in AngioCT This enables the user to rotate the table base either to free a patient or to expose off-center body parts (not motorized).
UNS.SmartSuite.TableConfiguration | Make use of the Allura/Azurion system in endovascular and hybrid procedures Context: For patient support at installation the user can choose between Angio Table and OR Table Magnus or Hillrom; Make use of the system in endovascular and hybrid procedures in one room and optimize for these kind of procedures; Use special patient accessories support that is required in open procedures with C-arc in imaging or in parking position. New, N.A. | SRS.Allura.Conf.Table-RelatedItems | Table: sub AD7XT AD7XNT Maquet tag item TruSystem notes . ptf Features -1 neuro tabletop # # -2 cardio tabletop # # -3 universal tabletop % -4 carbonfiber & neuro tabletops % variants A0…A7 for Maquet variants Carbon Floatline, Carbon Spine, SQ14-Xtra and universal for TruSystem -5 tabletop brake option o o VA kit: when power off, friction brake is applied to prevent movement of tabletop -6 table pivot option o o + see SRS.Allura.Func.PivotRotationTableBase -7 table swivel option o o only for Clea-F, see SRS.Allura.Func.SwivelMovementTableBase -8 motorized -float o o + motorized longitudinal and lateral tabletop movements -9 table tilt option o + see SRS.Allura.Func.TableTilting -10 table cradle option o + see SRS.Allura.Func.TableCradle -13 floor spacer o o ‘retention plate’, as mitigation against spillage in OR .pta AD7XT / 7XNT accessories -1 table-mounted radiation shield o o -2 peripheral filters o o -3 arm supports o o -4 head support o o -5 patient straps + + Mandatory in combination with AD7XT and AD7NXT -6 rail accessory clamp o o -7 tabletop accessory clamp o o -8 mattress o o -9 table covers (non-sterile) o o in hybrid OR, table-base, for cleaning purposes -10 additional OP-rail o o extended positioning for UI modules -11 auxiliary OP-rail kit o o positioning for third party equipment -13 neuro wedge o o for patient posturing -14 table rail cable guide + + -15 ratchet compressor o o compression belts with Ratchet-winding mechanism for symmetrical compression -16 handgrip and clamp set o o table clamps and handgrips -17 neuro head holder c c -18 Leg support board o o For use in AngioCT
UNS.SmartSuite.TableSetup | Position the table top in any orientation and height; Move the table top to position the region of interest in the center of the X-ray beam; Move the table top height (up/down) to optimize the working height for the physician; Move the table top to predefined positions. New, N.A. | SRS.Allura.Func.TableCradle ¦ SRS.Allura.Func.TableTilting ¦ SRS.Allura.Func.ChangingTableHeight ¦ SRS.Allura.Func.SwivelMovementTableBase ¦ SRS.Allura.Func.PositioningTableTop ¦ SRS.Allura.Qual.ClinicalResponseTimes.Movement ¦ SRS.Allura.Func.PivotRotationTableBase | The table rotates on the longitudinal axis (motorized). The cradle rotation center is ISO-centric within the limits set by the particular table type (Maquet/TruSystem OR Tables are excluded) The cradle movement automatically stops in horizontal position. This enables the user to tilt the tabletop to lift or lower the patient's head relative to his feet (motorized). The tilt rotation center is isocentric (except for TruSystem table) within the limits set by the particular table type. Customizable, the stand can follow the tilt movement in a synchronous way (except for TruSystem table). The tilting movement automatically stops in horizontal position. Isocentric tilt on the Magnus table requires a license. Moves the table up/down (motorized). The movement range depends on the table type and on the table tilt/cradle angle. Provides additional longitudinal span to bring the table closer to the stand (AD7-X(N)T only). The table base can be positioned in one of two places (movement is motorized). This applies to motorized or manual positioning of the tabletop in longitudinal and/or lateral direction. Table-type dependent manual positioning (so called 'float') is balanced (passive or motor-assist) to prevent the patient from sliding down (e.g. in case of tilt). Table: Categories/ items Max. notes: (all values in seconds) Move shutters & wedges 0.5 from UI up to & including image display. APC recall (start geometry movement) 1.0 1.5 for Maquet & TruSystem: 2.5 if the table is in hibernation mode Start all other movements 1.0 1.5 for Maquet & TruSystem: 2.5 if the table is in hibernation mode This enables the user to rotate the table base either to free a patient or to expose off-center body parts (not motorized).
UNS.SmartSuite.TableSideTextEntry | Allow table side text entry without compromising sterility/ hygiene. New, N.A. | SRS.Allura.Intgr.Flexvision.KeyboardMouse ¦ SRS.Allura.Intgr.Flexvision | With FlexVision, the system shall allow for remote control of auxiliary systems via on-screen keyboard and touchpad (on TSM) and / or wireless mouse. This requires a license. FlexVision enables flexible viewing for displaying multiple video sources (live images but also video from other systems) on one large screen display in the exam room with which the user can determine the layout, the position and size of multiple video sources on one large screen display in the exam room. Layouts are predefined. A layout divides the large screen in viewports. The user shall be able to select a preset (a preset contains the layout and for each viewport the specific video source shown in the viewport) to be applied on the large screen. The user shall be able to create, delete, change, save and rename a preset. The user shall be able to select which source is displayed in which viewport. The user shall be able to hide/unhide the video source display in a viewport. The user shall be able to swap the video sources of two viewports in a preset. The user shall be able to enlarge a specific viewport and restore it to its original size. The user shall be able to display video sources in native resolution. Display of 2k2 images shall be supported. Video can originate from the system itself and from auxiliary systems (configurable set, e.g. Interventional Workspot, IVUS). In combination with VoIP based infra, FlexVision requires a license. The display of 2K2 images requires an additional license.
UNS.SmartSuite.TaskGuidance | Receive displayed step-by-step guidance/ feedback for purpose of: Preparing/ performing clinical tasks efficiently (e.g. rotational scan); Minimizing the need for re-training; Allowing existing staff to train new staff; Show this taskguidance on large screen (yes/no) or only at TSM; Open this taskguidance yes or no. New, N.A. | SRS.Allura.Func.ProcedureCard ¦ SRS.Allura.UI.Help ¦ SRS.Allura.UI.Help.Customize | The user shall be offered a procedure based workflow by having procedure specific settings for a procedure step. A procedure card determines the procedure specific settings, being: The EPX selections the user can choose from during the procedure and a default EPX. The default preset for FlexVision. The default preset for FlexSpot. The default preset for Second FlexSpot. The patient orientation and surgical view Procedure guidance. Specific settings for compatible room specific auxiliary systems Predefined procedure cards shall be delivered with the system. The user can customize procedure cards and organize/group them. The user can set which procedure card shall be used when a procedure step is selected for acquisition. This can also be set automatically based on information from the WLM system (RIS code). The user shall be able to access integrated help and electronic Instructions for Use. This includes tooltips, task and procedure guidance and instruction material. The user shall be able to search for a specific topic within the Instructions for Use. The user shall be able to delete and modify a help header and link a new help header to an imported hospital protocol/checklist. This requires a license.
UNS.SmartSuite.TestandVerification | Collect evidence, by performing test and verify the system in the image-guided lab, on critical parameters related to safety and performance, to indicate that the system is operational according to manufacturer specification. *a functional test makes sure that the system, subsystem or Field Replaceable Units functions correctly. | SRS.Allura.Qual.FSResponseTimes ¦ SRS.Allura.UI.ServiceUILanguage ¦ SRS.Allura.Serv.AcceptanceTestReport ¦ SRS.Allura.Serv.UserInterface ¦ SRS.Allura.Intgr.Interfaces.ServiceDataRetrieval ¦ SRS.Allura.Serv.Adjust ¦ SRS.Allura.Serv.Logging ¦ SRS.Allura.Func.ServiceStudies ¦ SRS.Allura.Qual.IQFactory/ServiceMeasurements ¦ SRS.Allura.Serv.VerificationTests | Table: Categories/ items Max. notes Field-Service Tool sec Local start-up service framework: - from start-up to log-on screen 15 - from log-on screen to display framework 10 after entering service key password - start-up with already logged on service user 20 Remote start-up service framework 100 Local shut-down service framework: - from close to service conclusion screen 6 after removing service key or after closing the close confirmation screen - from service conclusion screen to shut-down 3 EPX Management sec Start-up EPX Management Tool 30 includes loading of database settings in editor Store modified settings for verification 30 up to modified settings active for try-out Batch verification (one modified EPX-Application/Procedure) 30 +10 sec per additional modified EPX-Application/Procedure Store modified changes to system 120 up to usable clinical system Calibration min Calibrate FD20 90 Calibrate FD15 90 Calibrate FD12 90 Frontal Stand calibration 340 calibrate geometry without X-ray Lateral Stand calibration 225 calibrate geometry without X-ray Table calibration 75 calibrate geometry without X-ray Generator calibration 90 Image transfer sec Image max. 5122 - 12 bit 150 over 40 Kbps link Image Quality This section describes the basic Image Quality requirements. The purpose is to describe the system in terms of Image Quality requirements such, that if a system meets the requirements in this specification, the image quality of that system is guaranteed (and acceptable for the end user). The service UI for service engineers is in English. QUALITY AND PERFORMANCE REQUIREMENTS In this chapter the requirements are given for the quality and performance requirements. X-ray Performance Figures The system shall be able to generate an acceptance test report that covers all necessary tests for a formal handover to the customer. Diagnostics, Faultfinding Faultfinding is used in corrective maintenance to localize a problem and to determine follow-up actions. The service UI shall provide access to the whole system. The service UI shall be based on a standardized service framework and the EPX management tool. The system shall provide on-line help information. The system shall provide communication interfaces to retrieve service data from compatible auxiliary systems and provide it to the (remote or local) service user Regularly saved log-files Diagnostic data saved for technical support Build-in correction measures shall provide adjustment functionality to bring the system within specified limits. The service framework shall provide adjustment criteria and procedures to the FSE and hospital technician. System information and events are stored and accessible e.g. for diagnostic/faultfinding purposes. Logged info comprises of system-info (e.g. configuration), system usage (e.g. user-actions, system-errors, wear-out indicators), and test results. FSE functions: read (select/filter/search), clear logfile, export logfile. Auxiliary Functions These are special studies, needed to support the field-service functionality. During clinical usage the service studies are not visible. During field-service, both clinical and service studies must be visible. Requirements for installation and manufacturing quality, related to guarantee reproduction of IQ. All system-wide level 1 measurements shall be executed and passed as part of manufacturing and/or installation and/or preventive and/or corrective maintenance.Additionally, lower-level measurement tests shall be available to be executed for diagnostic and regulatory purposes. Table: What is measured Level 1 Comments Dose per image detector Dose Test and Dose Calibration Noise estimate in ROI on FD Dose Test and Dose Calibration Stability noise value in ROI on FD Fast Dose Verification kV-Stabilized Energy and Signal flow 1 see SRS.Allura.Qual.IQKVStabilized Video stabilization level Energy and Signal flow 1 see SRS.Allura.Qual.IQWorkingPoint mAs-Stabilized Test Shot Energy and Signal flow 2 see SRS.Allura.Qual.IQmAsStabilized mAs-Stabilized exposure Energy and Signal flow 2 see SRS.Allura.Qual.IQmAsStabilized Calculated power to validate EDL function EDL verification Regulatory requirement described by FDA: 21CFR1020.32 (clause(d)). Contrast Monitor Performance see SRS.Allura.Qual.IQDisplaySignalTransfer Vignetting Monitor Performance see SRS.Allura.Qual.IQBrightnessUniformity Sharpness Monitor Performance see SRS.Allura.Qual.IQDisplaySignalTransfer Low contrast check Monitor Performance see SRS.Allura.Qual.IQDisplaySignalTransfer What is measured Lower Level DAP difference AEP meter and dose probe Skin dose difference AK displayed and dose probe DAP Measurement and AirKerma Verification Regulatory requirements described by IEC60601-2-43 (clauses:203.6.4.5 Dosimetric indications) and 21CFR1020.32 (clause (k)(6)) Test to check alignment of x-ray beam with detector field of view Field limitation Regulatory requirements described by IEC 60601-1-3 (clause 8.5.3), FDA 21CFR 1020.32 (clause (b)(2) etc), IEC 60601-2-43 (clause 203.8.5.3). Test for verifying x-ray beam vs detector entrance plane Xray Beam alignment see SRS.Allura.Qual.IQBeamAlignment & isolated part of Field Limitation test (without sagging and stand maximal application conditions). Horizontal focal spot Vertical focal spot Focal Spot size Focal spot size is conform IEC60336. Focal spot sizes are included in Performance Test Limits for fault finding. Small focus of x-ray tube Large focus of x-ray tube Limiting Resolution System Level SRS.Allura.Qual.IQLimitingResolutionSystemLevel Beam quality x-ray beam in mm Al Half Value Layer Regulatory requirements described by IEC 60601-1-3 (clause 7.1, half value layers in x-ray equipment). X-ray output x-ray tube X-ray Output SRS.Allura.Qual.IQX-rayOutput Check difference real kV and kV indication ∆kV for 6 measurements (5 manual exposure, 1 fluo technique) kV-Beam is kV-Desk Regulatory requirement described by IEC 61223-3-1 (clause 5.2.1, X-ray tube voltage). Qualitative image impression for Germany Monitor SMPTE Regulatory requirement described by German DIN6868-157:2014-11 What is measured Other Air Kerma verification Dose measurement for verification of the Entrance Doserate Limitation (EDL) mechanism Regulatory requirement described by FDA: 21CFR1020.32 (clause(d)). PHYSICAL REQUIREMENTS Mechanical No specific requirements apart from targets as derived from requirements stated in section 12.3 Standards, Legal Requirements. Tests are provided to verify proper functioning after installation, upgrading or repair. Tests shall apply to safety, function and quality related checks (e.g. X-ray dose, image quality, responsiveness). Checks shall provide objective verification results and do not require expert knowledge from the user. Built-in test patterns (e.g. SMPTE test image and images to support DIN 6868-157) shall provide a standard visual reference.
UNS.SmartSuite.TouchScreenModule | Place Touch Screen Module close to user without impact on sterility; Position the Touch Screen Module at head/side/foot side of table; Position the Touch Screen Module at a mounting device on a cart; Control the Touch Screen Module sterile/unsterile; Use up to 3 TSM's allowing both synchronized and parallel operation. Change contrast and brightness settings, zoom pan, subtraction on-off, select examination types (EPX) and recall previously stored geometry positions, APC control; Control (view and manipulate) 3rd party systems (incl. UltraSound) at table side with/without the use of a single mouse (in case of FlexVision (Pro), the mouse should control FlexVision (Pro) and (CWIS) third party; Have full control of large display FlexVision Pro (layout, applications, pre-sets and viewport sizes); On the TSM pro: View and navigate live clinical images on TSM in 2D; Play runs on TSM; Manually drag shutters/wedges on TSM. Control a mousepointer on the screen (s) to improve communication in ER and CR; View the user interface on the TSM Pro of the connected modality as presented by the 3rd party; Configure the order of Icon appearance on TSM Pro based on user preference Control (view and interact) 3rd party systems at table side using touch screen operation; Change contrast and brightness settings, zoom pan, subtraction on-off, vascular trace, select examination types (EPX) and recall previously stored geometry positions, APC control. New, N.A. | SRS.Allura.Func.PresettingFunctions ¦ SRS.Allura.Func.PanZoom ¦ SRS.Allura.Func.DynamicViewing ¦ SRS.Allura.Func.APC3DReference ¦ SRS.Allura.Intgr.Interfaces.ControlAuxiliarySystems ¦ SRS.Allura.Intgr.Interfaces.RealTimeImageLink ¦ SRS.Allura.Func.Move/ResetWedges ¦ SRS.Allura.Func.NormalSubtraction ¦ SRS.Allura.Conf.UI-RelatedItems ¦ SRS.Allura.Func.APCPredefined ¦ SRS.Allura.UI.UIModule ¦ SRS.Allura.UI.Workspot ¦ SRS.Allura.UI.TSM.Image ¦ SRS.Allura.Intgr.Flexvision.KeyboardMouse ¦ SRS.Allura.Func.Move/ResetShutters ¦ SRS.Allura.Func.APCReference ¦ SRS.Allura.Func.APCTable ¦ SRS.Allura.UI.Image.Pointer ¦ SRS.Allura.Intgr.Interfaces.ControlFromAuxiliarySystems ¦ SRS.Allura.Func.Cycle ¦ SRS.Allura.Intgr.Flexvision ¦ SRS.Allura.Func.SetImageEnhancement ¦ SRS.Allura.UI.TSM.Workspot ¦ SRS.Allura.Func.APCStore/Recall | Presetting is determined by the information in the study/procedure step that is current for acquisition. The user can modify the information. The procedure card to be used can be modified. The performing physician can be selected from a pre-defined physician list. The patient type can be entered/modified manually. The EPX application/procedure group/procedure can be selected from a list (which contains only items that can be performed within the system's current configuration). EPX application, procedure group and procedure can be selected in control room, as well as in exam room. EPX application, procedure group and procedure can be selected during x-ray. The selected settings will then be applied for the next exposure or fluoroscopy run. Enables the user to zoom the image on display and pan the zoom-center. The user can determine the zoom factor. The user can reset panning towards the image center. Annotation & Overlay Text This enables the user to view a run in cine mode, e.g. to show the dynamic behaviour of a beating heart. In cine mode, interactive image processing (e.g. zoom, subtraction) remains available. If the user selects two planes in biplane, the runs are displayed synchronously. The user shall be able to recall the frontal projection (rotation/angulation) that corresponds to a view on the Interventional Workspot (3DRA, XperCT, etc). Indicated by the Interventional Workspot, the SID shall go to minimum or to maximum (maximum to e.g. prevent hitting the needle) during the recall movement. APC 3D reference requires a license. The system shall provide communication interfaces to control compatible auxiliary systems from the systems user interface with wireless mouse on the auxiliary system’s graphical UI displayed on FlexVision with mouse and keyboard on the auxiliary system’s graphical UI displayed on Flexspot from TSM by a generic button based UI tailored by the auxiliary system or by a graphical UI streamed from the auxiliary system including (multi-)touch control. This requires a license control shall be established automatically when (mobile) compatible auxiliary system is connected The system shall provide communication interfaces to provide x-ray images of the frontal channel in real time to compatible auxiliary equipment during x-ray acquisition. This requires a license. Enables wedges to move in/out, or rotate. Reset puts the wedges outside the imaging area. The default position of wedges is EPX application/procedure determined. Minimal movement in acquisition is assumed, to obtain enhanced view with single mask subtraction. Enables that the user can step through the run to see the progressive filling of arteries or veins. The user can switch subtraction on/off, and/or select another mask from the run. Manually, registration errors due to patient movements can be corrected ("Pixel-shift" function). A background portion can be added manually for orientation purposes ("Landmarking" function). Subtracted imaging requires a license. Table: sub tag item monoplane biplane .xs X-ray switches -1 footswitch 3f + -2 footswitch 6f + -3 wireless footswitch 3f o -4 wireless footswitch 6f o -5 additional footswitch CR 3f o -6 additional footswitch CR 6f o -7 additional footswitch ER 3f o -8 additional footswitch ER 6f o -9 X-ray handswitch + + Table: sub tag item Monoplane Monoplane Monoplane XY (fm/fa) XY (fm/fa) XY (fm/fa) biplane biplane biplane sub tag item AD7 XT AD7 XNT Maq AD7 XT AD7 XNT Maq Tru AD7 XT AD7 XNT Maq ui UI-modules -1 ControlModule NT +1) -2 ControlModule Tilt +1) -3 ControlModule Flex +1) +1) -4 ControlModule ORT o2) o2) +1,3) -5 ControlModule ORT Flex o2) o2) +1,3) -6 ControlModule Geo NT7) +1) -7 ControlModule Geo Tilt7) +1) -8 ControlModule Geo ORT7) o2) o2) + -9 ControlModule Geo NT IBR +1) -10 ControlModule Geo Tilt IBR +1) -11 ControlModule Imaging o4) o4) o4) o4) o4) o4) + + + -12 speed controller o5) o5) o5) o5) o5) o5) -13 pan handle o o o o o o -14 TouchScreenModule +6) +6) +6) +6) +6) +6) +6) +6) +6) -15 review module + + + + + + + + + -16 viewpad (cardio) # # # # # # # # # -17 viewpad (mixed) # # # # # # # # # -18 ControlModule Geo ORT IBR o2) o2) + Notes: 1) Additional ControlModule optional for CR 2) Optional as additional ControlModule on pedestal for ER 3) On pedestal 4) ControlModule imaging optionally available for ER (not in combination with 2)) 5) Required for FDPA, not an option for TruSystem 6) Additional TSM optional for CR and (on pedestal) for ER 7) Not for NPI, applicable for modules already present in the installed base Table: sub item Table type Table type Table type notes tag AD 7XT AD 7XNT Maquet TruSystem .mi miscellaneous -1 TSM swingarm + + + -2 Pedestal o o + trolley to attach UI modules; for AD7X(N)T including Pedestal Wall Conn. Box; for ORT including Surgery Wall Conn. Box -3 Two-way Intercom o o o -4 Keyboard + Mouse in CR + + + in CR; additional second keyboard + mouse in case of additional FlexSpot -5 Mouse in ER o o o Requires a license -7 MultiVision Switch o o o see SRS.Allura.Intgr.Multivision -8 KVM-Switch o o o see SRS.Allura.Intgr.KVM-Switching, requires DVI based video infrastructure -10 MCS Auxiliary box + + + for viewpad-receiver and X-on light on 3rd party booms The system shall provide pre-defined sets each comprising a sequence of projections (defined by rotation/angulation, SID and Field of View). The sets shall be EPX defined and a set can have max. 20 projections. The user shall be able to customize the sets. Specific sets can be customized for different users. The user shall be able to recall the projections of a set in sequential order, but alternatively the user shall be able to recall a projection directly. APC predefined requires a license. The review and control modules shall support backlighting to allow easy visibility of buttons in e.g. dark examination rooms and give user guidance. The UI modules at or near the table shall allow control under sterile conditions (e.g. under a sterile cover). The user shall be able to detach the UI modules from the pedestal resp. table and attach them to the table resp. pedestal. Movement and X-ray generating functions shall be controllable at or near the table (e.g on a pedestal). The TSM shall display live X-ray images for image post-processing, image navigation and moving shutters and wedges on the image. Beside this live X-ray image, live images can be shown from applications running on other systems that are compatible with the multi-modality TSM. This requires a license With FlexVision, the system shall allow for remote control of auxiliary systems via on-screen keyboard and touchpad (on TSM) and / or wireless mouse. This requires a license. Enables shutters to move in/out. Opposite shutters move in pairs; horizontal and vertical pairs can be moved simultaneously. Opening stops at the edge of the current detector field (so becomes again visible at a larger field). Closing must be immediately visible (i.e. fast moving to edge of detector field, if initially larger). In case a user-defined position is overruled (by switching towards a smaller field-size), enlarging the field-size will automatically widen the shutters until the original user-defined position is reached. Reset puts the shutters at the edge of the detector field-size. The default position of shutters is EPX application/procedure determined. The user shall be able to recall the geometry positions (i.e. position of frontal and lateral stand, table, shutters and wedges) plus the receptor field size which corresponds to an image that is displayed in a live or reference viewport. In a biplane system, the user shall be able to recall a position for both frontal and lateral channels independently or simultaneously. Note: Table position recall is not supported for TruSystem tables in combination with FlexMove. On AD7X tables recall of longitudinal and lateral table position requires appropriate motors which are optional APC reference requires a license. After the user has brought a point-of-interest into the system’s iso-center, the user shall be able to store the corresponding table position (i.e. table longitudinal, lateral, height, tilt and cradle). The system shall provide a default point-of-interest (EPX defined). When a new acquisition study is selected the 'stored table position' is reset such that the default point-of-interest is in the iso-center. The user shall be able to recall the stored table position or only the table height. The system shall indicate to the user if the table is in the stored position. Note: Not applicable for Hillstrom TruSystem table with universal table top and Maquet Magnus table with universal table top. On TruSystem tables with the SQ14-Xtra table top only height can be recalled. On AD7X tables recall of longitudinal and lateral table position requires appropriate motors, which are optional. APC table requires a license. The user shall be able to show and move a pointer over an area of interest in an image to support collaboration between exam room and/or control room. This requires a license. The system shall provide communication interfaces to allow compatible auxiliary systems to select EPX procedures which may or may not be specific for the auxiliary system (i.e. they are selectable only by the auxiliary system) set APC target positions lock/unlock the table set specific acquisition parameters (e.g. injector coupling, patient orientation, biopsy mode, etc.) push a secondary capture image to a reference viewport This function enables to display a sequence of images in cine mode. In file-cycle all exposure and fluoro runs in the active file of the current study are cycled sequentially. In run-cycle the complete active run is cycled. User can set speed. Intra-Image Post-Processing Focuses on intra-image processing functions for customized enhanced display. General Concepts FlexVision enables flexible viewing for displaying multiple video sources (live images but also video from other systems) on one large screen display in the exam room with which the user can determine the layout, the position and size of multiple video sources on one large screen display in the exam room. Layouts are predefined. A layout divides the large screen in viewports. The user shall be able to select a preset (a preset contains the layout and for each viewport the specific video source shown in the viewport) to be applied on the large screen. The user shall be able to create, delete, change, save and rename a preset. The user shall be able to select which source is displayed in which viewport. The user shall be able to hide/unhide the video source display in a viewport. The user shall be able to swap the video sources of two viewports in a preset. The user shall be able to enlarge a specific viewport and restore it to its original size. The user shall be able to display video sources in native resolution. Display of 2k2 images shall be supported. Video can originate from the system itself and from auxiliary systems (configurable set, e.g. Interventional Workspot, IVUS). In combination with VoIP based infra, FlexVision requires a license. The display of 2K2 images requires an additional license. The user can interactively change image contrast, brightness and/or amount of edge-enhancement to adapt the subjective image quality to personal preference. The TSM shall offer a workspot in examination and/or control room to control system internal applications but also applications running on other systems. The touch screen shall be operable under table side conditions i.e. wearing gloves, covered with a sterile drape and spilled with fluids. The user shall be able to use an application on a certain TSM, while another user shall be able to use any other connected application on another TSM. The internal state of an application that is shown on multiple TSMs shall be synchronized. The system shall allow the user to store and recall multiple geometry positions (i.e. position of stand, table, shutters and wedges) including the receptor field size. When a new acquisition study is selected, all stored positions shall be removed. The system shall indicate to the user if the system has not reached a selected position. Note: Table position recall is not supported for TruSystem tables in combination with FlexMove. On AD7X tables recall of longitudinal and lateral table position requires appropriate motors which are optional
UNS.SmartSuite.TouchScreenModuleFluidresistance | Operate the TouchScreen Module normally when exposed to typical amount of fluids used during the procedure. New, N.A. | SRS.Allura.UI.TSM.Workspot | The TSM shall offer a workspot in examination and/or control room to control system internal applications but also applications running on other systems. The touch screen shall be operable under table side conditions i.e. wearing gloves, covered with a sterile drape and spilled with fluids. The user shall be able to use an application on a certain TSM, while another user shall be able to use any other connected application on another TSM. The internal state of an application that is shown on multiple TSMs shall be synchronized.
UNS.SmartSuite.TransferPatientFromTable | Lower the table top allowing the patient to get off the table in a comfortable way; Lower the table top allowing the staff to transfer the patient from the table in a comfortable way; Lock the table top preventing unintended table movement during patient transfer; To use the system in combination with an OR table; Easily put aside or remove and mount (without sliding) the flush mounted TSM (e.g. for allowing the transfer of the patient, working on other side). New, N.A. | SRS.Allura.Intgr.Magnus ¦ SRS.Allura.Func.TableLock ¦ SRS.Allura.Func.ChangingTableHeight | Compatibility with Magnus OR table of Maquet. Move table via system’s UI modules and Magnus UI modules to position table top, tilt, cradle and set height. Move table for reset geometry, APC iso-center and FDPA. Stop table movements in case of emergency stop. Detect collisions and allow override table movements. Synchronize table lock setting, patient orientation setting, table positions, status and capabilities. The user shall be able to lock/unlock: The table pivot movements. The table lateral movements. All table movements. For the TruSystem table in combination with FlexMove, the lateral table lock/unlock is not supported via the system UI. Beam Positioning This section focuses on moving and positioning the stand under direct user control. Moves the table up/down (motorized). The movement range depends on the table type and on the table tilt/cradle angle.
UNS.SmartSuite.UnpackandTransport | Have guidance to unpack (parts of) the system components. Have options to transport* (parts of) the system for installation at the hospital. *Transport is getting the equipment from unloading the truck at the hospital to the final destination in the hospital.The goal is to take all obstacles into account, like doors, elevator, corridor (bends), etc. | SRS.Allura.Serv.FSInstructions ¦ SRS.Allura.Phys.ClimaticEnvironmentalConditions ¦ SRS.Allura.Phys.MechanicalEnvironmentalConditions | Instructions are available on how to unpack, install, de-install, decommission, and packing and disposal of the system and components. Following are the expected climatic conditions to which systems classified for operation in "Indoor, non-regulated room" will be exposed. The system shall not be damaged and remain safe after exposure to these conditions, while in operation, transport or storage. Table: Conditions representative for: Description of requirement Environmental Conditions Operation Temperature (Performance) +10°C/+30°C Operation Temperature (Safety) +10°C/+35°C Operation Maximum Temperature for unit in rack Ambient T + 5-15 °C (performance; maximum 45 °C) Operation Temperature change rate ≤ 0.5 °C/min Operation Relative Humidity (Performance) 20% - 80% Operation Relative Humidity (Safety) 20% - 93% Operation Condensation No Transport & Storage Temperature Constant: -25°C (3 days) +70°C (4 days) Cycles 25°C/ 97% RH - 40°C/ 93% RH (6 cycles) -10°C - 45°C (1 cycle) Transport & Storage Relative Humidity 5% - 95% Transport & Storage Condensation Possible1) Operation, Transport & Storage Air Pressure 700 - 1100 hPa Note: 1) The packaging during transport and storage protects PC cabinets against moisture from condensation. Directly after exposure to transport and storage conditions, the exposed part or system does not need to be in normal operational mode, but must remain safe. It must return to normal operational mode after acclimatization, once the climatic conditions to which it is exposed are within the operational limits. There shall be no impact on system operation and safety after exposure to repetitive bumps and vibration, that are comparable to the expected mechanical conditions during transport.
UNS.SmartSuite.UseInterventionalTools | To overlay fluoroscopy images on reconstructed data on any 45 degrees position in a range of Z-rotation range of +135 to -135 degrees, taking into account movement compensation and interactions (excluding table tilt, table cradle and table pivot) New, N.A. | SRS.Allura.Intgr.Interfaces.RealTimeImageLink ¦ SRS.Allura.Intgr.Interfaces.DataExchange ¦ SRS.Allura.Intgr.Interfaces.ControlFromAuxiliarySystems ¦ SRS.Allura.Intgr.Interfaces.ControlAuxiliarySystems | The system shall provide communication interfaces to provide x-ray images of the frontal channel in real time to compatible auxiliary equipment during x-ray acquisition. This requires a license. The system shall provide communication interfaces to exchange data with compatible auxiliary systems Patient and examination related data Acquisition related data Dose data ECG data Geometry related data The system shall provide communication interfaces to allow compatible auxiliary systems to select EPX procedures which may or may not be specific for the auxiliary system (i.e. they are selectable only by the auxiliary system) set APC target positions lock/unlock the table set specific acquisition parameters (e.g. injector coupling, patient orientation, biopsy mode, etc.) push a secondary capture image to a reference viewport The system shall provide communication interfaces to control compatible auxiliary systems from the systems user interface with wireless mouse on the auxiliary system’s graphical UI displayed on FlexVision with mouse and keyboard on the auxiliary system’s graphical UI displayed on Flexspot from TSM by a generic button based UI tailored by the auxiliary system or by a graphical UI streamed from the auxiliary system including (multi-)touch control. This requires a license control shall be established automatically when (mobile) compatible auxiliary system is connected
UNS.SmartSuite.UserInterfaceAccess | Control the system at all tableside positions (head-end, foot-end, physician side, nurse side); Remotely control the system (e.g. via pedestal or from CR); View and interact with clinical images and runs on the TSM; Switch between modalities using touch and swipe functionality. New, N.A. | SRS.Allura.Func.SelectTableSide ¦ SRS.Allura.Conf.UI-RelatedItems ¦ SRS.Allura.Func.Concurrency ¦ SRS.Allura.Func.PriorityConstraints ¦ SRS.Allura.UI.TSM.Image ¦ SRS.Allura.Serv.RemoteDesktop.ClinicalMode | The user shall be able in table-oriented mode to maintain intuitive movement control from four table sides. The user shall be able to select the table side on every movement-control UI module individually. Combined Patient- And Beam Positioning This section focuses on movements under direct user control, which are not stand or table specific. Table: sub tag item monoplane biplane .xs X-ray switches -1 footswitch 3f + -2 footswitch 6f + -3 wireless footswitch 3f o -4 wireless footswitch 6f o -5 additional footswitch CR 3f o -6 additional footswitch CR 6f o -7 additional footswitch ER 3f o -8 additional footswitch ER 6f o -9 X-ray handswitch + + Table: sub tag item Monoplane Monoplane Monoplane XY (fm/fa) XY (fm/fa) XY (fm/fa) biplane biplane biplane sub tag item AD7 XT AD7 XNT Maq AD7 XT AD7 XNT Maq Tru AD7 XT AD7 XNT Maq ui UI-modules -1 ControlModule NT +1) -2 ControlModule Tilt +1) -3 ControlModule Flex +1) +1) -4 ControlModule ORT o2) o2) +1,3) -5 ControlModule ORT Flex o2) o2) +1,3) -6 ControlModule Geo NT7) +1) -7 ControlModule Geo Tilt7) +1) -8 ControlModule Geo ORT7) o2) o2) + -9 ControlModule Geo NT IBR +1) -10 ControlModule Geo Tilt IBR +1) -11 ControlModule Imaging o4) o4) o4) o4) o4) o4) + + + -12 speed controller o5) o5) o5) o5) o5) o5) -13 pan handle o o o o o o -14 TouchScreenModule +6) +6) +6) +6) +6) +6) +6) +6) +6) -15 review module + + + + + + + + + -16 viewpad (cardio) # # # # # # # # # -17 viewpad (mixed) # # # # # # # # # -18 ControlModule Geo ORT IBR o2) o2) + Notes: 1) Additional ControlModule optional for CR 2) Optional as additional ControlModule on pedestal for ER 3) On pedestal 4) ControlModule imaging optionally available for ER (not in combination with 2)) 5) Required for FDPA, not an option for TruSystem 6) Additional TSM optional for CR and (on pedestal) for ER 7) Not for NPI, applicable for modules already present in the installed base Table: sub item Table type Table type Table type notes tag AD 7XT AD 7XNT Maquet TruSystem .mi miscellaneous -1 TSM swingarm + + + -2 Pedestal o o + trolley to attach UI modules; for AD7X(N)T including Pedestal Wall Conn. Box; for ORT including Surgery Wall Conn. Box -3 Two-way Intercom o o o -4 Keyboard + Mouse in CR + + + in CR; additional second keyboard + mouse in case of additional FlexSpot -5 Mouse in ER o o o Requires a license -7 MultiVision Switch o o o see SRS.Allura.Intgr.Multivision -8 KVM-Switch o o o see SRS.Allura.Intgr.KVM-Switching, requires DVI based video infrastructure -10 MCS Auxiliary box + + + for viewpad-receiver and X-on light on 3rd party booms This addresses the possibility to use the same or different functions by multiple users at the same time. Administrative functions can be performed in the control room concurrent with activities in the exam room with exception of Bolus Chase Reconstruction and 2D-Quantitative Analysis. Local or remote service monitoring can be performed while the system is in use. Distribution/networking, printing and archiving can be performed with progress in the background. Based on switch-to licenses, auxiliary systems can be operated concurrently. The control room user can work on the same or a different patient and/or study from the exam room. (With FlexSpot) Two control room users can perform different functions on the same patient/study. Priority Handling In case of conflicting or ambiguous user requests the following constraints shall apply: Overruling priority (high to low): 1) exposure, 2) fluoro, 3) other functions (except movement requests and concurrent working in control room) that are applied to the current acquisition study. A function shall be executed as long as requested from at least one UI, with exception of motorized movement requests: parallel movement requests shall be served as long as they are not conflicting. External (surgery) tables are equipped with their own user interface modules. For Maquet Magnus, parallel table movement requests shall result in stopping of the motorized table movements. A movement request via a TruSystem UI will stop a table movement requested via the system UI, without starting the newly requested movement. A table request via the system UI is ignored as long as a table movement request via the TruSystem is active. Exception Handling The TSM shall display live X-ray images for image post-processing, image navigation and moving shutters and wedges on the image. Beside this live X-ray image, live images can be shown from applications running on other systems that are compatible with the multi-modality TSM. This requires a license It shall be possible to provide application support during a clinical procedure with remote desktop. A remote (clinical) expert shall be able to observe the case, give advice, and/or assist in controlling the system. It shall be possible to disable remote activity in the system configuration.
UNS.SmartSuite.UserInterfaceOrientationConcepts | Operate User Interface (UI) control devices (e.g. joystick) intuitively, meaning that the movement direction of table/stand and shutters/wedges should be: As much the same as the relative direction of the UI control device; As expected by taking into account: The location of the physician; The type of procedure (e.g. ‘diagnostic view’ for regular procedures ‘surgical view’ for biopsy procedures); The orientation of the displayed image. New, N.A. | SRS.Allura.Func.SelectTableSide ¦ SRS.Allura.Func.RotateBeam ¦ SRS.Allura.Func.PatientOrientation ¦ SRS.Allura.Func.BeamShapingControls ¦ SRS.Allura.Func.DisplayOrientation ¦ SRS.Allura.Func.SelectSurgicalView | The user shall be able in table-oriented mode to maintain intuitive movement control from four table sides. The user shall be able to select the table side on every movement-control UI module individually. Combined Patient- And Beam Positioning This section focuses on movements under direct user control, which are not stand or table specific. The user shall be able to rotate the beam around the longitudinal-table-axis (motorized). Intuitive control and display orientation rules (see SRS.Allura.Func.DisplayOrientation) also apply if rotation causes a reversed beam direction. In biplane systems, the frontal and lateral beam can be rotated simultaneously. During motorized movement the rotation position shall be displayed. The patient can be positioned on the table in one of 4 ways (nose up/down, legs down/up). The actual patient orientation (nose up/down, legs down/up) can be entered into the system. Default orientation is procedure card dependent. The orientation can be changed by the user; the override remains valid until a new procedure card is selected. This setting is also used for intuitive control of shutters and wedges. Controls are made intuitive in the sense that orientation with respect to displayed image is maintained. The patient shall be displayed ‘head-up’, facing towards the user, regardless of patient orientation on the table (nose/legs up/down), and stand/detector orientation. This ‘diagnostic view’ adheres to standard clinical practice and prevents diagnostic risks in mistaking left/right patient anatomy. For the lateral channel in a biplane system the same rule applies if the detector is rotated above the table, the image shall not be flipped if the lateral stand is rotated towards a lateral position. There are 2 exceptions: if ‘surgical view’ is chosen, the image will be left/right reversed in case a patient is positioned nose down (rationale: in order to have intuitive eye-hand coordination). in neuro applications (EPX controlled) the lateral-channel display will be left/right reversed (rationale: based on neuro preference; the lateral view has no risk of misinterpreting left/right anatomy). Enables the user to select between surgical view and diagnostic view. The surgical view is only available on a configuration with an FD20 and FD15. The reversed acquired images must be clearly marked (e.g. inserted icon) to prevent diagnostic mistakes later on.
UNS.SmartSuite.UserSpecificSettings | : Define system specific settings; Transfer specific settings from one system to another; Define procedure specific settings for the hospital; Customize hospital specific information e.g. physician list, ProcedureCards, dose in AK/DAP; Import/export ProcedureCards. New, N.A. | SRS.Allura.Func.ModifyPhysicianList ¦ SRS.Allura.Func.SystemCustomization ¦ SRS.Allura.Func.Display/SetTimeAndDate ¦ SRS.Allura.UI.Help.Customize | The user shall be able to create and change the physician list. Physician names can be added, deleted, and/or modified. If the system is connected to a hospital Worklist Management System, the name of the performing physician of a scheduled procedure step shall be automatically added to the physician list. Allows the user to customize system settings, that are available for the system's configuration, such as: Date, Time and Regional Settings User Administration and Administration settings List of annotation and performing physician Procedure cards and RIS code mapping Settings for viewing and quantitative analysis Media and DICOM export settings APC positions Linking DICOM export settings, APC positions and procedure help text to EPX procedures Remote support settings Defining APC pathways Saved settings are persistent over system restarts. System customization settings can be imported and exported. Administration Functions This section focuses on administrative functions to deal with patient data and related information. General Administration Concepts A local database includes all information that is entered manually and generated in acquisition. It can contain multiple patients/studies/procedure steps comprising patient data, images, etc. Local Administration Functions The local time is indicated to the user. Display format can be changed (e.g. d/m/yr, yr/m/d); standard PC practices. Time is set either via Network Time Synchronization or by the user. Miscellaneous General Functions The user shall be able to delete and modify a help header and link a new help header to an imported hospital protocol/checklist. This requires a license.
UNS.SmartSuite.VerticalArmConfiguration | Place the ceiling mounted system in high and low ceiling rooms with a resp. heights of 310 cm (FlexMove), 290 cm (Long L-arm) and 270 cm (Short L-arm). New, N.A. | SRS.Allura.Conf.Stand-RelatedItems | Table: sub Poly-G Poly-G Poly-G Clea Clea Clea Clea L-arc tag Item Ce XY fm Fl Ce XY fm XY fa Fl Bpl notes .std -1 Extended rail o o o o o o FlexArm: 4.3m; standard: 6.0m or 7.8m. -2 L-arm long # # # # -3 L-arm short # # # # -4 Flexarm long # -5 Flexarm short # -6 Ceiling rail covers o See SRS.Allura.Phys.LaminarAirflowCompatibility -7 Optical imaging system o o o see SRS.Allura.Intgr.OpticalImagingSystem -8 Coverbags c c c c c c c c Arc / Tube / Detector / TSM
UNS.SmartSuite.VeteransAdministration | Use a system that complies with the mandatory requirements from Veterans Administration (VA) New, N.A. | SRS.Allura.SLS.VA | Compatibility with mandatory requirements from Veterans Administration (VA) is required (see [REF-1]). This applies to positioning-related requirements for table and/or stand, requirements for radiation and image generation and DICOM interoperability aspects.
UNS.SmartSuite.VideoStreamIntegration | Stream video output signal to other rooms for purpose of i.e. communication, teaching and recording. New, N.A. | SRS.Allura.Intgr.VideoSlave | VideoSlave enables showing a copy of an X-ray video source (e.g. Live/Ref1, Ref2/Ref3, Review for monoplane system) or auxiliary system video source on a static video output. The static video output can be connected to a slave monitor (powered by the system) or to an external monitor/video system. In combination with IP based video infrastructure: maximum of 2 slave monitors in CR (close to the main workspot) and additionally maximum of 10 slave monitors at flexible location (in CR/ER outside of MCSs/Booms, part of total pool of 20 FHD monitors) maximum of 40 outputs for use outside the system. Every video output requires a counted license. The license does not actually limit the number of video outputs. DP is the video output standard Duplicating a X-ray video source that is already displayed on another monitor requires a license. The licenses are counted licenses for the number of duplicated video outputs. In combination with DVI based video infrastructure: maximum of 5 slave monitors (from that max 3 in CR) maximum of 8 outputs for use outside the system DVI-D is the video output standard
UNS.SmartSuite.VideoSwitching | Easily switch X-ray and external sources across the monitors in the Exam room (non-FlexVision). New, N.A. | SRS.Allura.Intgr.SwitchableMonitors ¦ SRS.Allura.Intgr.VideoDisplayInterface | The system shall provide an option to flexibly switch during the procedure all available video sources to monitors in the examination room (within or outside MCSs/booms) and control room. Video sources can be displayed on at most 16 monitors. Switchable Monitors requires a license. In combination with DVI based video infrastructure, if FlexVision is present at most 8 switchable monitors are supported. In VoIP based infrastructure the license is a counted license for the number of Switchable Monitors. The system shall support the display of its own video sources and/or sources provided by auxiliary systems on its examination/control room monitors and large screen.
UNS.SmartSuite.ViewingNavigation | Review the acquired images of the current examination or images from earlier examinations (from same and different patients); Run/replay image by image; Scan through available image data for overview and selection of relevant images; Clearly see when the displayed patient is not the acquisition patient. New, N.A. | SRS.Allura.Func.SwitchDisplayFocus ¦ SRS.Allura.Func.SelectViewingStudy ¦ SRS.Allura.Func.MosaicOverview ¦ SRS.Allura.Func.Cycle ¦ SRS.Allura.Func.DynamicViewing ¦ SRS.Allura.Func.ShowStudyInfo ¦ SRS.Allura.Func.DisplayStudyList ¦ SRS.Allura.Qual.ClinicalResponseTimes.Review ¦ SRS.Allura.Func.NavigateToImageData | This function switches focus from Exam Display towards Reference Display and vice-versa. Acquisition implicitly sets the display focus to the Exam Display. Enables the selection of a current viewing study (besides an already selected study for acquisition). The system shall indicate if the current viewing study differs from the study selected for acquisition. Images are minified and displayed as an N*N mosaic in one or more pages. In Run-Overview the mosaic displays each image in the run. In File-Overview the mosaic displays a single image from each run in the file. In mosaic display mode navigation functions remain available. This function enables to display a sequence of images in cine mode. In file-cycle all exposure and fluoro runs in the active file of the current study are cycled sequentially. In run-cycle the complete active run is cycled. User can set speed. Intra-Image Post-Processing Focuses on intra-image processing functions for customized enhanced display. General Concepts This enables the user to view a run in cine mode, e.g. to show the dynamic behaviour of a beating heart. In cine mode, interactive image processing (e.g. zoom, subtraction) remains available. If the user selects two planes in biplane, the runs are displayed synchronously. The user shall be able to view the information of a selected study, comprising at least: Study info: overview of runs, applied EPXs, cumulative dose, fluoro time. Run/image info: exposure parameters, projections. The user shall be able to filter the runs within the selected study. This function lists the clinical (not field-service) studies present in the system's database. The presentation is filtered (e.g. schedule list only). The list can be sorted and searched for easy navigation to the targeted patient/study. Navigation within patient/study is enabled through the display of representative images ("pictorial index"). Table: Categories/ items Max. notes: (all values in seconds ) Image stepping 0.2 Allowed 0.4 sec for run-subtract. Run stepping, non-subtract 0.3 Run stepping, subtract 0.6 Run overview 4*4, subtract & non-subtract 2 Run overview 4*4, run-subtract 4 File overview 4*4, non-subtract 2 File overview 4*4, subtract 4 Reviewing requirements are not applicable to mosaic overview modes, unless explicitly indicated otherwise. Enables the user to zoom-in progressively on a particular file and/or a run, down to a specific image. After selection of a study for viewing or for acquisition, the first exposure or fluoro run is active for viewing. Exposure overrules the selection and makes the latest acquired run active for viewing.
UNS.SmartSuite.ViewingROI | View from various points the region of interest to ensure a correct 3D impression of the anatomy with optimal imaging conditions. | SRS.Allura.Func.AngulateBeam ¦ SRS.Allura.Func.RotateBeam | Changes the angle between beam and longitudinal-table-axis (motorized). Angulation can be performed simultaneously with beam rotation. During motorized movement the angle position is displayed. The user shall be able to rotate the beam around the longitudinal-table-axis (motorized). Intuitive control and display orientation rules (see SRS.Allura.Func.DisplayOrientation) also apply if rotation causes a reversed beam direction. In biplane systems, the frontal and lateral beam can be rotated simultaneously. During motorized movement the rotation position shall be displayed.
UNS.SmartSuite.Workspot | Perform all activities, corresponding to his/her role in the hospital, from a single control point at the location of his/her preference; Control auxiliary systems on those work spots. New, N.A. | SRS.Allura.Intgr.Interfaces.RealTimeImageLink ¦ SRS.Allura.UI.TSM.Workspot ¦ SRS.Allura.Intgr.Interfaces.ControlFromAuxiliarySystems ¦ SRS.Allura.UI.Workspot ¦ SRS.Allura.Intgr.Interfaces.ControlAuxiliarySystems | The system shall provide communication interfaces to provide x-ray images of the frontal channel in real time to compatible auxiliary equipment during x-ray acquisition. This requires a license. The TSM shall offer a workspot in examination and/or control room to control system internal applications but also applications running on other systems. The touch screen shall be operable under table side conditions i.e. wearing gloves, covered with a sterile drape and spilled with fluids. The user shall be able to use an application on a certain TSM, while another user shall be able to use any other connected application on another TSM. The internal state of an application that is shown on multiple TSMs shall be synchronized. The system shall provide communication interfaces to allow compatible auxiliary systems to select EPX procedures which may or may not be specific for the auxiliary system (i.e. they are selectable only by the auxiliary system) set APC target positions lock/unlock the table set specific acquisition parameters (e.g. injector coupling, patient orientation, biopsy mode, etc.) push a secondary capture image to a reference viewport Movement and X-ray generating functions shall be controllable at or near the table (e.g on a pedestal). The system shall provide communication interfaces to control compatible auxiliary systems from the systems user interface with wireless mouse on the auxiliary system’s graphical UI displayed on FlexVision with mouse and keyboard on the auxiliary system’s graphical UI displayed on Flexspot from TSM by a generic button based UI tailored by the auxiliary system or by a graphical UI streamed from the auxiliary system including (multi-)touch control. This requires a license control shall be established automatically when (mobile) compatible auxiliary system is connected
UNS.SmartSuite.X-rayBeamShaping | Control the radiated area; Restrict X-ray radiation to the clinical region of interest e.g.: Manually positioning shutters; Manually closing/opening shutters (in horizontal and vertical direction); Automatically position wedges (depending on stand rotation and angulation angles) while moving the beam around the patient (should be configurable depending on the procedure); Manually adjust wedges (also if ‘Automatic Wedge Follow’ is active); Prevent over-exposure at clinical region of interest (e.g. manually adjusting wedges that reduce X-ray intensity); Blank part of the displayed image by means of image cropping; Set shutters (positioning, opening/closing) and wedges (moving), with or without radiation; View actual and new positions on display (with and without radiation); See the anticipated radiated area without radiation (e.g. displayed as overlay). Allow (graphical) control of shutters/wedges on the TSM; View actual and new positions on display and TSM Pro (with and without radiation) in relation to X-ray image. New, N.A. | SRS.Allura.Func.Move/ResetWedges ¦ SRS.Allura.Func.ShutterAndWedgeIndicationWithoutRadiation ¦ SRS.Allura.Func.Move/ResetShutters ¦ SRS.Allura.Func.BeamShapingControls ¦ SRS.Allura.Qual.ClinicalResponseTimes.Movement ¦ SRS.Allura.UI.TSM.Image ¦ SRS.Allura.Func.AutomaticWedgeFollowing ¦ SRS.Allura.Func.SetElectronicShutters | Enables wedges to move in/out, or rotate. Reset puts the wedges outside the imaging area. The default position of wedges is EPX application/procedure determined. This enables positioning before a run is started, in order to save patient radiation dose. Re-positioning causes the anticipated area to be displayed as an overlay for an EPX defined time. Enables shutters to move in/out. Opposite shutters move in pairs; horizontal and vertical pairs can be moved simultaneously. Opening stops at the edge of the current detector field (so becomes again visible at a larger field). Closing must be immediately visible (i.e. fast moving to edge of detector field, if initially larger). In case a user-defined position is overruled (by switching towards a smaller field-size), enlarging the field-size will automatically widen the shutters until the original user-defined position is reached. Reset puts the shutters at the edge of the detector field-size. The default position of shutters is EPX application/procedure determined. Controls are made intuitive in the sense that orientation with respect to displayed image is maintained. Table: Categories/ items Max. notes: (all values in seconds) Move shutters & wedges 0.5 from UI up to & including image display. APC recall (start geometry movement) 1.0 1.5 for Maquet & TruSystem: 2.5 if the table is in hibernation mode Start all other movements 1.0 1.5 for Maquet & TruSystem: 2.5 if the table is in hibernation mode The TSM shall display live X-ray images for image post-processing, image navigation and moving shutters and wedges on the image. Beside this live X-ray image, live images can be shown from applications running on other systems that are compatible with the multi-modality TSM. This requires a license Wedges can be positioned automatically, coupled to the position of the stand. The auto-follow position is restored (e.g. after manual overriding) each time the stand is moved. This auto-follow behaviour can be switched on or off by the user; default defined in EPX. Beam Conditioning Enables the user to interactively position 4 rectangular shutters to blank part of the displayed image.
UNS.SmartSuite.X-rayDetectorFieldSize | Zoom in on displayed image; View the region of interest by select/adjust/switch Fields of View (FoV); Identify the current level of magnification during FOV adjustment; Have the possibility to select a predefined Field Of View upon selection of a specific procedure. With the FD12: See the aortic arch and the valve in one view; See the whole coronary tree in one view (cardiac myopathy); See both the needle insertion point (subclavian or axillary vein) and the whole wire/lead towards the right atrium in one view. New, N.A. | SRS.Allura.Conf.SupportedConfigurations ¦ SRS.Allura.Func.SelectDetectorFieldsize ¦ SRS.Allura.Conf.XperFDSeriesConfigurations | The following FD configurations shall be supported: Table: Configuration Characteristic discriminators Azurion FD12 Poly-G Ceiling/Floor stand with FD12 detector Azurion FD15 Clea-Floor stand with FD15 detector Azurion FD20 Clea-Ceiling/Floor stand with FD20 detector Azurion FD12/12 Poly-G Floor with FD12 & Larc-C with FD12 Azurion FD20/12 Clea-Floor with FD20 & Larc-N1) or Larc-C with FD12 Azurion FD20/15 Clea-Floor with FD20 & Larc-N or Larc-C with FD15 Azurion FD20 OR-table Clea-Ceiling stand with FD20 detector Azurion FD12/12 OR-table Poly-G Floor with FD12 & Larc-C with FD12 Azurion FD20/12 OR-table Clea-Floor with FD20 & Larc-C with FD12 Azurion FD20/15 OR-table Clea-Floor with FD20 & Larc-N or Larc-C with FD15 Note 1): Azurion FD20/12 Clea-Floor with FD20 & Larc-N with FD12 is not supported for NPI. This configuration is only supported for the installed base. The following configurations are also supported as remanufactured systems: Table: Configuration Characteristic discriminators Azurion FD12 Poly-G Ceiling stand with FD12 detector and AD7 table Azurion FD20 Clea Floor stand with FD20 detector and AD7 table (no FlexMove) Detailed Configuration requirements Applied symbols: 1) In subsequent tables, additional detailing of configurations uses following abbreviations and indications: Mpl = Monoplane system Ce = Monoplane Ceiling system, suspended by longitudinal rail XY-fa = Monoplane Ceiling system, suspended by double-pivot arm (FlexArm) XY-fm = Monoplane Ceiling system, suspended by X-Y rails (FlexMove) Fl = Floor mounted Monoplane system Bpl = Biplane system (FD12/12 and FD20/12) NB = Neuro Biplane (FD20/15) Mag = Magnus Maquet Tru = Hillrom TruSystem 7500 2) Symbols are used in the tables have following meaning + = included in a system, not an option ++ = twice included in a system, not an option (in case of a Biplane system) # = each system comprises exactly one of the items marked with # in one cell % = each system comprises at least one of the items marked with % in one cell. o = option; is an extension based on extra hardware and/or based on a (software) license key c = compatibility based upon interfacing with the system 3) Use of sub-tags To deal with limited table space, sub-tags are used; to be appended to the specified table tag. Example: Table tag = SRS.Allura.Conf.XperFDSeriesConfigurations; sub-tag = .std-1 Resulting complete tag = SRS.Allura.Conf.XperFDSeriesConfigurations.std-1 The detector field-size is the detector area that is being acquired and displayed. The user shall be able to select a field-size; for fluoro also during acquisition. The available field-sizes and related image quality characteristics depend on the type of detector. The default shall be restored when a new procedure step is selected for acquisition, and EPX-dependent at EPX procedure selection. Table: sub tag Systems: FD12 FD12/12 FD12 FD12/12 FD12 FD12/12 FD20 FD20 FD20 FD20 FD 15 FD20/12 FD20/15 FD20/12 FD20/15 item Ce Fl Bpl Ce XY fa XY fm Fl Fl Bpl NB notes .std Geo Stand -1 Poly-G2-C # Poly-G2-C is supported for remanufacturing -2 Poly-G3-C # -3 Poly-G2-F + + -4 Clea-2-C + + Clea-2-C is supported for remanufacturing (not XY fm) -5 Clea-3-C + -6 Clea-F # # # # -7 Clea-3-F # Only in combination with AD7 table -8 Larc-N 5) # -9 Larc-C # # # -10 Larc-3-N # Only in combination with AD7 table .pt Patient Table -1 AD7XT # # o o o o # # o o -2 AD7XNT # # o o o # # o o -3 Maquet Magnus c c c c c c -4 Hillrom TruSystem c c See SRS.Allura.Intgr.TruSystem .ms Monitor Ceiling Support -1 Rail suspended MCS % % % % % % % % % -2 Spring-arm boom % % % % % % % % % % -3 Spring-arm boom XL % % % % % % % % % % -4 3rd party boom % % % % % % % % % % MCS Monitor integration kit -5 3rd party boom XL % % % % % % % % % % MCS Monitor integration kit XL .xrc X-ray chain ++= frontal & lateral -1 Tube MRC 200+ 0508 + + ++ +1) +1) -2 Tube MRC 200+ 04072) + + + + + +1) +1) -3 Collimator Nicol3) + + ++ + + + + + ++ ++ -5 AEP dose-meter o o o o o o o o o o AEP calculated by default, optionally hardware based -7 Detector FD124) + + ++ # # -8 Detector FD15 + # # -9 Detector FD20 + + + + # # Notes: 1) Frontal channel MRC0407; lateral channel MRC0508 2) MRC 200+ 0407 ROT-GS1008 for FD20 XY-fa; ROT-GS1004 for all other configurations 3) Nicol v4 on Clea-3, L-arc 3-N and Poly-G3 Ceiling. All other stands Nicol v3 4) Monoplane configurations PX2121S or PX2121C/V; biplane configurations PX2121C/V only 5) Azurion FD20/12 Clea-Floor with FD20 & Larc-N with FD12 is not supported for NPI. This configuration is only supported for the installed base
UNS.SmartSuite.processingfocus | Select if the processing applies to the images of both channels or to only one channel. New, N.A. | SRS.Allura.Func.SelectProcessingFocus | This function allows the user to select whether the processing applies to the images of both channels (in case of biplane runs) or only one.