## Sheet1
User Needs Specification Tag | System Requirements Specification Tag | UNS_Content
Not applicable, no user requirement | ['SRS.Allura.Phys.TechnicalRoomCabinetsSpace', 'SRS.Allura.Manf.LoadConfiguration', 'SRS.Allura.I&V.ManualTest', 'SRS.Allura.I&V.Environment', 'SRS.Allura.SLS.ECO.EnergyConsumption', 'SRS.Allura.Manf.VerifySystem', 'SRS.Allura.I&V.TestInterface'] | []
UNS.SmartSuite.2DQAAccuracyAndPrecision | ['SRS.Allura.Func.VesselAnalysis.Accuracy', 'SRS.Allura.Func.VentricularAnalysisAccuracy'] | [('text', 'Background:\xa0Established 2DQCA guidelines for systematic and random error values of a state-of-the-art QCA-system can be found in [Reiber-2009]. For 2DQCA applied to clinical images we require for pooled means and standard deviations [Reiber-1993]:'), ('table', [['Clinical images', 'Systematic error\xa0\n[mm]', 'Random error\xa0\n[mm]'], ['Inter-observer variabilities obstructions', 'n.a.', '< 0.15'], ['Diameter deviation from an equivalent marketed device', '< 0.13', 'n.a.']]), ('text', 'The above numbers are specified for measurements with an accurate Calibration Factor (CF from iso-center calibration or distance calibration).'), ('table', [['Legacy\n\n', 'The user wants to:\n< Clinical users/Physicists/ Service users, IR, ER/CR>\nQuantify vessels and obstructions with 2DQCA with accuracy and precision in line with established 2DQCA guidelines.'], ['New', 'N.A.']]), ('text', '\n')]
UNS.SmartSuite.2DQAAutomatedMeasurement | ['SRS.Allura.Func.VesselAnalysis', 'SRS.Allura.Func.VentricularAnalysis', 'SRS.Allura.Func.QA'] | [('text', ''), ('table', [['Legacy\n\n', 'The user wants to:\n< Clinical users/Physicists/ Service users, IR, ER/CR>\nHave automatic contour detection in 2DQCA and 2DQVA of a user-selected vessel and/or stenosis in an angiogram. The user needs to be able to manually edit the provided vessel contours and obstruction borders and trigger a recalculation of the results;\nAutomatically outline and estimate dimensions of user-indicated coronary segments (sizes 1-6 mm) and obstructions from non-subtracted angiograms (2DQCA);\nAutomatically outline and estimate dimensions of user-indicated aortic and peripheral artery\xa0 (sizes 5-50 mm) segments and obstructions from subtracted and non-subtracted angiograms (2DQVA).'], ['New', 'N.A.']]), ('text', '')]
UNS.SmartSuite.2DQACalibration | ['SRS.Allura.Func.QACalibration'] | [('text', ''), ('table', [['Legacy\n\n', 'The user wants to:\n< Clinical users/Physicists/ Service users, IR, ER/CR>\nAutomatic and manual calibrate to convert distances measured in the X-ray image to real-world distances.'], ['New', 'N.A.']]), ('text', ''), ('text', ' ')]
UNS.SmartSuite.2DQAEfficientToUse | ['SRS.Allura.Func.VesselAnalysis', 'SRS.Allura.Func.VentricularAnalysis', 'SRS.Allura.Func.QA'] | [('text', ''), ('table', [['Legacy\n\n', 'The user wants to:\n< Clinical users/Physicists/ Service users, IR, ER/CR>\nReceive automated assistance when execute time-consuming tasks like drawing a contour around a lumen or object;\nObtain automatically generated reports from user-reviewed results.'], ['New', 'N.A.']]), ('text', '\n')]
UNS.SmartSuite.3DAnatomyScan | ['SRS.Allura.Func.DRAFor3D', 'SRS.Allura.Func.DualPhaseXperCT', 'SRS.Allura.Func.ClassicDRA', 'SRS.Allura.Func.XperCT'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'N.A.'], ['New', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, ER/CR>\nCapture the 3D anatomy with a single scan.']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.3Dacquisition | ['SRS.Allura.Func.DRAFor3D', 'SRS.Allura.Func.DualPhaseXperCT', 'SRS.Allura.Func.ClassicDRA', 'SRS.Allura.Func.XperCT'] | [('text', '\xa0'), ('table', [['Legacy', 'The user wants to:\n< Clinical users/Physicists/ Service users, all disciplines, ER/CR>\nPerform rotational acquisitions from Head and Nurse or Doctor side of the table in the same procedure;\xa0\nPerform 3D rotational acquisitions from Head, Nurse and Doctor side of the table in the same procedure with the FlexArm system and re-use these images for navigation purposes in 7 working positions (with respect to the standard table position): 1) Head side, 2) Nurse side, 3) Doctor side, 4) +45 degrees, 5) -45 degrees, 6) +135 degrees, and 7) -135 degrees.'], ['New', 'Perform 3D rotational acquisitions from Head, and Nurse and Doctor side of the table in the same procedure with the Biplane floor system (20”) and re-use these images for navigation purposes in 6 working positions (with respect to the standard table position): 1) Head side, 2) Nurse side, 3) Doctor side, 4) +45 degrees, 5) -45 degrees and 6) +135 degrees (nurse side);\nPerform 3D rotational acquisitions of the head from the head side of the table along the predefined trajectory.\xa0']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.AcquisitionImageQualityAndDose | ['SRS.Allura.Func.FluoroFlavours', 'SRS.Allura.Qual.IQDisplaySignalTransfer', 'SRS.Allura.Func.SingleShotFootpedal', 'SRS.Allura.Qual.IQScatterReduction', 'SRS.Allura.Qual.IQArtifactsSubjective', 'SRS.Allura.Func.SelectDetectorFieldsize', 'SRS.Allura.Qual.IQWorkingPoint', 'SRS.Allura.Func.ActivateFluoro', 'SRS.Allura.Func.ActivateExposure', 'SRS.Allura.Qual.IQHalfValueLayer', 'SRS.Allura.Func.TimedVariableFramerates', 'SRS.Allura.Qual.IQmAsStabilized', 'SRS.Allura.Qual.IQX-rayOutput', 'SRS.Allura.Qual.IQKVStabilized', 'SRS.Allura.Qual.IQBrightnessUniformity', 'SRS.Allura.Func.ExposureFramespeed', 'SRS.Allura.Qual.IQLimitingResolutionSystemLevel', 'SRS.Allura.Func.AntiScatterGrid', 'SRS.Allura.Func.PresettingFunctions', 'SRS.Allura.Func.IntegratedExposure', 'SRS.Allura.Func.High-DoseFluoroProtection', 'SRS.Allura.Qual.IQContrastResolution2D', 'SRS.Allura.Func.ExposureTechniques', 'SRS.Allura.Func.ChangeFluoroProcessing', 'SRS.Allura.Qual.IQBeamAlignment', 'SRS.Allura.Qual.ClinicalResponseTimes.Acquisition', 'SRS.Allura.Func.DisplayFluoroTime', 'SRS.Allura.Qual.X-rayPerformance', 'SRS.Allura.Func.RunLogging', 'SRS.Allura.Func.X-rayFiltering', 'SRS.Allura.Func.DisplayAndReportPatientDose', 'SRS.Allura.Func.DisplayInfo', 'SRS.Allura.Conf.UI-RelatedItems', 'SRS.Allura.Qual.IQGeoStability', 'SRS.Allura.Func.ManualOverrideEPXParameterValues', 'SRS.Allura.Conf.ImageStorageOptions', 'SRS.Allura.Func.AutomaticElectronicShutters', 'SRS.Allura.Func.DoseControlMeasuringFields', 'SRS.Allura.Func.DoseControlTechniques'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, ER/CR>\nSelect, adjust and save image acquisition settings, for both fluoroscopy and exposures (incl. single shot):\nOn the fly\nFrom various work spots (at table side, in the control room, etc.)\nFor various items, such as: Patient Dose, Frame Rate\nDisplay real-time X-ray images;\nDisplay relevant (and regulatory required) acquisition parameters (e.g. kV, mA, DAP) in viewed images;\nSave relevant acquisition parameters in stored images/ image data.'], ['New', 'N.A.\xa0']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.ActivateFluoroExposure | ['SRS.Allura.Conf.UI-RelatedItems', 'SRS.Allura.Func.ActivateExposure', 'SRS.Allura.Func.ContrastInjection', 'SRS.Allura.Func.SetInjectorCoupling', 'SRS.Allura.Func.SingleShotFootpedal', 'SRS.Allura.Func.AcquisitionMethods', 'SRS.Allura.Func.SelectActiveChannel', 'SRS.Allura.Func.StoreFluoro', 'SRS.Allura.Func.ActivateFluoro'] | [('text', ''), ('table', [['Legacy\n\xa0', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, ER> \nActivate fluoroscopy and exposure;\nUse a hand-switch for exposure;\nUse a wired footswitch for fluoroscopy/exposure;\nUse a wireless footswitch for fluoroscopy/exposure;\nOperate with a single action the acquisition system together with the injector;\nSave a fluoroscopy run retrospectively (store);\nSave a fluoroscopy run prospectively (grab).\nIn case of biplane:\nActivate fluoroscopy from frontal and lateral channel separately as well as simultaneously\nActivate exposure from frontal and lateral channel separately as well as simultaneously\nContinue the examination with one channel if the other channel has an error.'], ['New', 'N.A.']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.AddPatientData | ['SRS.Allura.Intgr.Flexvision', 'SRS.Allura.Intgr.NonDICOMWLMSystems', 'SRS.Allura.Intgr.Flexvision.KeyboardMouse', 'SRS.Allura.Intgr.Interfaces.DataExchange', 'SRS.Allura.Func.ModifyStudy', 'SRS.Allura.Func.Workflow.Flexibility', 'SRS.Allura.Func.AddStudy', 'SRS.Allura.SLS.IHE', 'SRS.Allura.Intgr.Flexspot.AdditionalWorkspot', 'SRS.Allura.Func.Merge', 'SRS.Allura.Intgr.WLMSystems'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'The user wants to:\n<All users,\xa0\xa0all clinical areas, ER/CR>\nImport demographic patient/procedural information from the hospital information system (e.g. RIS/CIS);\nManually enter patient data;\nManually add missing information;\nModify added information;\nAutomatically push/pull information to/from other systems.\nAutomatically trigger the corresponding procedural information (Procedure Card)\xa0\xa0when the system is connected to the hospital information system(e.g. RIS/CIS);\nEnter patient data in the Exam Room.'], ['New', 'N.A.']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.AdditionalFlexSpot | ['SRS.Allura.Intgr.Flexspot.SecondFlexspot', 'SRS.Allura.Conf.Control-RoomDisplays', 'SRS.Allura.Intgr.Flexspot.AdditionalWorkspot'] | [('table', [['Legacy\n\n', 'The user needs to be able to:\n<Clinical users, all clinical areas , CR/ER>\nOccupy a second workspot (next to FlexSpot) in the CR or ER;\nCirculating nurse should be able to perform all tasks during a case (e.g. patient scheduling) without having to leave the ER.'], ['New', 'First and second FlexSpot with one or two 27’monitors(s) per workspot;\nWork in the same way has in the first FlexSpot;\nView and manipulate different sources on the primary and second workspot.']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.Administration | ['SRS.Allura.Serv.Backup/Restore', 'SRS.Allura.Serv.SiteLog', 'SRS.Allura.Serv.SystemID'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'N.A.'], ['New', 'The user needs to:\n<FSE/RSE/CAS, all service workflows>\nDo administration tasks and store evidence* after any service activity to indicate that the system is operational according to manufacturer specification.\n\n*, e.g. import to the Performance Assurance tool and installed base registration.']]), ('text', '\xa0'), ('text', 'Diagnosis'), ('text', ' ')]
UNS.SmartSuite.AmbientVideos | ['SRS.Allura.UI.ComfortThemes'] | [('text', '\xa0'), ('table', [['Legacy', 'The user wants to:\n<Clinical users/Physicists/ Service users, all clinical areas, ER>\nDisplay ComfortThemes videos on FlexVision controlled from TSM.'], ['New', 'N.A.']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.Annotations | ['SRS.Allura.Func.PredefinedAnnotation', 'SRS.Allura.Func.OverrideManualAdjustments', 'SRS.Allura.Func.SelectProcessingFocus', 'SRS.Allura.Func.ModifyStudy', 'SRS.Allura.Func.ManualAnnotation', 'SRS.Allura.Func.SelectProcessingScope'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, ER/CR>\nDefine/set annotations per image and run;\nSelect predefined standard annotations;\nReposition and modify annotations;\nControl the list of annotations;\nAdd free text to images/image selection during or shortly after the examination.\nEdit run name.\nDefine/set annotations on X-ray images, screenshots and reconstruction images.'], ['New', 'N.A.']]), ('text', '\xa0'), ('text', '')]
UNS.SmartSuite.AuthorizedUser | ['SRS.Security.PHD.PhysicalAccess', 'SRS.Security.MD.PhysicalAccess'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, CR>\nUse the system with or without authorization log on;\nLimit/restrict the use of the system by users;\nIn case of an emergency, unblock the use of primary system functionality by any user.'], ['New', 'N.A.']]), ('text', '\xa0'), ('text', 'Patient administration'), ('text', ' ')]
UNS.SmartSuite.AutomaticPositionControl | ['SRS.Allura.Func.APCTable', 'SRS.Allura.Func.SystemCustomization', 'SRS.Allura.Func.APCPathway', 'SRS.Allura.Func.APCPredefined', 'SRS.Allura.Func.APCCommon', 'SRS.Allura.Func.APCReference', 'SRS.Allura.Func.ModifyEPXParameters', 'SRS.Allura.Func.APCStore/Recall', 'SRS.Allura.Func.APC3DReference'] | [('text', ''), ('table', [['Legacy\n\n', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, ER/CR\xa0>\nStore and recall one table position;\nAutomatically re-acquire the exact same projection as the reference image currently displayed;\nSwitch between defined projections, mostly of different views of the same body part;\nCustomize and store pre-defined projections and sequences as part of the EPX, but without going to the EPX level;\nStore temporally and recall up to 20 projections of the system during a procedure (APC Store/Recall);\nMove the geometry to predefined projections (patient oriented) and along a sequence of projections defined at EPX level (APC Sequence);\nAutomatically set back the position of the system (beam carrier(s) and table) to the same position that was used during live, acquisition, reference images and/or manual storage, including the detector orientation, FOV, shutters and wedge positions;\nMove the position of the system (rotation and angulation) to the patient oriented projection defined in the interventional tools;\nAddition to APC with an APC along a path of positions of the longitudinal, transversal and z-rotation axes and of the table for the FlexArm configuration.'], ['New', 'N.A.']]), ('text', ''), ('text', ' ')]
UNS.SmartSuite.AuxiliarySystems | ['SRS.Allura.Intgr.Flexvision', 'SRS.Allura.Intgr.Flexspot.KeyboardMouse', 'SRS.Allura.Intgr.Interfaces.ControlAuxiliarySystems', 'SRS.Allura.Intgr.Interfaces.RealTimeImageLink', 'SRS.Allura.Func.SetInjectorCoupling', 'SRS.Allura.Intgr.OpenVideoInterface.LargeScreenDownscaled', 'SRS.Allura.Conf.RemainingComponents', 'SRS.Allura.Intgr.Flexspot', 'SRS.Allura.Intgr.VideoDisplayInterface', 'SRS.Allura.Intgr.Magnus', 'SRS.Allura.Intgr.OpenVideoInterface.LargeScreen', 'SRS.Allura.Intgr.Interfaces.DataExchange', 'SRS.Allura.Intgr.TruSystem', 'SRS.Allura.Intgr.Interfaces.AuxiliarySystemMounting', 'SRS.Allura.Intgr.SwitchableMonitors', 'SRS.Allura.Intgr.DoseAwareXtend', 'SRS.Allura.Intgr.KVM-Switching', 'SRS.Allura.Conf.MCS-RelatedItems', 'SRS.Allura.Conf.Table-RelatedItems', 'SRS.Allura.Intgr.FHDVideoSources', 'SRS.Allura.Conf.XperFDSeriesConfigurations', 'SRS.Allura.Func.PhysioAcquisition', 'SRS.Allura.Intgr.Interfaces.ControlFromAuxiliarySystems', 'SRS.Allura.Intgr.VideoSlave', 'SRS.Allura.Intgr.PhysioEquipment(ECG)', 'SRS.Allura.Intgr.ImageExportImport', 'SRS.Allura.Intgr.Flexspot.AdditionalWorkspot', 'SRS.Allura.Intgr.4KVideoInputOutput', 'SRS.Allura.Intgr.AblationEquipment', 'SRS.Allura.Intgr.Injectors', 'SRS.Allura.Conf.UI-RelatedItems', 'SRS.Allura.Intgr.Multivision', 'SRS.Allura.Intgr.Flexvision.KeyboardMouse', 'SRS.Allura.Intgr.EquipmentRack'] | [('text', '\n'), ('table', [['Legacy\n\n', 'The user wants to:\n<All users, all clinical areas, CR/ER>\nIntegrate and/or connect major brands of peripheral equipment and information systems;\nConnect the viewing system to a wide variety of third party systems (e.g. surgical knife, injector, video screens);\nView and interact with intra-vascular ultrasound (IVUS) images from table side;\nView and analyse images and volumes, originating from auxiliary systems, together with the images generated by the system;\nTransfer images and volumes, created by the system, to the auxiliary systems, as input for additional processing by these auxiliary systems;\nPosition catheters, as well as display and analyse catheter measurements;\nAcquire and display physiological signals;\nUse the system in a Hybrid OR environment by:\ncombining a special hybrid OR patient table with the system\ncreating a clean ceiling layout with the ability to position the equipment;\nCouple an injector to align (timing) the contrast injection with the runs;\nControl the injection of contrast fluid.\n'], ['New', 'Have flexibility in where to connect external equipment (e.g. different corners in the room, on the MCS);\nAlign (timing) the contrast injection with the runs.\n']]), ('text', '\n')]
UNS.SmartSuite.BiplaneContrast | ['SRS.Allura.Func.ContrastInjection', 'SRS.Allura.Func.SetInjectorCoupling', 'SRS.Allura.Func.BiplaneAcquisition'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'N.A.'], ['New', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, ER/CR>\nPerform imaging from two angles simultaneously to obtain an impression of the anatomy with a single contrast injection.']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.BiplaneFocusSelection | ['SRS.Allura.Func.SelectActiveChannel', 'SRS.Allura.Func.BiplaneAcquisition'] | [('text', ''), ('table', [['Legacy\n\n', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, ER>\nManually change the focus ( frontal, lateral or biplane) for imaging and acquisition, if the automatic focus selection does not apply.'], ['New', 'N.A.']]), ('text', ''), ('text', '')]
UNS.SmartSuite.BiplaneHeadSpace | ['SRS.Allura.Func.Z-RotateStandBase(L-ArmRotation)'] | [('text', 'Background: the -1350\xa0Rz movement of the frontal Arc allows the user to free up the head space with a biplane system. freeing up the headspace is for example relevant for the anesthesiologist in\xa0 order to monitor and respond to the patient’ situation adequately.'), ('table', [['Legacy\n\xa0', 'N.A.'], ['New', 'The user needs to be able to:\n< Clinical users/Physicists/ Service users, Neuro, ER>\nFree up the head space in a biplane system.']]), ('text', '\xa0'), ('text', 'Start examination'), ('text', ' ')]
UNS.SmartSuite.BiplaneRealTime | ['SRS.Allura.Conf.UI-RelatedItems', 'SRS.Allura.Func.ActivateExposure', 'SRS.Allura.Func.BiplaneAcquisition', 'SRS.Allura.Func.SingleShotFootpedal', 'SRS.Allura.Func.AcquisitionMethods', 'SRS.Allura.Func.SelectActiveChannel', 'SRS.Allura.Func.StoreFluoro', 'SRS.Allura.Func.ActivateFluoro'] | [('text', 'Background:\xa0in order to navigate interventional tools to, and at the region of interest, the clinical user wants to image from multiple angles to get an impression of the 3D anatomy in real-time. This is for example relevant in cases with torturous vessel structure (e.g. neurology case). '), ('table', [['Legacy\n\xa0', 'N.A.'], ['New', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, ER/CR>\nPerform imaging from two angles simultaneously to obtain multi-dimensional impression of the anatomy in real-time.']]), ('text', '\xa0')]
UNS.SmartSuite.BolusChaseReconstruction | ['SRS.Allura.Func.BolusChaseReconstruction'] | [('text', ''), ('table', [['Legacy\n\n', 'The user wants to:\n<All users, IR, ER/CR>\nAutomatically reconstruct Bolus Chase runs immediately after acquisition in order to create an overview of the peripheral arteries in one image;\nExport images to the hospital information system.'], ['New', 'N.A.']]), ('text', ''), ('text', 'Reporting'), ('text', ' ')]
UNS.SmartSuite.CardiacSwing | ['SRS.Allura.Func.Dual-AxisAcquisition'] | [('text', ''), ('table', [['Legacy\n\n', 'The user wants to\n<Clinical users, interventional cardiology, ER>:\nAcquire images of the patient’s anatomy from multiple views in a single run by rotating the system along curved trajectories around the patient;\nRotate the system along optimized predefined trajectories for adequate image quality;\nRotate the system along dedicated trajectories for the left and the right coronary arteries;\nRotate the system while staying automatically within safe boundaries in order to avoid any collisions between the patient/bystander and System (‘BodyGuard’).'], ['New', 'N.A.']]), ('text', ''), ('text', ' ')]
UNS.SmartSuite.Carriage | ['SRS.Allura.Conf.Stand-RelatedItems'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0\n\xa0', 'The user wants to:\n<Clinical users/Physicists/ Service users, all clinical areas, ER>\nTo select best treatment position and pass the table at head end from nurse to doctor side and vice versa with the FlexArm system.'], ['New', 'N.A.']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.CeilingCompatibility | ['SRS.Allura.Conf.Stand-RelatedItems'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0\n\xa0', 'The user wants to:\n<Clinical users/Physicists/ Service users, all clinical areas, ER>\nInstall the system in existing rooms;\nReuse the existing Allura/Azurion family ceiling construction as much as possible to avoid reconstruction and save cost.'], ['New', 'N.A.']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.ChangeDetectorOrientation | ['SRS.Allura.Func.ImageBeamRotation', 'SRS.Allura.Func.RotateFDXDDetector'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, ER>\nImage in portrait and landscape detector position independent of the gantry position;\nAlign the detector to the anatomy of the patient independent of the gantry position*.\n*this option is specific to the FlexArm and Azurion 7 B20/15LN.\n'], ['New', 'N.A.']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.ChangeFluoroExposure | ['SRS.Allura.Conf.UI-RelatedItems', 'SRS.Allura.Func.ActivateExposure', 'SRS.Allura.Func.ContrastInjection', 'SRS.Allura.Func.SetInjectorCoupling', 'SRS.Allura.Func.SingleShotFootpedal', 'SRS.Allura.Func.AcquisitionMethods', 'SRS.Allura.Func.SelectActiveChannel', 'SRS.Allura.Func.StoreFluoro', 'SRS.Allura.Func.ActivateFluoro'] | [('text', ''), ('table', [['Legacy', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, ER> \nChange the X-ray protocol during exposure and fluoroscopy.'], ['New', 'N.A.']]), ('text', ' ')]
UNS.SmartSuite.Cleaning | ['SRS.Allura.Phys.Cleanability', 'SRS.Allura.SLS.patient-contamination', 'SRS.Allura.UI.Help', 'SRS.Allura.UI.Help.Customize'] | [('text', '\n'), ('table', [['Legacy\n\n', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, ER>\nClean/disinfect system parts in the immediate vicinity of patients;\nEasily clean the cable covers.\nReceive or have direct access to instructions for appropriate cleaning/disinfection regimes of system parts (e.g. procedure, frequency, materials, and fluids).'], ['New', 'N.A.']]), ('text', ''), ('text', ''), ('text', ''), ('text', 'System shut-down'), ('text', ' ')]
UNS.SmartSuite.ClinicalImagesOnTSM | ['SRS.Allura.UI.TSM.Image'] | [('table', [['Legacy\n\n', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, ER/CR >\nProcess and view images on the TSM*);\nNavigate images on TSM without using remote control;\nZoom/pan images on the TSM without using remote control;\nReset the zoom/panning level on TSM;\nDisplay semi-live the acquired X-ray, IVUS and 3D (SmartCT) images on the TSM.\n\n*) Image display on TSM is not for diagnostic purposes.'], ['New', 'N.A.']]), ('text', ''), ('text', ' ')]
UNS.SmartSuite.CollisionDetection | ['SRS.Allura.SLS.collision-avoidance', 'SRS.Allura.SLS.unexpected-manual-behavior', 'SRS.Allura.Intgr.Magnus', 'SRS.Allura.SLS.movement-integrity', 'SRS.Allura.SLS.careless-manual-operation', 'SRS.Allura.Intgr.TruSystem', 'SRS.Allura.SLS.collision-unintend-ui', 'SRS.Allura.SLS.collision-harm-reduction', 'SRS.Allura.Func.OverrideMovementsBlockage'] | [('table', [['Legacy\n\n', 'The user needs to be able to:\n< Clinical users/Physicists/ Service users, all clinical areas, ER >\nRely on a collision avoidance system that:\nMaintains rotation and angulation speeds until it comes close to the patient;\nIs compatible with released Philips native tables and released integrated ORT solutions.'], ['New', 'Rely on a collision avoidance system that:\nPrevents collision with the patient;\nAllows the user to override the stop and continue at a slow speed in close proximity to the patient and/or system parts.']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.CompatibilityLAF | ['SRS.Allura.Phys.LaminarAirflowCompatibility'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0\n\xa0', 'The user wants to:\n<Clinical users/Physicists/ Service users, all clinical areas, ER>\nFit FlexArm in ORs offering a free floor and such that Raumklasse 1A (DIN 1946) can be established in an examination room with a Laminar Air Flow system with the FlexArm in parking position and in working position (the latter preferred, but not required);\nThe user should not experience any jet stream when carriage moves under LAF area.'], ['New', 'N.A.']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.CompensatePatientMovements | ['SRS.Allura.Func.NormalSubtraction', 'SRS.Allura.Func.DICOM.Export', 'SRS.Allura.Func.Roadmap.AutomaticPixelShift', 'SRS.Allura.Func.X-rayImageProcessing', 'SRS.Allura.Func.ExposureSubtract'] | [('text', ''), ('table', [['Legacy\n\n', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, ER/CR>\nPerform diagnosis on patients even when small patient and/or table unintended movements occur during acquisition of images (automatic image correction in case of small patient/table movements).\nExport motion correction parameters to peripheral systems (such as PACS, workstations);\nDo/undo automatic motion correction (toggle) to check and set quality of result.'], ['New', 'N.A.\xa0\xa0']]), ('text', '\n')]
UNS.SmartSuite.ControlModule | ['SRS.Allura.Conf.UI-RelatedItems', 'SRS.Allura.UI.Workspot', 'SRS.Allura.Func.AngulateBeam', 'SRS.Allura.UI.UIModule', 'SRS.Allura.SLS.ListOfApplicableStandards', 'SRS.Allura.Func.RotateBeam'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'The user wants to (at table side):\n< Clinical users/Physicists/ Service users, all clinical areas, ER>\nProtect the system from ingress of fluids (e.g. IPX4);\nControl the system at table side using a compact, simple and waterproof TSO, which allows for sufficient patient access and workspace.\nHave visibility and guidance in the dark (e.g. backlighting);\nMove the Control Module from the pedestal to the table and vice versa.'], ['New', 'N.A.']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.ControlRoomDesign | ['SRS.Allura.Intgr.Flexspot', 'SRS.Allura.Intgr.AuxiliarySystems', 'SRS.Allura.UI.KeyboardMouse', 'SRS.Allura.Intgr.Flexspot.KeyboardMouse'] | [('text', '\xa0'), ('table', [['Legacy', 'The user wants to:\n<Clinical users/Physicists/ Service users, all clinical areas, CR>\nLook over the top of screens in control room to see the exam-room;\nMinimize the amount of monitors;\nOperate different system applications with harmonized keyboard and mouse interfaces.'], ['New', 'Have the same monitor configuration in the control room for the biplane and monoplane systems.']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.CreateWorklist | ['SRS.Allura.SLS.IHE', 'SRS.Allura.Func.AddStudy', 'SRS.Allura.Func.ShowStudyInfo'] | [('text', ''), ('table', [['Legacy\n', 'The user wants to:\n<All users,\xa0\xa0all clinical areas, CR>\nCreate a work list of scheduled patients and associated procedures;\nManually create a work list and add patients;\nAutomatically create a work list when the system is connected to the hospital information system (e.g. RIS/CIS).'], ['New', 'N.A.']]), ('text', ''), ('text', ' ')]
UNS.SmartSuite.DataHandler | ['SRS.Allura.Func.CopyImages', 'SRS.Allura.Func.ModifyStudy', 'SRS.Allura.Func.DICOM.Export', 'SRS.Allura.Func.ShowStudyInfo', 'SRS.Allura.Func.RunLogging', 'SRS.Allura.Func.StoreToMedia', 'SRS.Allura.Func.DisplayStudyList'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'The user wants to:\n<All users, all clinical areas, CR/ER>\nSave reference images in specific destinations (e.g. reference file, Photo file, PACS, Information Systems) and in various standard formats.\nView/display small pictorials of image series (e.g. by using a data handler) coming from previous examinations (e.g. stored on the system database or retrieved from PACS) or the current examination;\nView/display small pictorials of image series arranged by filter settings (e.g. middle image;\nView/display relevant details of image series (e.g. pictorial pop-up);\xa0\nEasily access applications (e.g. Measurements, QA, BCR) from the DataHandler;\nCreate/edit names for a single or for multiple fluoroscopy/exposure run(s).\nHave the last used series filter settings persistently available.'], ['New', 'N.A.']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.Decommission | ['SRS.Allura.UI.Help', 'SRS.Allura.Serv.FSInstructions'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'N.A.'], ['New', "The user needs to:\n<FSE/RSE/BM, Deinstallation>\nDecommission* the system in the image-guided lab.\n\xa0\n*, e.g. remove from PRS, data destruction, 'un-setting to work' of the system, Philips Remote Services, etc."]]), ('text', '\xa0'), ('text', 'Packing and disposal'), ('text', ' ')]
UNS.SmartSuite.DeleteRun | ['SRS.Allura.Func.DeleteRun'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'The user wants to:\n<Clinical users/Physicists/ Service users, all clinical areas, ER/CR>\nManually delete a complete run;\nPrevent accidental deletion of a complete run.'], ['New', 'N.A.']]), ('text', '\xa0'), ('text', 'Park system'), ('text', ' ')]
UNS.SmartSuite.DeleteStudy | ['SRS.Allura.Func.DeleteStudy', 'SRS.Allura.Func.Protect/UnprotectStudy'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'The user wants to:\n<Clinical users/Physicists/ Service users, all clinical areas, CR/ER>\nFree up system storage space by deleting a study/studies and all related data from the database;\nProtect a study from automatic deletion by the system;\nReceive a message when the respective study cannot be deleted (e.g. study is protected or tasks are pending).\n'], ['New', 'N.A.']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.Diagnosis | ['SRS.Allura.Serv.Logging', 'SRS.Allura.Qual.FSResponseTimes', 'SRS.Allura.Serv.SystemMonitoring&Alerting', 'SRS.Allura.Serv.NetworkDiagnosis', 'SRS.Allura.Serv.HardwareTests', 'SRS.Allura.Serv.RemoteConnectivityStatus', 'SRS.Allura.Serv.FunctionalTests', 'SRS.Allura.Serv.SystemInformation.Remote', 'SRS.Allura.Serv.UserInterface', 'SRS.Security.MD.RemoteAccessPoint', 'SRS.Allura.Serv.RemoteSFTP', 'SRS.Allura.Serv.Maintenance.Predictive', 'SRS.Allura.Serv.SystemInformation', 'SRS.Allura.Intgr.Interfaces.RemoteConnectionSettings', 'SRS.Allura.Serv.Adjust', 'SRS.Allura.UI.ServiceUILanguage'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'N.A.'], ['New', 'The user needs to:\n<FSE/RSE/CAS/BM, Corrective Maintenance>\nDiagnose problems on the system in the image-guided lab*.\nUse diagnosis to determine correct action to solve the problems.\n\xa0\n*All diagnostic capabilities and information that is needed to determine the corrective action to solve the (future) issue, e.g. diagnostic capabilities and information needed: correct actual lab configuration, fingerprint, logging, health status,image-guided system FRUs, tests, adjustments, EPX database, ServiceHub, heamo system, ultrasound (that can be moved to different lab), 3rd\xa0party devices, usage history (who used the device and for what reason), to enable condition/usage based preventive maintenance.']]), ('text', '\xa0'), ('text', 'Escalate'), ('text', ' ')]
UNS.SmartSuite.DicomNetworkExporting | ['SRS.Allura.Qual.ClinicalResponseTimes.Arch', 'SRS.Allura.Func.SystemCustomization', 'SRS.Allura.Func.Diagnostics', 'SRS.Allura.Func.Workflow.Flexibility', 'SRS.Allura.Func.FinishStudy', 'SRS.Allura.Intgr.ImageExportImport', 'SRS.Allura.Func.DICOM.Query', 'SRS.Allura.Func.DICOM.Export', 'SRS.Allura.Func.DICOM.Export/ImportJobManagement', 'SRS.Allura.SLS.IHE', 'SRS.Allura.Func.DICOM.Import', 'SRS.Allura.Func.DICOM.Consistency'] | [('table', [['Legacy\n\n', 'The user wants to:\n<All users, all clinical areas, CR>\nManually and automatically export images from the X-ray system to PACS and/or an external workstation in the hospital;\nAutomatically transfer patient data (e.g. images, annotations) to the ISP;\nSchedule automatic transfer of images.\nReproduce (grey levels, up/downscaling and (un)subtraction) exported images on a wide variety of image archives, such as external viewers as PACS.\nAutomatically transfer cardiac data to cardio-PACS (with different settings, e.g. downscale, processed) and radiology data to radiology PACS (with different settings, e.g. no downscale, unprocessed) to optimally support mixed procedures.\nConfigure transfer actions based on procedures used with different settings to destinations (workstations, PACS), customizable by the user.\nLimit/ restrict export settings configuration for unauthorized users;\nShow network and status of connected systems with modality related information (e.g. IP address, AE title).'], ['New', 'Group single images from a study and export as one series to a DICOM destination.\nAzurion reduces the DICOM export time from the system. The export time of both cardiac (multi-frame) and radiology data (single-frame) is at least 2x faster compared to Allura Xper FD R8.2*\nOptimized interpolation;\nIQ subtraction for paedriatic patients.\n*Multi-frame: export of 10 runs * 100 images of 1K. Single-frame: export of 40 images of 1K. ']]), ('text', ''), ('text', ' ')]
UNS.SmartSuite.DisplayConfiguration | ['SRS.Allura.Intgr.Flexvision', 'SRS.Allura.UI.TSM.Image', 'SRS.Allura.Intgr.Flexspot.KeyboardMouse', 'SRS.Allura.Intgr.OpenVideoInterface.LargeScreenDownscaled', 'SRS.Allura.UI.PositionMCSFrame', 'SRS.Allura.Func.ActivateFluoro', 'SRS.Allura.Intgr.Flexspot', 'SRS.Allura.Intgr.VideoDisplayInterface', 'SRS.Allura.Intgr.3rdAnd4thMultiView', 'SRS.Allura.Func.ActivateExposure', 'SRS.Allura.Intgr.OpenVideoInterface.LargeScreen', 'SRS.Allura.SLS.ListOfApplicableStandards', 'SRS.Allura.Intgr.SwitchableMonitors', 'SRS.Allura.Func.Roadmap.DevicePhase.VesselSubtract', 'SRS.Allura.Func.Roadmap.DevicePhase.VesselUnSubtract', 'SRS.Allura.Func.NavigateToImageData', 'SRS.Allura.Conf.MCS-RelatedItems', 'SRS.Allura.Func.Roadmap.DevicePhase.LiveSubtract', 'SRS.Allura.Intgr.VideoSlave', 'SRS.Allura.Intgr.AdditionalMultiView', 'SRS.Allura.Func.Select/DeselectStopwatch', 'SRS.Allura.Conf.Control-RoomDisplays', 'SRS.Allura.Intgr.Flexspot.AdditionalWorkspot', 'SRS.Allura.Func.DisplayInfo', 'SRS.Allura.Func.Roadmap.DevicePhase', 'SRS.Allura.Intgr.Multivision', 'SRS.Allura.Intgr.Flexvision.KeyboardMouse', 'SRS.Allura.Intgr.MultiView'] | [('text', ''), ('table', [['Legacy\n\n', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, ER/CR>\nTo see all relevant images for his tasks from his/her working position;\nIn case of FlexVision (Pro), choose screen size (heamo live screen needs to remain in original resolution);\nView at selected screen location without having to lean towards the screen (images and UI should be clearly visible and readable);\nSlave screens for monitoring purposes;\nSee a source on multiple monitors;\nConnect 3rd party screens;\nConnect the system to 3rd party video switching and streaming solutions;\nActivate or de-activate a general purpose on-screen stopwatch.'], ['New', 'Display and control up to 8 sources (one at a time) on the additional MultiSwitch without the dependency on the physical connection to the KVM switch;\nIncrease switching flexibility when compared to the current DVI infrastructure, in the exam room and control room by enabling a flexible division between inputs, outputs and displays, with the routing flexibility of the network switch;\nEnable multiple signal layouts from external sources.\nMake sure some sources are always displayed (e.g. always see the live screen(s) and live heamo screen);\nChoose which image (live, ref) is shown on which screen in the exam room.\nEnables multiple signal layouts accordingly to the user preferences, on systems with FlexVision (Pro) and FlexSpot, due to the ability to connect up to 20 external video signals (not generated directly by the X-Ray system);\nSystems with 16 Switchable Monitors can connect up to 20 external video signals (not generated directly by the X-Ray system);\nSystems with FlexSpot can have up to 3 integrated workstations.']]), ('text', ''), ('text', ' ')]
UNS.SmartSuite.DoorContact | ['SRS.Allura.Func.DoorContact'] | [('text', ''), ('table', [['Legacy\n\n', 'The user wants to:\n<All users, all clinical areas, CR>\nInhibit X-ray from the system while the door is open between Exam Room and adjacent rooms.'], ['New', 'N.A.']]), ('text', ''), ('text', ' ')]
UNS.SmartSuite.DoseControl | ['SRS.Allura.Conf.UI-RelatedItems', 'SRS.Allura.Func.Move/ResetWedges', 'SRS.Allura.Func.FluoroFlavours', 'SRS.Allura.Func.Move/ResetShutters', 'SRS.Allura.Conf.RemainingComponents', 'SRS.Allura.Func.DisplayAndReportPatientDose'] | [('text', '\xa0'), ('table', [['Legacy', 'N.A.'], ['New', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, ER/CR>\nChoose fluoroscopy flavours (e.g. Low, Medium, Normal) and different acquisition settings for same procedure (low vs high dose, low vs high quality) to adjust image quality;\nPerform acquisition with As-Low-As-Reasonably-Achievable (ALARA) X-ray dose to patient and staff, i.e. optimize balance between radiation dose and IQ by adjusting relevant settings (e.g. ProcedureCard, DAP, frame rate, FoV).\nActivate acquisition from all sides of table (head, side (left/right), arm, feet side, remotely (trolley, control room);\nOperate system at a safe distance/shielded (with regard to user exposure dose) during fluoroscopy and acquisition;\nSee the pre-acquisition time;\nConnect multiple footswitches and hand switches in ER/CR to the system; receive dose information; both fluoroscopy and exposure.']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.DoseReporting | ['SRS.Allura.Func.DICOM.DoseStructuredReport', 'SRS.Allura.Func.Reporting.Dose', 'SRS.Allura.Func.DICOM.Export', 'SRS.Allura.Func.DICOM.Import', 'SRS.Allura.Func.StoreToMedia'] | [('text', ''), ('table', [['Legacy\n', 'The user wants to:\n<All users, all clinical areas, ER/CR>\nCreate/obtain\xa0a report with dose data:\nCumulative AK in skin area (in case of Cardiac/Thorax zones)\nCumulative Air Kerma (AK)\nCumulative DAP\nTotal Fluoro Time\nSave/ export dose data (e.g. by using DICOM Radiation Dose Structured Report);\nSave/export the report on/to the information system or an external destination in a commonly used\xa0file formats.'], ['New', 'N.A.']]), ('text', ''), ('text', ' ')]
UNS.SmartSuite.DualFluoroscopy | ['SRS.Allura.Func.AcquisitionZoom', 'SRS.Allura.Func.DualFluoro', 'SRS.Allura.Func.StoreReferenceRun/Image'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, ER/CR>\nView the fluoroscopy images subtracted and un-subtracted simultaneously (‘Dual Fluoroscopy’);\nFor biplane; View one reference image per channel at the same time with the live Frontal & Lateral image next to dual fluoroscopy;\nDisplay subtracted and un-subtracted exposures simultaneously;\nDigitally zoom in/out the subtracted and un-subtracted live image during Dual Fluoroscopy;\nView live image, reference image and Dual Fluoroscopy simultaneously.'], ['New', 'N.A.']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.DynamicRotationalAngiography | ['SRS.Allura.Func.DRAFor3D', 'SRS.Allura.Func.DualPhaseXperCT', 'SRS.Allura.Func.ClassicDRA', 'SRS.Allura.Func.XperCT'] | [('table', [['Legacy\n\n', "The user needs to be able to:\n< Clinical users, all clinical areas , ER/CR>\nGain 3-D insight of the patient’s anatomy (by making a dynamic rotational angiography scan where the X-ray beam rotates around the patient's area of interest placed within the iso-center);\nControl a rotational scan from all table sides/pedestal (head, foot, nurse, physician, specific);\nMake a rotational scan with a wide range of motion to cover the complete patient’s anatomy (from head to toe);\nMake a rotational scan with a wide range of motion to cover the complete anatomy of obese (BMI>30) patients."], ['New', 'N.A.']]), ('text', '\n')]
UNS.SmartSuite.ECGTrigger | ['SRS.Allura.Intgr.PhysioEquipment(ECG)', 'SRS.Allura.Func.EPTriggering'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'The user wants to:\n< Clinical users/Physicists/ Service users, PCI/EP, ER/CR>\nSynchronize fluoroscopy and acquisition timing with patient’s ECG signal (e.g. ‘ECG triggering’\xa0to acquire images in the same phase of the heart cycle).'], ['New', 'N.A.\xa0\xa0']]), ('text', '\xa0'), ('text', '')]
UNS.SmartSuite.EMFieldImageFiltering | ['SRS.Allura.Func.StripeFilter'] | [('text', ''), ('table', [['Legacy\n\n', 'The user wants to:\n< Clinical users/Physicists/ Service users, EP, ER>\nApply fluoroscopy and acquisition without artefacts that are introduced by electromagnetic tracking devices.'], ['New', 'N.A.']]), ('text', '\n')]
UNS.SmartSuite.EPXtree | ['SRS.Allura.Func.PresettingFunctions', 'SRS.Allura.Func.ProcedureCard', 'SRS.Allura.Func.EPX', 'SRS.Allura.Func.ModifyEPXParameters', 'SRS.Allura.Func.ModifySelectionTree'] | [('text', '\xa0'), ('table', [['Legacy', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, CR/ER>\nPerform a simple procedure selection, coupled to the selection of the ProcedureCard.'], ['New', 'Perform a simple procedure selection, outside the content of a\xa0the ProcedureCard.']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.EasyUSBExport | ['SRS.Allura.Intgr.Media', 'SRS.Allura.Func.Import/ExportOfClinicalImages', 'SRS.Allura.Func.StoreToMedia'] | [('text', ''), ('table', [['Legacy\n\n', 'The user wants to:\n<Clinical users/Physicists/ Service users, all clinical areas, CR>\nExport image data through a USB connected destination;\nTo select the format (DICOM or standard multi-media formats) and anonymize the data before exporting;\nTo include a DICOM viewer to the exported data;\nGive the exported data a new name.'], ['New', 'N.A.']]), ('text', ''), ('text', ' ')]
UNS.SmartSuite.Emergency | ['SRS.Allura.Func.ImplicitSelectionDefaultStudy', 'SRS.Allura.Func.TableLock', 'SRS.Allura.Func.MoveBeam', 'SRS.Allura.Qual.ClinicalResponseTimes.System', 'SRS.Allura.Func.EmergencyPowerDown', 'SRS.Allura.Func.EmergencyStop', 'SRS.Allura.Func.MoveBeamXY'] | [('text', ''), ('table', [['Legacy\n\n', 'The user needs to be able to:\n<All users, all clinical areas , ER/CR>\nIn case of an emergency, unblock the use of primary system functionality by any user;\nOperate the system without entering a user name and password;\nStarting imaging without having entered any patient data (e.g. set-up one of the X-ray protocols as default for emergency cases);\nManually move the stand in XY out of way in emergency and/or power off situation;\nAccess a system emergency stop from the ER;\nPerform CPR on the table without table top floating (lock table);\nFree up the C-arm space around the patient manually, even in case of a power down (excluding 15”and 20” detector floor system).'], ['New', 'N.A.']]), ('text', ''), ('text', 'Service'), ('text', 'Service workflows'), ('text', ' '), ('text', 'There are five service activities (installation, corrective maintenance, preventive maintenance, upgrade/field change order, and deinstallation) that each have a workflow that represents the different service activities that are performed by users from the service user group during the system life cycle. Figure 1 shows the complete set of tasks in the workflows. Most tasks have to be done in multiple service workflows; each unique task has the same color.'), ('text', ' \n'), ('text', ''), ('text', 'Fig. 1:\xa0Service workflows.'), ('text', 'Service activities'), ('text', ' '), ('text', 'This section summarizes the unique user needs that can be encountered in the service workflows.\xa0For all tasks related to the installation activity the user should have all the information that is needed to ensure that the right preparations for the tasks can be made such as mains, network connections, remote connectivity, hospital IT impact, building construction requirements, etc. Also, other systems in the hospital need to be prepared by their provider to work correctly with the Philips Image Guided System e.g. heamo system, ultrasound or 3rd\xa0party devices.'), ('text', 'Site planning'), ('text', ' ')]
UNS.SmartSuite.EndExamination | ['SRS.Allura.Func.ReopenFinishedStudy', 'SRS.Allura.Intgr.Flexspot.AdditionalWorkspot', 'SRS.Allura.Func.FinishStudy', 'SRS.Allura.Func.SelectStudyFromSuspendedList'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'The user wants to:\n<Clinical users/Physicists/ Service users, all clinical areas, ER/CR>\nSuspend an examination (or intervention) on one patient, continuing with an examination on another patient, and resume the examination on the previous patient a later moment in time;\nClose the Exam in the ER from the Additional FlexSpot.'], ['New', 'N.A.']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.EnvironmentalCare | ['SRS.Allura.SLS.ECO.ProductLabeling', 'SRS.Allura.SLS.ListOfApplicableStandards'] | [('text', '\n'), ('table', [['Legacy\n\n', 'The user wants to:\n<All users, all clinical areas, ER/CR>\nUse a system that complies with the international environmental standards.'], ['New', 'N.A.']]), ('text', '\n'), ('text', ' ')]
UNS.SmartSuite.Escalation | ['SRS.Allura.Serv.SystemInformation.Remote', 'SRS.Allura.Serv.RetrieveInformation', 'SRS.Allura.Serv.RemoteAccessAndConnection', 'SRS.Allura.Serv.RemoteSFTP', 'SRS.Allura.Serv.Field-ServiceAccessPoints', 'SRS.Allura.Serv.RemoteImageTransfer', 'SRS.Allura.Serv.RemoteDesktop.ServiceMode'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'N.A.'], ['New', 'The user needs to:\n<FSE/RSE/CAS/BM, Corrective Maintenance>\nCollect data and send to Philips Remote Service to enable remote support to be able to escalate problems to a higher support level (tier-1 to 4).']]), ('text', '\xa0'), ('text', 'Mechanical deinstallation'), ('text', ' ')]
UNS.SmartSuite.ExaminationSelectionAndPresetting | ['SRS.Allura.Func.EPX', 'SRS.Allura.Func.PhysicianSpecificSettings', 'SRS.Allura.Qual.ClinicalResponseTimes.Admin', 'SRS.Allura.Func.PresettingFunctions'] | [('text', '\n'), ('table', [['Legacy\n\n', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, CR/ER>\nSelect the settings for an examination depending on procedure type, patient type and X-ray operator.\nChange the procedure type during the examination at tableside without leaving the ER.'], ['New', 'N.A.']]), ('text', '\n')]
UNS.SmartSuite.ExceptionHandling | ['SRS.Allura.Func.GracefulDegradation', 'SRS.Allura.SLS.xray-imaging-failure-detection-and-fault-tolerant', 'SRS.Allura.Func.DisplayUserMessages', 'SRS.Allura.SLS.movements-failure-detection-and-fault-tolerant', 'SRS.Allura.Func.ExceptionHandling', 'SRS.Allura.Func.TubeLoadProtectionAndIndication', 'SRS.Allura.Qual.ClinicalResponseTimes.System', 'SRS.Allura.SLS.maintain-minimal-functionality', 'SRS.Allura.Func.DisplayServiceGuidance'] | []
UNS.SmartSuite.ExportApp | ['SRS.Allura.Func.DICOM.Export', 'SRS.Allura.Func.Flagging'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'The user wants to:\n<All users, all clinical areas, CR>\nManage images/runs selected for export (including selection at table side) at table side;\nExport selected images/runs to DICOM destinations;\nExport in original raw format, in DICOM processed/unprocessed format;\nExport imported data.'], ['New', 'N.A.']]), ('text', '\xa0'), ('text', '')]
UNS.SmartSuite.EyeHandCoordination | ['SRS.Allura.Qual.ClinicalResponseTimes.Acquisition', 'SRS.Allura.SLS.image-latency'] | [('table', [['Legacy\n\n', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, ER/CR>\nPerceive that the fluoroscopy images and exposure are displayed in real time and without noticeable delay.'], ['New', 'N.A.']]), ('text', ''), ('text', ' ')]
UNS.SmartSuite.FD12DetectorSize | ['SRS.Allura.Conf.XperFDSeriesConfigurations', 'SRS.Allura.Conf.SupportedConfigurations'] | [('text', '\n'), ('table', [['Legacy', 'The user wants to:\n< Clinical users/Physicists/ Service users, PCI/SHD/EP, ER>\nAchieve a larger field of view when compared to with 10”detectors:\nat a selected SID while maintaining projection flexibility;\nto support viewing of the aortic valve and a good part of the aortic arch, preferably in one view with minimal table panning, so that one can follow the device when navigating the aortic arch and see the valve at the same time;\nto support viewing of the whole coronary tree in one view with minimal need for table panning. This can be challenging especially during procedures of patient with a dilated heart (cardiac myopathy);\nto view both the needle insertion point (subclavian or axillary vein) and the whole wire/lead towards the right atrium in one view with minimal table panning and C-arm repositioning. This can be challenging especially for patient with a dilated heart and left ventricle;\nSee small details by having a small pixel size, resulting in high limiting resolution. '], ['New', 'N.A.']]), ('text', '')]
UNS.SmartSuite.FD15DetectorSize | ['SRS.Allura.Conf.XperFDSeriesConfigurations', 'SRS.Allura.Conf.SupportedConfigurations'] | [('text', '\xa0'), ('table', [['Legacy', 'The user wants to:\n<Clinical users/Physicists/ Service users, all clinical areas, ER>\nAchieve a 154% larger field of view than the 10” detector, to support viewing of larger vascular structures in routine procedures such as renal artery procedures, carotid artery procedures or iliac procedures.'], ['New', 'N.A.']]), ('text', '\xa0')]
UNS.SmartSuite.FD20DetectorSize | ['SRS.Allura.Conf.XperFDSeriesConfigurations', 'SRS.Allura.Conf.SupportedConfigurations'] | [('text', '\xa0'), ('table', [['Legacy', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, ER>\nMake images up to 48cm diagonal.'], ['New', 'N.A.']]), ('text', '\xa0')]
UNS.SmartSuite.FDPA | ['SRS.Allura.Func.BolusChaseReconstruction', 'SRS.Allura.Func.FreeInteractiveFDPA', 'SRS.Allura.Func.RunSubtract'] | [('text', ''), ('table', [['Legacy\n\n', 'The user wants to:\n< Clinical users/Physicists/ Service users, IR, ER/CR>\nAcquire (subtracted and non-subtracted) images from the lower abdomen to the feet for showing a survey of the contrast bolus as it progresses over time (‘bolus tracking’);\nPerform a single injection run off in Flexible Dynamic Peripheral Angiography (FDPA):\nPerform multiple runs for subtraction;\nReview the whole peripheral vasculature in one image (e.g. image stitching);\nAcquire the FDPA images with the frontal stand or lateral stand in case of PolyG floor frontal stand;'], ['New', 'N.A.']]), ('text', '\n')]
UNS.SmartSuite.FactoriumLogging | ['SRS.Allura.Serv.Logging'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'The user wants to:\n<Clinical users/Physicists/ Service users, all clinical areas s, CR/ER>\nAutomatically log events that happened while operating the system;\nReuse this data for data-mining.'], ['New', 'N.A.']]), ('text', '\xa0'), ('text', 'Integration'), ('text', ' ')]
UNS.SmartSuite.Flagging | ['SRS.Allura.Func.Flagging'] | [('text', ''), ('table', [['Legacy\n\n', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, ER/CR>\nFlag images/runs and sort on flagged images;'], ['New', 'N.A.']]), ('text', ''), ('text', ' ')]
UNS.SmartSuite.FlexMove | ['SRS.Allura.Func.APCStore/Recall', 'SRS.Allura.Func.MoveBeamXY'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'The user wants to:\n< Clinical users/Physicists/ Service users, VASU/SHD, ER>\nMove the stand (instead of the table) in XY direction (lateral & longitudinal movements) to position for imaging;\nStore and recall multiple XY projections.'], ['New', 'N.A.']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.FlexMoveControl | ['SRS.Allura.Conf.Stand-RelatedItems', 'SRS.Allura.Conf.UI-RelatedItems', 'SRS.Allura.Func.MoveBeamXY'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0\n\xa0', 'The user wants to:\n<Clinical users/Physicists/ Service users, VASU/SHD, ER>\nControl combined XY movement with a single hand in an intuitive way;\nStay clear of the head-end area for anesthesia;\nMove from patient left to patient right with the stand, without moving anesthesia equipment;\nStay and work in between table and stand in lateral position;\nControl all geometry movements from a single control panel;\nOperate the system with a XY joystick with proportional speed-control in concurrent X-Y directions.'], ['New', 'N.A.']]), ('text', '\xa0'), ('text', 'Customization'), ('text', ' ')]
UNS.SmartSuite.FlexMoveDesign | ['SRS.Allura.SLS.patient-contamination', 'SRS.Allura.SLS.patient-contamin-airborne', 'SRS.Allura.Func.MoveBeamXY'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0\n\xa0', 'The user wants to:\n<Clinical users/Physicists/ Service users, VASU/SHD, ER>\nPrevent interference with monitors/lamps as much as possible (integrated cabling);\nObtain a high level of sterility/hygiene in a Hybrid OR (e.g. minimum open cabling at monitor connection);\nMinimize the interruption of laminar airflow caused by the system.'], ['New', 'N.A.']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.FlexMovePark | ['SRS.Allura.Func.ParkorStandbyBeam', 'SRS.Allura.Func.MoveBeamXY'] | [('text', ''), ('table', [['Legacy\n\n', 'The user wants to:\n<Clinical users/Physicists/ Service users, VASU/SHD, ER>\nPark the stand everywhere where the user wants to, within the system’s limitations.'], ['New', 'N.A.']]), ('text', '')]
UNS.SmartSuite.FlexSpot | ['SRS.Allura.Intgr.Flexspot.SecondFlexspot', 'SRS.Allura.Intgr.VideoDisplayInterface', 'SRS.Allura.Intgr.Flexspot', 'SRS.Allura.Intgr.VideoInputInterface', 'SRS.Allura.Intgr.Flexspot.KeyboardMouse', 'SRS.Allura.Intgr.OpticalImagingSystem', 'SRS.Allura.Conf.Control-RoomDisplays', 'SRS.Allura.Intgr.Flexspot.AdditionalWorkspot', 'SRS.Allura.Intgr.AuxiliarySystems', 'SRS.Allura.Intgr.VideoSourcesMaintainIQ'] | [('table', [['Legacy\n\n', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, CR/ER>\nSeamlessly control (view and manipulate) all viewports within a single view and with a single mouse and keyboard;\nQuickly access applications without the need to switch between applications and without different mouse control;\nFully customize the layout (e.g. which applications on screen, size of viewports, number of viewports);\nCreate presets, adjust these on-the-go and quickly save changes;\nOrganize presets in groups for easy storage and retrieval;\nDisplay any (3rd party) system;\nConnect at least 11 external video signals (that are not generated directly by the X-ray system);\nChoose from at least 4 different FlexSpot configurations to accommodate the hospital workflow (single work spot with single monitor, single work spot with two monitors, two work spots with one monitor each, two work spots (one with two monitors and the other with one monitor);\nReduce the footprint of display monitors in the Control Room;\nChange the layout of the FlexVision (Pro) from the Control Room;\nCreate a snapshot of a viewport/all displayed video source with one click, export it as secondary capture and store the snapshot in the patient file;\nVisualize the references monitors in separate window (parallel to live).'], ['New', 'Have flexibility in positioning workspots (e.g. between CR/ER and multiple workspots in the CR);\nLayouts can be changed during the procedure without having to go into the system settings;']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.FlexVision(Pro) | ['SRS.Allura.Intgr.Flexspot', 'SRS.Allura.Intgr.Flexvision', 'SRS.Allura.Intgr.VideoDisplayInterface', 'SRS.Allura.Intgr.VideoInputInterface', 'SRS.Allura.Intgr.Flexvision.KeyboardMouse', 'SRS.Allura.Intgr.Flexspot.KeyboardMouse', 'SRS.Allura.Intgr.OpenVideoInterface.LargeScreen', 'SRS.Allura.Intgr.Interfaces.ControlFromAuxiliarySystems', 'SRS.Allura.Intgr.OpticalImagingSystem', 'SRS.Allura.Intgr.Interfaces.ControlAuxiliarySystems', 'SRS.Allura.Intgr.Interfaces.RealTimeImageLink', 'SRS.Allura.Intgr.Flexvision.Snapshot', 'SRS.Allura.Intgr.OpenVideoInterface.LargeScreenDownscaled', 'SRS.Allura.UI.ComfortThemes', 'SRS.Allura.Intgr.NativeIQ', 'SRS.Allura.Intgr.VideoSourcesMaintainIQ'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'The user needs to be able to:\n< Clinical users/Physicists/ Service users, all clinical areas , ER>\nDefine and customize the ER monitor configuration and layout of images on screen;\nAdd and connect a second/backup screen for live/heamo when a single large monitor is used;\nSimultaneously see:\nLive X-ray and Ref X-ray in same size;\nHeamo in proper size, resolution, non-scaled, pixel for pixel;\nScaled 3rd\xa0party/Interventional Tools/QA input (optional and switchable, based on operator request);\nStart the system with Heamo in proper size/resolution/non scaled/pixel for pixel and have the option to scale the Heamo back and up again;\nSimultaneously see (optional) 3rd\xa0party monitors in addition;\nRescale/rezoom display container to its demand;\nSeamlessly control (view and manipulate) all viewports at table side with a single means / mouse;\nFully customize the layout (e.g. which applications on screen, size of viewports, number of viewports, position of the status bar) using the table side mouse, TSM or from the FlexSpot;\nCreate presets, adjust these on-the-go and quickly save changes;\nOrganize presets in groups for easy storage and retrieval;\nDisplay any (3rd party) system;\nConnect at least 11 external video signals (that are not generated directly by the X-ray system;\nView X-ray images in native resolution without downscaling;\nView external sources in their native resolution to prevent signal distortion due to scaling (up/down);\nCreate a snapshot of a viewport/all displayed video source with one click, export it as secondary capture and store the snapshot in the patient file;\nFully customize the layout (e.g. which applications on screen, size of viewports, number of viewports, position of status bar) using the table side mouse, TSM or from the FlexSpot;\nVisualize two full HD video sources in original pixel format, at the same time, mainly for EP procedures.'], ['New', 'Improved image quality when downscaling EP sources.']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.FlexibleAngles | ['SRS.Allura.Func.Z-RotateStandBase(L-ArmRotation)', 'SRS.Allura.Func.PatientOrientedMovement', 'SRS.Allura.Func.RotateFDXDDetector'] | [('text', ''), ('table', [['Legacy', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, ER>\nTo position FlexArm C-arc in any angle for 2D imaging between -135 and +135 degrees with respect to the standard table position to facilitate access to the patient'], ['New ', 'N.A.']]), ('text', ' ')]
UNS.SmartSuite.FlexibleViewing | ['SRS.Allura.Intgr.Flexvision', 'SRS.Allura.Intgr.Flexspot', 'SRS.Allura.Intgr.VideoDisplayInterface', 'SRS.Allura.Intgr.Flexvision.KeyboardMouse', 'SRS.Allura.Intgr.Flexspot.KeyboardMouse', 'SRS.Allura.Intgr.Flexspot.AdditionalWorkspot'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, CR/ER>\nSelect from the available image sources which ones will be displayed on which part of the screen (viewports); and save this display;\nChange the display layout and rearrange the position and size of multiple viewports; and save this display;\nSwap/move keyboard and mouse focus over selected viewports;\nEnable/disable whether the display will be visible in the ER or not;\nConnect a dedicated keyboard to an external workstation modality (e.g. Heamo and/or EP).'], ['New', 'Have flexibility in scaling the size of multiple sources;\nHave flexibility in displaying different sources on different monitors;\nShow multiple sources with different resolutions on Philips flexible monitors (e.g. Flexvision, Flexspot).']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.FluoroIndicators | ['SRS.Allura.Func.RoomLightAndOperationLamp', 'SRS.Allura.Func.FluoroTimeBuzzer', 'SRS.Allura.Func.IndicateRadiationOn'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'N.A.'], ['New', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, CR/ER>\nSwitch on/off the volume of sounds for acquisition and fluoroscopy;\nStop the 5 minute fluoro time buzzer; \nIndicate visually if actual X-ray is given or initiated e.g. “X-ray on” indicator light).']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.FreeEndPositions | ['SRS.Allura.Func.ImageBeamRotation', 'SRS.Allura.Func.RotateBeam'] | [('text', '\xa0'), ('table', [['Legacy', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, ER>\nMove the FlexArm system from/to working position without interference with equipment/persons located at all sides of the table (e.g. anaesthesiologists).'], ['New', 'N.A.']]), ('text', '\xa0'), ('text', 'Position patient'), ('text', ' ')]
UNS.SmartSuite.FromStandbyToWorkingPosition | ['SRS.Allura.Func.ParkorStandbyBeam', 'SRS.Allura.Func.XperCT', 'SRS.Allura.Func.DRAFor3D', 'SRS.Allura.Func.DualPhaseXperCT', 'SRS.Allura.Func.ClassicDRA'] | [('text', 'Background information: e.g. after a 3D acquisition, going to standby position and going back to working position.'), ('table', [['Legacy', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, ER>\nFor FlexArm configuration return to the last X-ray Position from a Standby Position for any procedure while maintaining previous registration on Philips clinical applications (e.g. Interventional Workspot, EchoNav and Coronary tools. Example 3D roadmap, VesselNavigator overlay. Assuming patient and table remain in same position).'], ['New ', 'Have the largest lateral standby distance on the market while maintaining 3D registration.']]), ('text', ' ')]
UNS.SmartSuite.FullBodyCoverage | ['SRS.Allura.Func.MoveBeamXY'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'The user wants to:\n< Clinical users/Physicists/ Service users, all\xa0clinical areas, ER>\nImage (2D) the entire body with FlexArm from foot to head from both nurse and doctor sides of the table, without the need to move the patient table combination*.\n\n*not applicable when using a leg support board'], ['New', 'N.A.']]), ('text', '\xa0'), ('text', 'Position system'), ('text', ' ')]
UNS.SmartSuite.HCA | ['SRS.Security.PHD.HospitalNetwork', 'SRS.Security.HN.RemovableMediaBreached', 'SRS.Security.PHD.PatientDataCorrupted', 'SRS.Security.MD.DataBreach', 'SRS.Security.MD.UnavailabilityOfMedicalDevice', 'SRS.Security.MD.HospitalNetwork', 'SRS.Security.PHD.PhysicalDismounting', 'SRS.Security.MD.MalwareOnRemovable media', 'SRS.Security.PHD.PatientDataLoss', 'SRS.Security.PHD.MedicalDevice', 'SRS.Security.PHD.PhysicalAccess', 'SRS.Security.HN.AttackViaMedicalDevice', 'SRS.Security.MD.AttackViaPhysical DeviceModification ', 'SRS.Security.PHD.RemoteAccessPoint', 'SRS.Security.MD.RemoteAccessPoint', 'SRS.Security.PHD.DataBreach', 'SRS.Security.MD.PhysicalAccess', 'SRS.Security.MD.ReputationDamage', 'SRS.Security.HN.HospitalNetworkBreached', 'SRS.Security.RM.AttackViaMedicalDevice'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'The user wants to:\n<All users, all clinical areas, ER/CR>\nUse a system that complies with the security requirements from HCA.\xa0'], ['New', 'N.A.']]), ('text', '\xa0'), ('text', ''), ('text', 'Annex A: Compatibility Overview'), ('text', ' '), ('text', 'Legacy Azurion will deliver compatibility statements for:'), ('text', 'Philips clinical applications (e.g. Interventional Workspot, EchoNav and Coronary tools)'), ('text', 'Philips Ultrasound systems; video display integration and table side control'), ('text', 'Philips heamo dynamic systems and applications; with table side controlPhilips information systems and multi-modality imaging clinical applications, with table side control'), ('text', 'Departmental Image Management Systems /PACS systems'), ('text', 'Measurement systems (e.g. 2D quantitative analysis)'), ('text', 'Export of Philips clinical applications'), ('text', 'Monitor booms'), ('text', 'Universal trolley'), ('text', 'Injectors'), ('text', 'OR tables'), ('text', 'Robotic steering systems'), ('text', 'Video distribution systems (2 ways)'), ('text', 'Video display integration systems, e.g. ImageStream'), ('text', 'EP recorders'), ('text', 'EP mapping, navigation and planning systems'), ('text', 'EM tracking systems'), ('text', 'PCI guidance systems (e.g. Corindus)'), ('text', 'Intravascular imaging systems, with table side control (e.g. TVC, IVUS, …)'), ('text', 'Neuro head holder'), ('text', 'Patient dose monitoring systems'), ('text', ''), ('table', [['\n\n'], [''], [''], ['Philips’ proprietary information. Unauthorized use is prohibited.']])]
UNS.SmartSuite.Handover | ['SRS.Allura.Serv.AcceptanceTestReport', 'SRS.Allura.Serv.SiteLog', 'SRS.Security.MD.RemoteAccessPoint', 'SRS.Allura.Serv.SystemInformation.Export', 'SRS.Allura.Serv.SystemID'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'N.A.'], ['New', 'The user needs to:\n<FSE/RSE/BM, all service workflows>\nHandover evidence indicating that the system is operational according to manufacture specification after any service activity;\nGet proof that the customer accepts the system in the image-guided lab.']]), ('text', '\xa0'), ('text', 'Administration'), ('text', ' ')]
UNS.SmartSuite.HospitalITConformance | ['SRS.Allura.Func.SystemCustomization'] | [('text', ''), ('table', [['Legacy\n\n', 'The user wants to:\n<Hospital IT, all clinical areas, CR/ER>\nTo bring the hospital infrastructure and associated network protocols up-to-date and connect the system to the updated configuration.'], ['New', 'N.A.']]), ('text', ' '), ('text', '')]
UNS.SmartSuite.IQ | ['SRS.Allura.Func.AcquisitionMethods', 'SRS.Allura.Func.X-rayImageProcessing'] | [('text', 'Background: the system should provide image quality that allows the user to perform procedures. This includes, but is not limited to, being able to see anatomy, vasculature, devices and catheters.'), ('table', [['Legacy', 'N.A.'], ['New', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, CR/ER>\nGet images that can be used by the physician to perform procedures.\n']]), ('text', '\xa0'), ('text', 'Review images'), ('text', ' ')]
UNS.SmartSuite.ISPWorkflowIntegration | ['SRS.Allura.Intgr.Interfaces.DataExchange'] | [('text', ''), ('table', [['Legacy\n\n', 'The user wants to:\n<All users, all clinical areas, CR/ER>\nExchange data between Azurion and IntelliSpace Portal.'], ['New', 'Automatically launch relevant images on IntelliSpacePortal* ;\nPrefetch of data upon selection of the patient acquisition on Azurion.\n*Available from ISP R8 onwards.\n']]), ('text', ''), ('text', ' ')]
UNS.SmartSuite.ImageEnhancement | ['SRS.Allura.Func.OverrideManualAdjustments', 'SRS.Allura.Func.VideoInvert', 'SRS.Allura.Func.SetImageEnhancement', 'SRS.Allura.Func.PanZoom', 'SRS.Allura.Qual.ClinicalResponseTimes.Review'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, ER/CR>\nAdjust image to enhance image review, e.g.:\nZoom/pan the displayed image;\nChange image contrast, brightness and/or amount of edge enhancement;\nInvert image polarity on display (e.g. invert black vessels to white).'], ['New', '\xa0N.A.']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.ImageOffCenter | ['SRS.Allura.Func.MoveBeamXY'] | [('text', '\xa0'), ('table', [['Legacy', 'The user wants to:\n< Clinical users/Physicists/ Service users, all disciplines, ER>\nPosition the FlexArm gantry centre laterally over at least 52 cm with respect to the table longitudinal centre line without moving staff or equipment which is positioned next to the table or the patient table combination;\nCreate workspace, at least 118 cm, between the FlexArm stand centre and table longitudinal centre line for example to stay between stand and table, maintaining sterility of the user.'], ['New', 'Perform radial access procedures on both sides of the table.\xa0']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.ImageOverlay | ['SRS.Allura.Func.PredefinedAnnotation', 'SRS.Allura.Func.OverrideManualAdjustments', 'SRS.Allura.Func.CO2Tracing', 'SRS.Allura.Func.VideoInvert', 'SRS.Allura.Func.NormalSubtraction', 'SRS.Allura.Func.ChangeFluoroProcessing', 'SRS.Allura.Func.Tracing', 'SRS.Allura.Func.DICOM.Export', 'SRS.Allura.Func.SetImageEnhancement', 'SRS.Allura.Qual.ClinicalResponseTimes.Review', 'SRS.Allura.Func.StoreToMedia', 'SRS.Allura.Func.Roadmap.DevicePhase.Processing', 'SRS.Allura.Func.ManualAnnotation', 'SRS.Allura.Func.SelectProcessingScope', 'SRS.Allura.Func.OverlayText', 'SRS.Allura.Func.SelectPhysioDisplay'] | [('text', '\n'), ('table', [['Legacy\n\n', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, CR>\nAdjust post-processing parameters of the image; at image level, at run level and for fluoroscopy;\nAdd portion of background information to processed image (e.g. Contrast Brightness and Edge enhancement (CBE), Landmark, ViewTraced, video invert etc.)\nAdd/remove/adjust/configure/view/save/export the following data/ information to images:\nRelevant physiology data (e.g. ECG signals);\nImage acquisition related information (e.g. detector size, run time stamp);\nAnnotations (predefined as well as added manually).\nCustomize the default information shown on the live and review image.'], ['New', 'N.A.']]), ('text', '\n')]
UNS.SmartSuite.ImageSubtraction | ['SRS.Allura.Func.Roadmap.DevicePhase.DeviceMask', 'SRS.Allura.Func.Roadmap.DevicePhase.LiveSubtract', 'SRS.Allura.Func.Roadmap.DevicePhase.Subtraction', 'SRS.Allura.Func.NormalSubtraction', 'SRS.Allura.Func.Roadmap.DevicePhase.VesselSubtract', 'SRS.Allura.Func.AcquisitionMaskAveraging', 'SRS.Allura.Qual.ClinicalResponseTimes.Acquisition', 'SRS.Allura.Func.RunSubtract', 'SRS.Allura.Func.ExposureSubtract'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'The user wants to\n<Clinical users/Physicists/ Service users, all clinical areas, ER/CR >\nMake vessels visible without disturbing background; at image level, at run level and during fluoroscopy;\nToggle between subtracted and un-subtracted display.'], ['New', 'N.A.']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.ImageTranfer | ['SRS.Allura.Func.ReopenFinishedStudy', 'SRS.Allura.Func.FinishStudy', 'SRS.Allura.Func.DICOM.Export', 'SRS.Allura.Func.DICOM.Export/ImportJobManagement', 'SRS.Allura.SLS.IHE', 'SRS.Allura.Func.StoreToMedia'] | [('text', ''), ('table', [['Legacy\n\n', 'The user wants to:\n<Clinical users/Physicists/ Service users, all clinical areas, CR>\nTransfer image data to multiple destinations in multiple formats;\nReview the status (e.g. scheduled, in-progress, completed), of image data transfer;\xa0\nReceive confirmation (e.g. successful, non-successful) of image data transfer;\nClose examination and receive notification of closure;\nRe-open and re-close a closed examination.'], ['New', 'N.A.']]), ('text', ''), ('text', ' ')]
UNS.SmartSuite.ImagingControl | ['SRS.Allura.Func.FluoroFlavours', 'SRS.Allura.Func.Move/ResetShutters', 'SRS.Allura.Func.SelectDetectorFieldsize', 'SRS.Allura.Func.EmergencyStop', 'SRS.Allura.Func.APCReference', 'SRS.Allura.Func.TableLock', 'SRS.Allura.Func.SwingStandBase', 'SRS.Allura.Func.PositioningTableTop', 'SRS.Allura.Func.ChangeSIDByDetector', 'SRS.Allura.Func.MoveBeam', 'SRS.Allura.Func.StoreFluoro', 'SRS.Allura.Func.MoveBeamXY', 'SRS.Allura.Func.APC3DReference', 'SRS.Allura.Func.Roadmap.SmartMask', 'SRS.Allura.Func.RotateFDXDDetector', 'SRS.Allura.Func.AngulateBeam', 'SRS.Allura.Func.FluoroTimeBuzzer', 'SRS.Allura.Func.BiplaneAcquisition', 'SRS.Allura.Qual.ClinicalResponseTimes.Movement', 'SRS.Allura.Func.TableTilting', 'SRS.Allura.Func.Roadmap', 'SRS.Allura.Func.ResetGeometry', 'SRS.Allura.Func.APCStore/Recall', 'SRS.Allura.Func.APCTable', 'SRS.Allura.Conf.UI-RelatedItems', 'SRS.Allura.Func.Z-RotateStandBase(L-ArmRotation)', 'SRS.Allura.Func.Move/ResetWedges', 'SRS.Allura.Func.RotateBeam', 'SRS.Allura.Func.ChangingTableHeight', 'SRS.Allura.Func.APCPredefined', 'SRS.Allura.Func.TableCradle', 'SRS.Allura.Func.SelectTableSide'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'The user wants to:\n<Clinical users/Physicists/ Service users, all clinical areas, ER>\nControl the imaging settings at table side;\nSelect/adjust/control the detector field size, FOV, position of shutters/wedges, Source Image Distance (SID), X-ray channel, fluoro-buzzer, fluoro-grab, fluoro-flavour, Roadmap Pro on/off.'], ['New', 'N.A.']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.InjectorCoupling | ['SRS.Allura.Func.ContrastInjection', 'SRS.Allura.Func.SetInjectorCoupling', 'SRS.Allura.Conf.RemainingComponents', 'SRS.Allura.Intgr.Injectors'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, ER/CR>\nCouple certain automated injector systems;\nPerform timed administration (injection) of contrast medium (precise timing depends on patient’s circulation, injection protocol, and imaging protocol);\nSet X-ray delay;\nPerform manual injection;\nPerform automatic injection;\nCouple/decouple automatic injection to/from acquisition at EPX level;\nMaintain coupling after acquisition;\nPerform acquisition with manual injection while coupling is on;\nAdjust contrast injector parameters at table side.'], ['New', 'N.A.']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.InspectCleanLubricate | ['SRS.Allura.Serv.Maintenance.Planned'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'N.A.'], ['New', 'The user needs to:\n<FSE/RSE/BM, Preventive Maintenance>\nInspect, clean, lubricate, and adjust accordingly to the critical to safety and quality aspects of the system in the image-guided lab.']]), ('text', '\xa0'), ('text', 'Decomission'), ('text', ' ')]
UNS.SmartSuite.IntegratedChecklistandProtocols | ['SRS.Allura.Func.ProcedureCard', 'SRS.Allura.UI.Help', 'SRS.Allura.UI.Help.Customize'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'The user wants to:\n<Clinical users/Physicists/ Service users, all clinical areas, CR/ER>\nMake a scan of their checklist(s) (time-out document, compromising text and images) and display it on the screen;\nAdd Checklists and Protocols (comprising text and images of e.g. injector protocols, kind of disposables etc.) to a Procedure Card;\nRecall Checklists and Protocols from a Procedure Card before or during the procedure.'], ['New', 'N.A.']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.IntegratedDoseAware | ['SRS.Allura.Intgr.DoseAwareXtend'] | [('text', ''), ('table', [['Legacy\n\n', 'The user wants to:\n<Clinical users/Physicists/ Service users, all clinical areas, CR/ER>\nAutomatically receive real-time feedback of radiation exposure to user and staff.'], ['New', 'N.A.']]), ('text', '\n ')]
UNS.SmartSuite.IntegratedHelp | ['SRS.Allura.UI.Help'] | [('text', ''), ('table', [['Legacy\n\n', 'The user wants to:\n<Clinical users/Physicists/ Service users, all clinical areas, CR/ER>\nElectronically search through the user manual for help on specific topics/tasks/problems (e.g. digital help function) integrated into the system.'], ['New', 'N.A.']]), ('text', ''), ('text', ' ')]
UNS.SmartSuite.IntegratedIFU | ['SRS.Allura.UI.Help', 'SRS.Allura.SLS.ListOfApplicableStandards', 'SRS.Allura.Serv.RemoteSWInstall'] | [('text', ''), ('table', [['Legacy\n\n', 'The user wants to:\n<Clinical users/Physicists/ Service users, all clinical areas, CR/ER>\nHave access to the latest/up-to-date Instructions for Use electronically on the system.'], ['New', 'N.A.']]), ('text', ''), ('text', ' ')]
UNS.SmartSuite.IntegratedTraining | ['SRS.Allura.UI.Help'] | [('table', [['Legacy\n\n', 'The user wants to:\n<Clinical users/Physicists/ Service users, all clinical areas, CR/ER>\nAccess training materials on/onto the system (e.g. help materials, integrated applications/ simulations).'], ['New', 'N.A.']]), ('text', ''), ('text', ' ')]
UNS.SmartSuite.Intercom | ['SRS.Allura.Func.Intercom'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'The user wants to:\n<Clinical users/Physicists/ Service users, all clinical areas, CR/ER>\nCommunicate verbally between ER and CR (e.g. by using an Intercom).'], ['New', 'N.A.']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.IntuitiveCollisionDetectionFeedback | ['SRS.Allura.Func.ExceptionHandling'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, ER>\nTake appropriate measures in order to avoid and/or resolve collision based on the information and guidance received;\nReceive visual feedback (e.g. icons) showing what is causing a collision and where.'], ['New', 'N.A.']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.LIHTablePanning | ['SRS.Allura.Func.PositioningIndicationWithoutRadiation'] | [('text', ''), ('table', [['Legacy\n\n', 'The user wants to:\n<Clinical users/Physicists/ Service users, all clinical areas, ER>\nReposition the patient without using fluoroscopy (Zero Dose Positioning) including an indication of the new position on both the lateral as the frontal display, e.g. by using a graphical image area overlay of the resulting image area when the user moves the table, shutter or wedge or changes the detector field size.'], ['New', 'N.A.']]), ('text', ''), ('text', ' ')]
UNS.SmartSuite.Languages | ['SRS.Allura.UI.Clinical-UILanguage'] | [('text', '\xa0'), ('table', [['Legacy', 'The user wants to:\n<Clinical users/Physicists/ Service users, all clinical areas, CR/ER>\nRead from a user interface in his/her native or at least in familiar language;\nSelect one user interface language from a range of user interface languages.'], ['New', 'N.A.']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.LastImageHoldOrAutoCycle | ['SRS.Allura.Func.Cycle', 'SRS.Allura.Func.StoreFluoro', 'SRS.Allura.Func.LastImageHoldOrAutoCycle'] | [('text', ''), ('table', [['Legacy\n\n', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, ER/CR>\nView/review after fluoroscopy and acquisition the last acquired images;\nSelect that the last acquisition run is displayed in a cycle.'], ['New', 'N.A.']]), ('text', '\n')]
UNS.SmartSuite.LightsAndOperationLamp | ['SRS.Allura.Func.RoomLightAndOperationLamp'] | [('table', [['Legacy', 'N.A.'], ['New', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, CR/ER>\nSwitch on/off or dim the room-light and/or operation lamp, optionally synchronized to acquisition.']]), ('text', ''), ('text', ' ')]
UNS.SmartSuite.LinkWorkstation | ['SRS.Allura.Intgr.Interfaces.ControlFromAuxiliarySystems', 'SRS.Allura.Intgr.Interfaces.ControlAuxiliarySystems', 'SRS.Allura.Intgr.Interfaces.RealTimeImageLink', 'SRS.Allura.Func.Dual-AxisAcquisition', 'SRS.Allura.Func.DRAFor3D', 'SRS.Allura.Func.XperCT', 'SRS.Allura.Func.DualPhaseXperCT', 'SRS.Allura.Func.ClassicDRA'] | [('text', ''), ('table', [['Legacy\n\n', 'The user needs to be able to:\n<all users, all\xa0clinical areas, ER/CR>\nHave a direct (real time link) connection with the Interventional WorkSpot and\xa0 Coronary Tools and EchoNavigator and access to relevant Interventional Tools (e.g. StentBoost);\nHave adequate image quality for (3D) reconstructions/information at their Philips workstation(s). '], ['New', 'N.A.']]), ('text', ''), ('text', ' ')]
UNS.SmartSuite.LiveImageView | ['SRS.Allura.Func.StoreReferenceRun/Image', 'SRS.Allura.Func.ActivateExposure', 'SRS.Allura.Intgr.NativeIQ', 'SRS.Allura.Func.DVDRecordControl', 'SRS.Allura.Func.SingleShotFootpedal', 'SRS.Allura.Qual.ClinicalResponseTimes.Acquisition', 'SRS.Allura.Intgr.FHDVideoSources', 'SRS.Allura.Intgr.4KVideoInputOutput', 'SRS.Allura.Func.ActivateFluoro'] | [('table', [['Legacy\n\n', 'The user needs to be able to:\n< Clinical users/Physicists/ Service users, all clinical areas, ER/CR>\nDisplay real-time images from fluoroscopy as well as exposure;\nView live and previously acquired images simultaneously;\nRecord the image stream on a DVD during fluoroscopy/exposure;\nFor monoplane: Save images for reference up to three different reference screens;\nFor biplane: Save images for reference up to two different reference screens per channel. In monoplane mode save images up to 3 reference screens for the active channel.\n'], ['New', 'Have the ability to view multiple previously acquired images (i.e. reference screens) simultaneously with live;\nShow an external source in its native resolution on a 3rd party monitor;\nShow multiple sources in their native resolution;\nUse 4K/UHD input connectivity to display 3rd party sources;\nDeliver artefact-free images distributed with near-zero latency via optical fiber.']]), ('text', ''), ('text', ' ')]
UNS.SmartSuite.LocalStorage | ['SRS.Allura.Intgr.DVDRecorders', 'SRS.Allura.Conf.ImageStorageOptions', 'SRS.Allura.Intgr.Media', 'SRS.Allura.Func.StoreToMedia', 'SRS.Allura.Func.RetrieveFromMedia'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'The user wants to:\n<Clinical users/Physicists/ Service users, all clinical areas, ER/CR>\nStore/ retrieve fluoroscopy and exposure images with DICOM format on/from external destinations (e.g. PACS, USB, CD, DVD);\nStore fluoroscopy and exposure images with DICOM and non-DICOM format on local storage media/devices (e.g. USB, CD, DVD).'], ['New', 'N.A.']]), ('text', '')]
UNS.SmartSuite.ManoeuvreBeam | ['SRS.Allura.Func.MoveBeam', 'SRS.Allura.Func.SwingStandBase', 'SRS.Allura.Func.Z-RotateStandBase(L-ArmRotation)', 'SRS.Allura.Func.SelectTableSide'] | [('text', ''), ('table', [['Legacy\n\n', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, ER>\nManoeuvre the system out of current working area;\nManoeuvre the system around other equipment;\nManoeuvre the system to current working area, either manually, via controls or automated.'], ['New', 'N.A.']]), ('text', ' ')]
UNS.SmartSuite.Marker | ['SRS.Allura.Func.Marker'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, ER/CR>\nPut multiple marks on the live viewport, from TSM, ER and CR;\nMarkers can be adjusted or erased;\nMarkers stay visible during fluoroscopy and exposure;\nMarkers adapt to changes of magnification (zoom, SID or detector format);\nMarkers will stay when copy to ref;\nMarkers will not be stored on image.\n'], ['New', 'N.A.']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.Measurements | ['SRS.Allura.Func.Measurement'] | [('text', ''), ('table', [['Legacy\n\n', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, ER/CR>\nPerform a measurement on an image (angle, ratio, length);\nPerform manual and automatic calibration;\nPerform a distance measurement on TSM Pro.'], ['New', 'N.A.']]), ('text', ''), ('text', ' ')]
UNS.SmartSuite.MechanicalInstallation | ['SRS.Allura.Serv.FSTools', 'SRS.Allura.UI.Help', 'SRS.Allura.Serv.FSInstructions'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'N.A.'], ['New', 'The user needs to:\n<FSE/BM, Installation/Corrective Maintenance/ Upgrade/FCO\xa0>\nMechanically install the system using the appropriate instruction guide in the image-guided lab. ']]), ('text', ''), ('text', ''), ('text', 'Setting to work'), ('text', '')]
UNS.SmartSuite.MechnicalDeinstallation | ['SRS.Allura.UI.Help', 'SRS.Allura.Serv.SingleLabConnection', 'SRS.Allura.Serv.FSInstructions'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'N.A.'], ['New', 'The user needs to:\n<FSE/BM, Corrective Maintenance, Upgrade/FCO/ Deinstallation>\nMechanically deinstall the system in the image-guided lab based on the appropriate instructions;\nMechanically deinstall parts of the system for corrective maintenance & Upgrade/FCO.']]), ('text', '\xa0'), ('text', 'Software installation'), ('text', ' ')]
UNS.SmartSuite.MonitorConfiguration | ['SRS.Allura.UI.PositionMCSFrame', 'SRS.Allura.Conf.MCS-RelatedItems'] | [('text', 'Background: Adjust the system to the optimal viewing position during procedure. To optimize the viewing solution to his needs the user can choose between a set of smaller displays or large screen and adjust the position of the displays on viewing distance and height.'), ('table', [['Legacy\n\xa0\n\xa0', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, ER>\nAdjust the system to his particular viewing needs;\nUse the system in hybrid labs with minimal interference with Laminar Airflow;\nBe able in interventional labs to position the viewing support at 3 sides of the table, doctor, nurse and foot end;\nAdjust the screens on viewing distance and viewing height;\nChoose a wide range of screen configurations to match the intended procedure mix of the lab. Screen configurations range from separate loose screens to screens combined on single booms to large screens or combinations of those.\n'], ['New', 'N.A.']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.MoveToPark | ['SRS.Allura.Func.ParkorStandbyBeam'] | [('text', '\xa0'), ('table', [['Legacy', 'The user wants to:\n<Clinical users/Physicists/ Service users, all clinical areas, ER>\nIn case the FlexArm stand is not needed inter procedural, have no hinder of the system and have it parked/placed outside the working area, for example for cleaning, during the night, or for during a procedures that do not require imaging\nPark the FlexArm in central rail configuration away against the wall at foot side or head side of the room'], ['New', 'N.A.']]), ('text', '\xa0'), ('text', 'Transfer patient from table'), ('text', ' ')]
UNS.SmartSuite.MoveToStandby | ['SRS.Allura.Func.ParkorStandbyBeam'] | [('text', 'Background information: Free up space around the table in case if temporarily X-ray is not needed during the procedure.'), ('table', [['Legacy', 'The user wants to:\n<Clinical users/Physicists/ Service users, all clinical areas, ER>\nMove the FlexArm system to and from a Standby Position by table side controls.\nA standby position is an optimal* position of the system in the room, when imaging is temporary not needed during a procedure.\n* Optimal is defined as where the gantry is out of the way during the procedure'], ['New ', 'N.A.']]), ('text', '\n')]
UNS.SmartSuite.NewIdentityUserInterface | ['SRS.Allura.UI.Design.Application', 'SRS.Allura.Func.PresettingFunctions', 'SRS.Allura.UI.KeyboardMouse', 'SRS.Allura.Func.ProcedureCard', 'SRS.Allura.UI.UIModule', 'SRS.Allura.UI.Design.FlexibleViewing'] | [('text', ''), ('table', [['Legacy', 'The user wants to:\n<Clinical users/Physicists/ Service users, all clinical areas, CR/ER>\nAdjust visual display settings depending on the type of procedure (only relevant features should be displayed);\nHave focus on the clinical task and especially on the images (e.g. the legibility of visual display through better icon size, color, contrast, backlight, etc.).'], ['New', 'N.A.']]), ('text', '\n')]
UNS.SmartSuite.ORTable | ['SRS.Allura.Conf.UI-RelatedItems', 'SRS.Allura.Intgr.Magnus', 'SRS.Allura.UI.UIModule', 'SRS.Allura.Intgr.TruSystem', 'SRS.Allura.Conf.Table-RelatedItems'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0\n\xa0', 'N.A.'], ['New', 'The user needs to be able to\n< Clinical users/Physicists/ Service users, all clinical areas, ER>\nControl all geometry movements from a single control panel;\nAttach and detach the TSO and TSM from the table rail during the clinical procedure;\nChoose different table tops;\nStore and recall table position and set the table to stored positions;\nTo use the system with reverse table top and 180 deg. pivot.']]), ('text', ' ')]
UNS.SmartSuite.OperateSystem | ['SRS.Allura.SLS.movement-integrity', 'SRS.Allura.Func.MovementDisable/Enable', 'SRS.Allura.SLS.unexpected-manual-behavior'] | [('text', 'Background information: ‘In control’ means you feel the system doesn’t move too fast and you lose control. ‘Effective’ meaning the system doesn’t move to slow so you have to wait too long and it takes the ‘smartest’ route to the desired position.'), ('table', [['Legacy\n\xa0', 'N.A.'], ['New', 'The user wants to:\n<Clinical users/Physicists/ Service users, all clinical areas, ER/CR>\nOperate the system such that the movements of the system are effective and in control, both manually and via the operator controls.']]), ('text', '\xa0'), ('text', ' '), ('text', ''), ('text', ''), ('text', ''), ('text', '')]
UNS.SmartSuite.OptimalWorkingPosition | ['SRS.Allura.Func.RotateBeam', 'SRS.Allura.Func.Z-RotateStandBase(L-ArmRotation)', 'SRS.Allura.Func.RotateFDXDDetector'] | [('table', [['Legacy', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, ER>\nMove the FlexArm system to an optimal working position continuous in a Rz rotation range of +135 to -135 degrees while maintaining the selected projection position to create space for e.g. physician, anaesthesia, equipment or visibility on the display.\n\n*optimal = a position in the range of 270 degrees minimizing interference with patient accessibility, equipment and viewing angles of the monitors.'], ['New ', '\xa0N.A.']]), ('text', '\n')]
UNS.SmartSuite.PackingandDisposal | ['SRS.Allura.UI.Help', 'SRS.Allura.Serv.FSInstructions'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'N.A.'], ['New', 'The user needs to:\n<FSE/BM, Deinstallation>\nDispose of the system accordingly to the standards/guidelines.']]), ('text', '\xa0'), ('text', ''), ('text', 'Regulatory, Safety, Security and Legal Needs'), ('text', ' ')]
UNS.SmartSuite.ParallelWorking | ['SRS.Allura.Intgr.Flexvision', 'SRS.Allura.Intgr.Multivision', 'SRS.Allura.Func.PriorityConstraints', 'SRS.Allura.Intgr.SwitchableMonitors', 'SRS.Allura.Func.Concurrency'] | [('text', ''), ('table', [['Legacy\n\n', 'The user wants to:\n<All users, all clinical areas, ER/CR>\nSimultaneously run image acquisitions and work on other tasks in the Control Room on the same patient during X-ray generation;\nReview and process previously and intra-procedurally acquired images\nPerform QA (without interrupting an examination)\nFilm (i.e. select/flag images) and export images;\nPatient administration.\nSimultaneously run image acquisitions and work on other tasks in the Control Room on different patients during X-ray generation;\nSchedule an examination on another patient in advance to the current examination;\nSchedule an examination on another patient during (and without interrupting) the current examination;\nReview, post process, measure, print and transfer images of one patient in parallel to performing an examination on another patient;\nPerform QA during (without interrupting an examination);\nFilm (i.e. select/flag images) and export images;\nDisplay data of the current patient in the ER, coming from an auxiliary system (e.g. PACS) to enable reviewing previous studies of this patient;\nUse the display(s) in the ER to look at prior images of the current patient or to perform a QA or Bolus Chase Reconstruction (BCR) for the current patient;\nAlways see the Live X-ray while simultaneously performing patient administration/scheduling/QA task;\nWork simultaneously on different touch screen modules in different applications in Exam Room and Control Room.'], ['New', 'Tailored configuration of the workspot with the possibility of adding additional monitors (up to 6) with a multiswitch capability.']]), ('text', '\n')]
UNS.SmartSuite.ParkBeam | ['SRS.Allura.Func.ParkorStandbyBeam'] | [('table', [['Legacy\n\n', 'The user wants to:\n<Clinical users/Physicists/ Service users, all clinical areas, ER>\nManually and motorized (user controlled) park/position the system (if temporarily not used) out of the way to improve access to the patient.'], ['New', 'N.A.']]), ('text', ''), ('text', ' ')]
UNS.SmartSuite.PatientAndBeam | ['SRS.Allura.Func.APCTable', 'SRS.Allura.Func.Z-RotateStandBase(L-ArmRotation)', 'SRS.Allura.Func.SwingStandBase', 'SRS.Allura.Func.RotateFDXDDetector', 'SRS.Allura.Func.AngulateBeam', 'SRS.Allura.Qual.ClinicalResponseTimes.Movement', 'SRS.Allura.Func.ChangeSIDByDetector', 'SRS.Allura.Func.DisplayActualPosition', 'SRS.Allura.Func.RotateBeam', 'SRS.Allura.Func.MoveBeam', 'SRS.Allura.Func.XperCT', 'SRS.Allura.Func.APCStore/Recall', 'SRS.Allura.Func.MoveBeamXY'] | [('table', [['Legacy\n\n', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, ER>\nOperate the system from multiple user positions (e.g. Table Side controlled, pedestal side controlled and remotely controlled from the Control Room);\nRotate the frontal and lateral beam simultaneously in biplane mode;\nGet user guidance/feedback about geometry positions, e.g. height, SID, cradle/tilt, shutter and wedge position.\n'], ['New', 'The frontal Arc to rotate independently of the detector and collimator, guaranteeing patient-oriented images in every angulation (Azurion 7 B20/15 LN).']]), ('text', '')]
UNS.SmartSuite.PatientComfort | ['SRS.Allura.Conf.Table-RelatedItems'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'The user wants to:\n< Clinical users/Physicists/ Service users, all\xa0clinical areas, ER>\nProvide patients with a comfortable position on the table (e.g. ergonomic mattress).'], ['New', 'N.A.']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.PatientData | ['SRS.Allura.Func.ProcedureCard', 'SRS.Allura.Func.ImplicitSelectionDefaultStudy', 'SRS.Allura.Func.Merge', 'SRS.Allura.SLS.consistent-patient-and-image-data'] | [('text', ''), ('table', [['Legacy\n', 'The user wants to:\n<All users, all clinical areas, CR>\nPrevent mix-up of patient data/information (e.g. identify viewports with name of current patient, or a warning; resolve patient mix);\nStarting imaging without having entered any patient data (e.g. set-up one of the X-ray protocols as default for emergency cases).'], ['New', 'N.A.']]), ('text', ''), ('text', 'Exam preparation'), ('text', ' ')]
UNS.SmartSuite.PatientDoseFeedback | ['SRS.Allura.Func.DisplayAndReportPatientDose'] | [('text', ''), ('table', [['Legacy\n\n', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, ER/CR>\nMonitor Dose Area Product (DAP) and Air-Kerma (AK; skin-entrance-dose);\nReceive a warning if the accumulated AK is above a configurable limit.'], ['New', 'N.A.']]), ('text', '\n ')]
UNS.SmartSuite.PatientOrientation | ['SRS.Allura.Func.SelectSurgicalView', 'SRS.Allura.Conf.SupportedConfigurations', 'SRS.Allura.SLS.consistent-image-orientation', 'SRS.Allura.Func.PatientOrientation'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, CR/ER>\nTo display the patient’s anatomy (or clinical image) default in such a way that it adheres to the clinical practice (head on top side; left/right as if facing the patient) as commonly used for the various examinations;\nPrevent diagnostic risks in mistaking left/right patient anatomy;\nTo deviate from image display default e.g. adjusting the display orientation at table side, during the case. '], ['New', 'Indicate the correct patient positioning with respect to the system to ensure proper image display of the region of interest.']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.PatientOrientedMovements | ['SRS.Allura.Func.PatientOrientedMovement'] | [('text', ''), ('table', [['Legacy', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, ER>\nTo control the FlexArm projections and movements of the gantry (rotate, angulate, translate)\xa0 by moving the X-ray beam (detector + tube) relative to the longitudinal axis of the table / patient;\nNavigate to the region of interest through automated FlexArm movements: Image beam rotation, roll, prop, dual Rz rotation, and rail movement in such a way that the user has the perception to steer the X-ray beam.'], ['New ', 'N.A.']]), ('text', '\n')]
UNS.SmartSuite.PatientPositioning | ['SRS.Allura.Func.TableLock', 'SRS.Allura.Func.PositioningTableTop', 'SRS.Allura.Func.DisplayOrientation', 'SRS.Allura.Func.DisplayActualPosition', 'SRS.Allura.Func.ChangingTableHeight', 'SRS.Allura.Func.SelectTableSide'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'The user needs to be able to\n< Clinical users/Physicists/ Service users, all clinical areas ER>\nPosition the patient on the table;\nLock the table to prevent unintended table movement;\nAdjust the height of the table for ergonomic purposes;\nPerform interventions via radial, subclavian and groin access; from head, arm and groin position; user standing at left or right side;\nIndicate the proper patient positioning with respect to the system to ensure proper image display;'], ['New', 'Adjust the table position to optimize the clinical setup.\xa0']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.Pedestal | ['SRS.Allura.Conf.UI-RelatedItems', 'SRS.Allura.UI.Workspot', 'SRS.Allura.SLS.patient-contamination'] | [('table', [['Legacy\n\n', 'The user wants to:\n< Clinical users/Physicists/ Service users, VASU/SHD, ER>\nStand in between the pedestal and the table during the procedure;\nMove the pedestal (in relation to footswitch);\nOperate the system with minimum footprint of the pedestal;\nOperate the system with minimum loose cables in the working area;\nOperate the system with minimal cables on the floor (e.g. between wall and pedestal) during imaging procedures and open surgery (to prevent tripping hazard, and blocking transport of equipment);\nOperate the system with minimum interference of wires of the control units;\nEasily and effectively clean the pedestal and related wires;\nMove the Control Module, Imaging Module and TSM from the pedestal to the table and vice versa.'], ['New', 'N.A.']]), ('text', ''), ('text', ' ')]
UNS.SmartSuite.PhysioAcquisition | ['SRS.Allura.Func.PhysioAcquisition', 'SRS.Allura.Intgr.PhysioEquipment(ECG)', 'SRS.Allura.Func.SelectPhysioDisplay'] | [('text', ''), ('table', [['Legacy\n\n', 'The user wants to:\n< Clinical users/Physicists/ Service users, PCI/SHD/EP, ER/CR>\nObtain X-ray image information in relation to the heart cycle e.g. by acquisition, display and storage of physiological signals (limited to electrical signals, i.e. ECG) together with the X-ray images;\nManually switch on/off the overlay of the physiological signals on the X-ray images.\xa0'], ['New', 'N.A.']]), ('text', ''), ('text', ' ')]
UNS.SmartSuite.PointerOnViewingContainer | ['SRS.Allura.UI.Image.Pointer'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, ER/CR> \nControl a visible pointer on a viewport to indicate to others an area of interest of an image (e.g. finger on TSM in ER, or mouse in CR);\nSee the pointer on all instances of that viewport.'], ['New', 'N.A.']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.PowerFailure | ['SRS.Allura.Intgr.ExternalThreePhaseUPS', 'SRS.Allura.Func.PowerFailureBehaviour', 'SRS.Allura.Conf.RemainingComponents', 'SRS.Allura.Intgr.BackupPowerActive', 'SRS.Allura.Conf.Table-RelatedItems', 'SRS.Allura.Intgr.AuxiliarySystems'] | []
UNS.SmartSuite.Pre-procedurePlanning | ['SRS.Allura.SLS.IHE', 'SRS.Allura.Func.DICOM.Import', 'SRS.Allura.Qual.ClinicalResponseTimes.Arch'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'The user should be able to:\n<All users, all clinical areas, CR/ER>\nUse/reuse pre-procedure planning from other systems (e.g. recall projection angles planned on HeartNavigator) for more complex cases.'], ['New', 'N.A.']]), ('text', '\xa0')]
UNS.SmartSuite.Preparation | ['SRS.Allura.Serv.NetworkDiagnosis', 'SRS.Allura.Serv.HardwareTests', 'SRS.Allura.Serv.RemoteConnectivityStatus', 'SRS.Allura.Serv.FunctionalTests'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'N.A.'], ['New', 'The user needs to:\n<FSE/RSE/CAS/BM, Installation/Preventive Maintenance/ Upgrade/FCO/Deinstallation>\nObtain information that is needed to ensure the rightful preparation for a service activity before visiting the system.']]), ('text', '\xa0'), ('text', 'Site readiness check'), ('text', ' ')]
UNS.SmartSuite.Printing | ['SRS.Allura.Func.PrintJobControl', 'SRS.Allura.Func.PrintSheetCreation', 'SRS.Allura.Qual.ClinicalResponseTimes.Arch', 'SRS.Allura.Intgr.LocalPaperPrinters', 'SRS.Allura.Intgr.NetworkFilmPrinters'] | [('text', ''), ('table', [['Legacy\n', 'The user wants to:\n<All users, all clinical areas, ER/CR>\nSelect images and create a film sheet from these images;\nPrint the film sheet;\nPreview the film sheet prior to printing;\nSelect printer and modify printer settings;\nSubmit print job and manage print queues;\nCreate and print film sheet at table side prior/during/after an examination;\nPrint via DICOM and non DICOM printer.'], ['New', 'N.A.']]), ('text', ''), ('text', ' ')]
UNS.SmartSuite.ProcedureCards | ['SRS.Allura.Func.PredefinedAnnotation', 'SRS.Allura.Func.SystemCustomization', 'SRS.Allura.Func.ProcedureCard', 'SRS.Allura.Func.PatientOrientation', 'SRS.Allura.Func.Workflow.Flexibility', 'SRS.Allura.UI.Help.Customize'] | [('text', ''), ('table', [['Legacy\n\n', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, CR/ER>\nSelect the system set-up, edit, customize, and create their own specific ProcedureCards including quick-save during procedure and the ability to save under different groups (e.g. per clinical discipline, hospital department or per user);\nSelect the system set-up (based on the above parameters) from the ProcedureCard with a single click;\nSelect a ProcedureCard from at least 20 default ProcedureCards; at least 3 per clinical area;\nAutomatically select the configured ProcedureCards, based on the RIS/HIS/CIS code of the scheduled procedure;\nSelect a ProcedureCard within a maximum of 3 clicks once the user selects the patient and there is no RIS/HIS/CIS code coupling;\nImport, modify/adapt, configure and export (default) procedural information (ProcedureCards) including at least 6 parameters:\nShort description of procedure;\nSet of X-ray protocols which should be directly available related to the clinical procedure (incl. geometry position, APC position and injector coupling);\nPatient orientation;\nA pre-set screen layout which will be used as default for viewing in ER (FlexVision);\nA pre-set screen layout which will be used as default for viewing in CR (FlexSpot);\nOn-screen display of hospital specific checklists & protocols s in exam room and control room;\nInjection Protocol (only in case off SmartCT).'], ['New', 'N.A.']]), ('text', '\n')]
UNS.SmartSuite.ProvideInformation | ['SRS.Allura.Func.FinishStudy', 'SRS.Allura.SLS.IHE', 'SRS.Allura.Func.ShowStudyInfo', 'SRS.Allura.Intgr.WLMSystems', 'SRS.Allura.Func.SelectStudyFromScheduledList'] | [('table', [['Legacy\n\n', 'The user wants to:\n<All users, all clinical areas, CR>\nProvide the hospital administration with patient/ procedure information during/at the end of the examination (e.g. export onto the hospital information system (e.g. RIS/CIS);\nAccess patient data after closing the examination\nExport information to the hospital system about what (events and setting) actually has been performed by the system during/at the end of the procedure (e.g. MPPS manager).'], ['New', 'N.A.']]), ('text', ''), ('text', ' ')]
UNS.SmartSuite.QA | ['SRS.Allura.Func.VesselAnalysis', 'SRS.Allura.Func.VentricularAnalysis', 'SRS.Allura.Func.QA'] | [('text', 'Background: “2D Quantitative Analysis" (2D-QA) is a post processing software intended to assist physicians through providing quantitative information as additional input for their comprehensive diagnosis decision making process and planning during cardiovascular procedures and for post procedural evaluation. 2D-QA consists of six applications:'), ('text', '\xa0'), ('text', 'The "2D Quantitative Coronary Analysis" application is intended to be used for quantification of coronary artery dimensions (approximately 1 to 6 mm) from 2D angiographic images. It allows the user to quantify the dimensions of a coronary vessel and/or stenosis from a non-subtracted angiogram.'), ('text', '\xa0'), ('text', 'The "2D Quantitative Vascular Analysis" application is intended to be used for quantification of aortic and peripheral artery dimensions (approximately 5 to 50 mm) from 2D angiographic images. It allows the user to quantify the dimensions of an aorta, a peripheral artery, for example but not limited to the carotid, renal, iliac and femoral arteries, or a stenosis in these from a subtracted or a non-subtracted angiogram.'), ('text', '\xa0'), ('text', 'The "2D Left Ventricle Analysis" and the "Biplane 2D Left Ventricle Analysis" applications are intended to be used for quantification of left ventricular volumes and local wall motion from monoplane and from biplane angiographic series, respectively. It allows the user to compute the left-ventricular Ejection Fraction, Cardiac Output and Wall Motion from user-reviewed 2D contours in selected ventriculograms.'), ('text', '\xa0'), ('text', 'The "2D Right Ventricle Analysis" and the "Biplane 2D Right Ventricle Analysis" applications are intended to be used for quantification of right ventricular volumes and local wall motion from monoplane and from biplane angiographic series, respectively. It allows the user to compute the right-ventricular Ejection Fraction, Cardiac Output and Wall Motion from user-reviewed 2D contours in selected ventriculograms.'), ('table', [['Legacy', 'The user needs to be able to:\n< Clinical users/Physicists/ Service users, all clinical areas, ER/CR>\nStart Quantitative Analysis;\nOutline and quantify of contrast-medium filled vessels (2DQCA and 2DQVA) and ventricles (2DLVA and 2DRVA for monoplane and biplane), respectively, in angiographic X-ray images.'], ['New', 'N.A.']]), ('text', '\xa0 '), ('text', ''), ('text', ''), ('text', ''), ('text', ''), ('text', ''), ('text', ''), ('text', ''), ('text', ''), ('text', '')]
UNS.SmartSuite.QualityControl | ['SRS.Allura.SLS.ListOfApplicableStandards'] | [('text', ''), ('table', [['Legacy\n\n', 'The user wants to:\n<Physicist, all clinical areas, CR/ER>\nPerform system quality control checks (in order to assure safety, regulatory compliance and validated system performance (as standardized by NEMA).'], ['New', 'N.A.']]), ('text', '\n ')]
UNS.SmartSuite.QuickSaveDuringProcedure | ['SRS.Allura.Intgr.Flexvision', 'SRS.Allura.Func.APCTable', 'SRS.Allura.Intgr.Flexvision.Snapshot', 'SRS.Allura.Func.APCCommon', 'SRS.Allura.Func.APCReference', 'SRS.Allura.Func.APCStore/Recall', 'SRS.Allura.Func.APC3DReference'] | [('text', ''), ('table', [['Legacy\n\n', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, ER/CR>\nSave image(s);\nSave a changed screen layout and save under a different name;\nStore temporarily the table position with a single click ;\nStore temporarily the position of the system (beam carrier(s) and table), including the detector orientation, FOV, shutters and wedge positions with a single click.'], ['New', 'N.A.']]), ('text', '\n ')]
UNS.SmartSuite.Raillength | ['SRS.Allura.Conf.Stand-RelatedItems'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0\n\xa0', 'The user wants to:\n<Clinical users/Physicists/ Service users, all clinical areas, ER>\nAdapt the system to the size of the room by the option of different rail lengths;\nBe able to move the FlexArm system to the far end of the room to free up space around the table.'], ['New', 'N.A.']]), ('text', '\xa0'), ('text', ' '), ('text', ''), ('text', ''), ('text', '')]
UNS.SmartSuite.ReducePatientMovement | ['SRS.Allura.Func.Roadmap.AutomaticPixelShift'] | [('table', [['Legacy\n\n', 'N.A.'], ['New', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, ER >\nHave means to compensate for the patient movements during a scan thus minimizing image artifacts.']]), ('text', ''), ('text', 'Additional'), ('text', ' ')]
UNS.SmartSuite.ReviewModule | ['SRS.Allura.Conf.UI-RelatedItems', 'SRS.Allura.Func.MosaicOverview', 'SRS.Allura.UI.UIModule', 'SRS.Allura.Func.Cycle', 'SRS.Allura.Qual.ClinicalResponseTimes.Review', 'SRS.Allura.Func.NavigateToImageData'] | [('text', ''), ('table', [['Legacy\n\n', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, CR>\nPerform controlled image review:\nWork with one hand on the review module and one hand on the mouse in order to create an efficient workflow for reviewing and image processing (e.g. ergonomic module design).\nHave improved visibility of icons in the dark (e.g. backlighting);'], ['New', 'N.A.']]), ('text', '\n'), ('text', 'Post-processing'), ('text', ' ')]
UNS.SmartSuite.Roadmap | ['SRS.Allura.Func.Roadmap.DevicePhase.DeviceMask', 'SRS.Allura.Func.Roadmap.SmartMask', 'SRS.Allura.Func.Roadmap.DevicePhase.Misregistration', 'SRS.Allura.Func.Roadmap.DevicePhase', 'SRS.Allura.Func.Roadmap.DevicePhase.LiveSubtract', 'SRS.Allura.Func.Roadmap.DevicePhase.Subtraction', 'SRS.Allura.Func.Roadmap', 'SRS.Allura.Func.Roadmap.DevicePhase.VesselSubtract', 'SRS.Allura.Qual.ClinicalResponseTimes.Acquisition', 'SRS.Allura.Func.Roadmap.DevicePhase.VesselUnSubtract', 'SRS.Allura.Func.Roadmap.AutomaticPixelShift', 'SRS.Allura.Func.Roadmap.Mode', 'SRS.Allura.Func.Roadmap.DevicePhase.Processing', 'SRS.Allura.Func.Roadmap.VesselPhase'] | [('text', ''), ('table', [['Legacy\n\n', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, ER/CR>\nSuperimpose vessel image during live fluoroscopy (on both the frontal as the lateral channel) to guide a device (catheter, guide wire or other interventional device) through the vascular structure during fluoroscopy in all patient positions;\nSuperimpose vessel image in (un-)subtracted mode.'], ['New', 'N.A.']]), ('text', '\n ')]
UNS.SmartSuite.RollSpeed | ['SRS.Allura.Qual.ClinicalResponseTimes.3D-Acquisition'] | [('text', 'Background: To reduce image breathing artifacts in e.g. the abdominal region.'), ('table', [['Legacy\n\xa0\n\xa0', 'The user wants to:\n<Clinical users/Physicists/ Service users, all clinical areas, ER/CR>\nPerform a 3D roll scan fast enough for the patient to hold his/her breath.'], ['New', 'N.A.']]), ('text', '\xa0'), ('text', 'Ease of use'), ('text', ' ')]
UNS.SmartSuite.RoomSetup | ['SRS.Allura.Phys.LaminarAirflowCompatibility', 'SRS.Allura.Func.SwingStandBase', 'SRS.Allura.Func.Z-RotateStandBase(L-ArmRotation)', 'SRS.Allura.Func.MoveBeam', 'SRS.Allura.SLS.patient-contamin-airborne', 'SRS.Allura.SLS.patient-contamination', 'SRS.Allura.Conf.MCS-RelatedItems', 'SRS.Allura.Func.SelectActiveChannel', 'SRS.Allura.UI.PositionMCSFrame', 'SRS.Allura.Func.MoveBeamXY'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'The user wants to:\n< Clinical users/Physicists/ Service users, all\xa0clinical areas, ER>\nPosition system easy and efficiently (manually/from TSO);\nMaintain the sterile field when moving the FD system, while using the system over time;\nPosition system so as to minimally disrupt laminar airflow in OR settings;\nPosition system so that the physician can see other staff during the procedure;\nIn case of biplane, work with both channels or work with either the frontal or lateral channel \nWork with the frontal channel if the lateral channel is in parked position.\nPosition image monitors at a comfortable viewing angle for physician when standing in head/side/arm/left/right position;\nHave a clear view on the monitor from work position.'], ['New', 'Choose a wide range of screen configurations to match the intended procedure mix of the lab. Screen configurations range from separate loose screens to screens combined on single booms to large screens or combinations of those.']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.Safe3DScan | ['SRS.Allura.Func.ClassicDRA'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'N.A.'], ['New', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, ER >\nBe sure that 3D scan can be done without any collisions.']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.SafeSystemUse | ['SRS.Allura.SLS.collision-avoidance', 'SRS.Allura.SLS.fse-contamin', 'SRS.Allura.SLS.unexpected-manual-behavior', 'SRS.Allura.SLS.careless-operator-patient-fall', 'SRS.Allura.Phys.ClimaticEnvironmentalConditions', 'SRS.Allura.SLS.collision-harm-reduction', 'SRS.Allura.SLS.radiation-user-dose', 'SRS.Allura.Phys.MechanicalEnvironmentalConditions', 'SRS.Allura.SLS.prevent-using-incorrect-in-patient-measurements', 'SRS.Allura.SLS.thermal-safety', 'SRS.Allura.SLS.intended-image-quality', 'SRS.Allura.SLS.prevent-lih-live-mixup', 'SRS.Allura.SLS.patient-move-collision', 'SRS.Allura.SLS.replay-path-free', 'SRS.Allura.SLS.DoD', 'SRS.Allura.Phys.MainsPower', 'SRS.Allura.Serv.RemoteInterference', 'SRS.Allura.SLS.ListOfApplicableStandards', 'SRS.Allura.SLS.collision-unintend-ui', 'SRS.Allura.SLS.table-movement-integrity', 'SRS.Allura.SLS.safe-surfaces', 'SRS.Allura.Func.RemoteServiceActiveIndication', 'SRS.Allura.SLS.electr-shock-prevention-fse', 'SRS.Allura.SLS.peripheral-equipment', 'SRS.Allura.SLS.patient-move-fall', 'SRS.Allura.SLS.xray-imaging-failure-detection-and-fault-tolerant', 'SRS.Allura.SLS.image-latency', 'SRS.Allura.SLS.acoustic-noise', 'SRS.Allura.SLS.consistent-image-orientation', 'SRS.Allura.SLS.tipping-equipment', 'SRS.Allura.SLS.careless-manual-operation', 'SRS.Allura.SLS.movements-failure-detection-and-fault-tolerant', 'SRS.Allura.SLS.radiation-patient-dose', 'SRS.Allura.Phys.ElectromagneticCompatibility(EMC)', 'SRS.Allura.Func.SwitchSystemPowerOn', 'SRS.Allura.SLS.electr-shock-prev-staff-patient', 'SRS.Allura.SLS.consistent-patient-and-image-data', 'SRS.Allura.Func.EnableRemoteServiceAccess', 'SRS.Allura.SLS.structural-integrity', 'SRS.Allura.SLS.unintend-ui-patient-falling', 'SRS.Allura.SLS.correct-overlay', 'SRS.Allura.SLS.movement-integrity', 'SRS.Allura.SLS.patient-contamin-airborne', 'SRS.Allura.SLS.patient-contamination', 'SRS.Allura.SLS.fse-safety', 'SRS.Allura.SLS.maintain-minimal-functionality', 'SRS.Allura.SLS.injector-safety'] | [('text', ''), ('table', [['Legacy\n', 'The user wants to:\n<All users, all clinical areas, CR/ER>\nUse a system that is safe to use and complies with the applicable international standards.\nPrevent harm to patients or hospital by:\nUsing a system that provides protection against unacceptable safety risks in hazardous situations (for example, protection against X-ray, high temperature);\nReceiving guidance and instructions on minimizing the risk\xa0of causing harm in hazardous situations.'], ['New', 'N.A.']]), ('text', ''), ('text', ' ')]
UNS.SmartSuite.SecurityProtection | ['SRS.Security.PHD.HospitalNetwork', 'SRS.Security.HN.RemovableMediaBreached', 'SRS.Security.PHD.PatientDataCorrupted', 'SRS.Security.MD.DataBreach', 'SRS.Security.MD.UnavailabilityOfMedicalDevice', 'SRS.Security.MD.HospitalNetwork', 'SRS.Security.PHD.PhysicalDismounting', 'SRS.Security.MD.MalwareOnRemovable media', 'SRS.Allura.SLS.DoD', 'SRS.Allura.Serv.RemoteInterference', 'SRS.Security.PHD.PatientDataLoss', 'SRS.Allura.Func.RemoteServiceActiveIndication', 'SRS.Security.PHD.MedicalDevice', 'SRS.Security.PHD.PhysicalAccess', 'SRS.Security.HN.AttackViaMedicalDevice', 'SRS.Security.MD.AttackViaPhysical DeviceModification ', 'SRS.Security.PHD.RemoteAccessPoint', 'SRS.Security.MD.RemoteAccessPoint', 'SRS.Security.PHD.DataBreach', 'SRS.Allura.Func.EnableRemoteServiceAccess', 'SRS.Allura.Serv.AccessRightsAndAuthorization', 'SRS.Security.MD.PhysicalAccess', 'SRS.Security.MD.ReputationDamage', 'SRS.Security.HN.HospitalNetworkBreached', 'SRS.Security.RM.AttackViaMedicalDevice'] | [('table', [['Legacy\n\n', 'The user wants to:\n<Hospital IT, all clinical areas, CR/ER>\nProtect the patient privacy as well as to prevent loss of information or inability to perform clinical procedures, as a result of security breaches;\nProvide protection against all known security threats;\nReceive guidance and instructions how to protect the system against security threats and how to minimize the vulnerability of the system;\nCreate and manage user accounts.'], ['New', 'Have protection over patient and personal data across the network by using user authentication and encryption protocols;\nComply with the industry security and privacy standards.']]), ('text', '\n'), ('text', ' ')]
UNS.SmartSuite.SettingtoWork | ['SRS.Allura.Func.GracefulDegradation', 'SRS.Allura.Qual.FSResponseTimes', 'SRS.Allura.Serv.Configure', 'SRS.Allura.Func.SystemRestart', 'SRS.Allura.SLS.xray-imaging-failure-detection-and-fault-tolerant', 'SRS.Allura.Qual.IQFactory/ServiceMeasurements', 'SRS.Allura.Func.DisplayUserMessages', 'SRS.Allura.SLS.movements-failure-detection-and-fault-tolerant', 'SRS.Allura.Func.ExceptionHandling', 'SRS.Allura.Serv.UserInterface', 'SRS.Allura.Func.TubeLoadProtectionAndIndication', 'SRS.Allura.Qual.ClinicalResponseTimes.System', 'SRS.Allura.Serv.Customize', 'SRS.Allura.Func.DisplayServiceGuidance', 'SRS.Allura.Serv.Adjust', 'SRS.Allura.SLS.maintain-minimal-functionality', 'SRS.Allura.UI.ServiceUILanguage'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'N.A.'], ['New', 'The user needs to:\n<FSE/RSE/CAS/BM, Installation/ Corrective Maintenance/ Upgrade/FCO>\nPerform the main tasks* that are needed to make sure that the system is operational according to specifications in the image-guided lab.\n\xa0\n*power on for the first time, configuration, adjustments.']]), ('text', '\xa0'), ('text', 'Test and verification'), ('text', ' ')]
UNS.SmartSuite.ShutDown | ['SRS.Allura.Func.SwitchSystemPowerOff', 'SRS.Allura.Qual.ClinicalResponseTimes.System'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, CR>\nShut-down the system.'], ['New', 'N.A.']]), ('text', '\xa0'), ('text', 'Generic'), ('text', 'Hardware'), ('text', ' ')]
UNS.SmartSuite.SitePlanning | ['SRS.Allura.Serv.SingleLabConnection'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'N.A.'], ['New', 'The user needs to:\n<FSE/RSE/BM, Installation/Upgrade/FCO>\nObtain all information that is needed to prepare for the installation/upgrade. ']]), ('text', '\xa0'), ('text', 'Preparation'), ('text', ' ')]
UNS.SmartSuite.SiteReadinessCheck | ['SRS.Allura.Serv.Maintenance.Planned'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'N.A.'], ['New', 'The user needs to:\n<FSE/RSE/BM, Installation>\nObtain information that is needed to ensure that the prerequisites of system installation are met. ']]), ('text', '\xa0'), ('text', 'Unpack and transport'), ('text', ' ')]
UNS.SmartSuite.SmallRoom | ['SRS.Allura.Conf.MinimalRoomSize'] | [('text', 'Background: Context: The user wants to optimize room efficiency and therefore needs to know how much the Reach Over Available Room (ROAR) is. The formula for ROAR is:'), ('text', ''), ('text', '\xa0'), ('table', [['Legacy\n\xa0\n\xa0', 'The user wants to:\n<Clinical users/Physicists/ Service users, all clinical areas, ER>\nFit FlexArm in smaller rooms (35m² and more);\nCalculate the ROAR by having Imaging Reach of the System data available.'], ['New', 'N.A.']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.SmoothInteraction | ['SRS.Allura.SLS.movement-integrity', 'SRS.Allura.Func.PriorityConstraints', 'SRS.Allura.SLS.unexpected-manual-behavior', 'SRS.Allura.Func.SelectTableSide'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0\n\xa0', 'The user wants to:\n<Clinical users/Physicists/ Service users, all clinical areas, ER>\nMove the system in such a way that the acceleration and de-acceleration, the speed and the logic of automated movements is perceived by the user as being effective and in a controlled manner (not too fast, not too slow);\nMove the system with minimal* user interaction.\n\n*) minimal is with the least possible amount of user interaction.'], ['New', 'N.A.']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.SoftwareInstallation | ['SRS.Allura.Serv.RemoteConnectivityStatus', 'SRS.Allura.Serv.Upgrading', 'SRS.Allura.Serv.InstallSoftware'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'N.A.'], ['New', 'The user needs to:\n<FSE/RSE/BM, Corrective Maintenance/ Upgrade/FCO>\nInstall software on the system that match the workflow prerequisites*.\n\xa0\n*hardware and software compatibility.']]), ('text', '\xa0'), ('text', 'Inspect, clean, and lubricate'), ('text', ' ')]
UNS.SmartSuite.SoundsRoom | ['SRS.Allura.SLS.acoustic-noise', 'SRS.Allura.Phys.Noise'] | [('text', ''), ('table', [['Legacy\n\n', 'N.A.'], ['New', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, CR/ER>\nHave a low level of noise in the room (no disturbing noise in the room, except when immediately attention required,\xa0like heart stops).']]), ('text', '\n')]
UNS.SmartSuite.StartExamination | ['SRS.Allura.Func.ReopenFinishedStudy', 'SRS.Allura.Func.SelectStudyFromSuspendedList', 'SRS.Allura.Intgr.Flexspot.AdditionalWorkspot', 'SRS.Allura.Qual.ClinicalResponseTimes.Admin', 'SRS.Allura.Func.SelectStudyFromScheduledList'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, ER/ CR> \nStart the examination from the Control Room;\nStart the examination from the Exam Room from the Additional FlexSpot.'], ['New', 'N.A.']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.StartUp | ['SRS.Allura.Func.SwitchSystemPowerOn', 'SRS.Allura.Phys.MainsPower', 'SRS.Allura.Func.SwitchSystemPowerOn.VideoOnly', 'SRS.Allura.Qual.ClinicalResponseTimes.System', 'SRS.Allura.Func.X-RayDisable/Enable'] | [('text', ''), ('table', [['Legacy\n\n', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, CR>\nStart-up the whole system;\nEnable/disable X-ray to prevent accidental exposures (e.g. during cleaning, patient transfer etc.).\nTurn on the displays only in FlexSpot configuration, FlexVision (Pro) and Switchable Monitors (i.e. to access the non-X-ray sources).'], ['New', 'N.A.']]), ('text', '\n')]
UNS.SmartSuite.Sterility | ['SRS.Allura.Conf.Stand-RelatedItems', 'SRS.Allura.UI.TSM.Workspot', 'SRS.Allura.UI.UIModule'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'The user needs to be able to:\n< Clinical users/Physicists/ Service users, all\xa0clinical areas, ER>\nCover system parts with sterile covers in the immediate vicinity of patients.\nHave a draping solution for TSM with regard to:\nSterility: positioning and control over sterile sections\nVisibility: transparent without reflections\nResponsiveness: response rate using (dirty) gloves comparable to TSM without draping.'], ['New', 'N.A.']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.SystemNoise | ['SRS.Allura.Phys.Noise'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0\n\xa0', 'The user wants to:\n<Clinical users/Physicists/ Service users, all clinical areas, ER/CR>\nHave a low level of noise in the room of the system when in use (no disturbing system noise when moving the system, except when immediately attention required).'], ['New', 'N.A.']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.TSMNavigation | ['SRS.Allura.UI.TSM.Workspot', 'SRS.Allura.Intgr.Interfaces.ControlAuxiliarySystems'] | [('text', '\xa0'), ('table', [['Legacy', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, CR/ER>\nSwitch between modalities displayed on the TSM Pro within a maximum of 2 user interactions.'], ['New', 'N.A.']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.TSMPlugandPlay | ['SRS.Allura.Intgr.Interfaces.ControlAuxiliarySystems'] | [('text', ''), ('table', [['Legacy', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, ER/CR>\nAutomatically display 3rd\xa0party modality information on the TSM when the 3rd\xa0party is connected to the network.'], ['New', 'N.A.']]), ('text', '\n ')]
UNS.SmartSuite.TSMtouchpad | ['SRS.Allura.Intgr.Flexvision.KeyboardMouse'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'The user wants to:\n<Clinical users/Physicists/ Service users, all clinical areas, ER>\nUse the TSM as touchpad/ keyboard for manipulating/ navigating the FlexVision (Pro) at table side, also under a sterile bag.'], ['New', 'Perform new interactions at tableside when compared to previous generations of Allura systems such as:\nSee live clinical images in 2D;\nNavigate images in 2D;\nUse mouse pointer which is visible in all screens showing X-ray images, to improve communication in ER/CR;\nControl the shutters and wedges from TSM.']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.Table | ['SRS.Allura.Func.PivotRotationTableBase', 'SRS.Allura.Func.MovementDisable/Enable', 'SRS.Allura.Func.PositioningTableTop', 'SRS.Allura.Func.SwivelMovementTableBase', 'SRS.Allura.Func.TableTilting', 'SRS.Allura.Func.ChangingTableHeight', 'SRS.Allura.Conf.Table-RelatedItems', 'SRS.Allura.Func.TableCradle'] | [('text', '\xa0'), ('table', [['Legacy \n\xa0', 'The user wants to:\n< Clinical users/Physicists/ Service users, all\xa0clinical areas, ER>\nPosition the table top in any orientation and height;\nMove the table top to position the region of interest in the center of the X-ray beam;\nMove the table top height (up/down) to optimize the working height for the physician;\nMove the table top to predefined positions.'], ['New', 'N.A.']]), ('text', '\xa0'), ('text', ''), ('text', ''), ('text', '')]
UNS.SmartSuite.TableConfiguration | ['SRS.Allura.Conf.Table-RelatedItems'] | [('text', ''), ('table', [['Legacy\n\xa0', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, ER>\nMake use of the Allura/Azurion system in endovascular and hybrid procedures\nContext: For patient support at installation the user can choose between Angio Table and OR Table Magnus or Hillrom;\nMake use of the system in endovascular and hybrid procedures in one room and optimize for these kind of procedures;\nUse special patient accessories support that is required in open procedures with C-arc in imaging or in parking position.\xa0'], ['New', 'N.A.\xa0']]), ('text', ''), ('text', ' ')]
UNS.SmartSuite.TableSetup | ['SRS.Allura.Func.PivotRotationTableBase', 'SRS.Allura.Func.PositioningTableTop', 'SRS.Allura.Func.SwivelMovementTableBase', 'SRS.Allura.Qual.ClinicalResponseTimes.Movement', 'SRS.Allura.Func.TableTilting', 'SRS.Allura.Func.ChangingTableHeight', 'SRS.Allura.Func.TableCradle'] | [('text', '\xa0'), ('table', [['Legacy \n\xa0', 'The user wants to:\n< Clinical users/Physicists/ Service users, all\xa0clinical areas, ER>\nPosition the table top in any orientation and height;\nMove the table top to position the region of interest in the center of the X-ray beam;\nMove the table top height (up/down) to optimize the working height for the physician;\nMove the table top to predefined positions.'], ['New', 'N.A.']]), ('text', '\xa0'), ('text', ''), ('text', ''), ('text', '')]
UNS.SmartSuite.TableSideTextEntry | ['SRS.Allura.Intgr.Flexvision', 'SRS.Allura.Intgr.Flexvision.KeyboardMouse'] | [('table', [['Legacy\n\n', 'The user wants to:\n< Clinical users/Physicists/ Service users, all\xa0clinical areas, ER>\nAllow table side text entry without compromising sterility/ hygiene.'], ['New', 'N.A.']]), ('text', ''), ('text', 'End examination'), ('text', ' ')]
UNS.SmartSuite.TaskGuidance | ['SRS.Allura.Func.ProcedureCard', 'SRS.Allura.UI.Help', 'SRS.Allura.UI.Help.Customize'] | [('text', ''), ('table', [['Legacy\n\n', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, CR/ER>\nReceive displayed step-by-step guidance/ feedback for purpose of:\nPreparing/ performing clinical tasks efficiently (e.g. rotational scan);\nMinimizing the need for re-training;\nAllowing existing staff to train new staff;\nShow this taskguidance on large screen (yes/no) or only at TSM;\nOpen this taskguidance yes or no.'], ['New', 'N.A.']]), ('text', ''), ('text', ' ')]
UNS.SmartSuite.TestandVerification | ['SRS.Allura.Serv.Logging', 'SRS.Allura.Qual.FSResponseTimes', 'SRS.Allura.Func.ServiceStudies', 'SRS.Allura.Qual.IQFactory/ServiceMeasurements', 'SRS.Allura.Serv.AcceptanceTestReport', 'SRS.Allura.Intgr.Interfaces.ServiceDataRetrieval', 'SRS.Allura.Serv.VerificationTests', 'SRS.Allura.Serv.UserInterface', 'SRS.Allura.Serv.Adjust', 'SRS.Allura.UI.ServiceUILanguage'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'N.A.'], ['New', 'The user needs to:\n<FSE/RSE/CAS/BM, Installation/Corrective Maintenance/ Upgrade/FCO>\nCollect evidence, by performing test and verify the system in the image-guided lab, on critical parameters related to safety and performance, to indicate that the system is operational according to manufacturer specification.\n\n\xa0*a functional test makes sure that the system, subsystem or Field Replaceable Units functions correctly.']]), ('text', '\xa0'), ('text', 'Handover'), ('text', ' ')]
UNS.SmartSuite.TouchScreenModule | ['SRS.Allura.Intgr.Flexvision', 'SRS.Allura.UI.Workspot', 'SRS.Allura.UI.TSM.Image', 'SRS.Allura.Intgr.Interfaces.ControlAuxiliarySystems', 'SRS.Allura.Intgr.Interfaces.RealTimeImageLink', 'SRS.Allura.Func.NormalSubtraction', 'SRS.Allura.Func.Move/ResetShutters', 'SRS.Allura.Func.Cycle', 'SRS.Allura.UI.TSM.Workspot', 'SRS.Allura.UI.Image.Pointer', 'SRS.Allura.Func.SetImageEnhancement', 'SRS.Allura.Func.APC3DReference', 'SRS.Allura.Func.PresettingFunctions', 'SRS.Allura.Intgr.Interfaces.ControlFromAuxiliarySystems', 'SRS.Allura.UI.UIModule', 'SRS.Allura.Func.DynamicViewing', 'SRS.Allura.Func.PanZoom', 'SRS.Allura.Func.APCStore/Recall', 'SRS.Allura.Func.APCTable', 'SRS.Allura.Conf.UI-RelatedItems', 'SRS.Allura.Func.Move/ResetWedges', 'SRS.Allura.Intgr.Flexvision.KeyboardMouse', 'SRS.Allura.Func.APCPredefined', 'SRS.Allura.Func.APCReference'] | [('text', ''), ('table', [['Legacy', "The user wants to:\n<Clinical users/Physicists/ Service users, all clinical areas, ER>\nPlace Touch Screen Module close to user without impact on sterility;\nPosition the Touch Screen Module at head/side/foot side of table;\nPosition the Touch Screen Module at a mounting device on a cart;\nControl the Touch Screen Module sterile/unsterile;\nUse up to 3 TSM's allowing both synchronized and parallel operation.\nChange contrast and brightness settings, zoom pan, subtraction on-off, select examination types (EPX) and recall previously stored geometry positions, APC control;\nControl (view and manipulate) 3rd\xa0party systems (incl. UltraSound) at table side with/without the use of a single mouse (in case of FlexVision (Pro), the mouse should control FlexVision (Pro) and (CWIS) third party;\nHave full control of large display FlexVision Pro (layout, applications, pre-sets and viewport sizes);\nOn the TSM pro:\nView and navigate live clinical images on TSM in 2D;\nPlay runs on TSM;\nManually drag shutters/wedges on TSM.\nControl a mousepointer on the screen (s) to improve communication in ER and CR;\nView the user interface on the TSM Pro of the connected modality as presented by the 3rd\xa0party;\nConfigure the order of Icon appearance on TSM Pro based on user preference\nControl (view and interact) 3rd\xa0party systems at table side using touch screen operation;\nChange contrast and brightness settings, zoom pan, subtraction on-off, vascular trace, select examination types (EPX) and recall previously stored geometry positions, APC control."], ['New', 'N.A.']]), ('text', ''), ('text', ' ')]
UNS.SmartSuite.TouchScreenModuleFluidresistance | ['SRS.Allura.UI.TSM.Workspot'] | [('text', ''), ('table', [['Legacy', 'The user wants to:\n<Clinical users/Physicists/ Service users, all clinical areas, ER>\nOperate the TouchScreen Module normally when exposed to typical amount of fluids used during the procedure.'], ['New', 'N.A.']]), ('text', ''), ('text', ' ')]
UNS.SmartSuite.TransferPatientFromTable | ['SRS.Allura.Func.ChangingTableHeight', 'SRS.Allura.Func.TableLock', 'SRS.Allura.Intgr.Magnus'] | [('text', '\n'), ('table', [['Legacy\n\n', 'The user wants to:\n<Clinical users/Physicists/ Service users, all clinical areas, ER>\nLower the table top allowing the patient to get off the table in a comfortable way;\nLower the table top allowing the staff to transfer the patient from the table in a comfortable way;\nLock the table top preventing unintended table movement during patient transfer;\nTo use the system in combination with an OR table;\nEasily put aside or remove and mount (without sliding) the flush mounted TSM (e.g. for allowing the transfer of the patient, working on other side).'], ['New', 'N.A.']]), ('text', ''), ('text', 'Cleaning'), ('text', ' ')]
UNS.SmartSuite.UnpackandTransport | ['SRS.Allura.Phys.MechanicalEnvironmentalConditions', 'SRS.Allura.Serv.FSInstructions', 'SRS.Allura.Phys.ClimaticEnvironmentalConditions'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'N.A.'], ['New', 'The user needs to:\n<FSE/BM, Installation/Corrective Maintenance/ Upgrade/FCO>\nHave guidance to unpack (parts of) the system components.\nHave options to transport* (parts of) the system for installation at the hospital.\n*Transport is getting the equipment from unloading the truck at the hospital to the final destination in the hospital.The goal is to take all obstacles into account, like doors, elevator, corridor (bends), etc.']]), ('text', '\xa0'), ('text', 'Mechanical installation'), ('text', ' ')]
UNS.SmartSuite.UseInterventionalTools | ['SRS.Allura.Intgr.Interfaces.ControlFromAuxiliarySystems', 'SRS.Allura.Intgr.Interfaces.ControlAuxiliarySystems', 'SRS.Allura.Intgr.Interfaces.RealTimeImageLink', 'SRS.Allura.Intgr.Interfaces.DataExchange'] | [('text', ''), ('table', [['Legacy', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, ER/CR>\nTo overlay fluoroscopy images on reconstructed data on any 45 degrees position in a range of Z-rotation range of +135 to -135 degrees, taking into account movement compensation and interactions (excluding table tilt, table cradle and table pivot)'], ['New', 'N.A.']]), ('text', '\n')]
UNS.SmartSuite.UserInterfaceAccess | ['SRS.Allura.Conf.UI-RelatedItems', 'SRS.Allura.UI.TSM.Image', 'SRS.Allura.Serv.RemoteDesktop.ClinicalMode', 'SRS.Allura.Func.PriorityConstraints', 'SRS.Allura.Func.SelectTableSide', 'SRS.Allura.Func.Concurrency'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0\n\xa0', 'The user wants to:\n<Clinical users/Physicists/ Service users, all clinical areas, ER>\nControl the system at all tableside positions (head-end, foot-end, physician side, nurse side);\nRemotely control the system (e.g. via pedestal or from CR);\nView and interact with clinical images and runs\xa0on the TSM;\nSwitch between modalities using touch and swipe functionality.'], ['New', 'N.A.']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.UserInterfaceOrientationConcepts | ['SRS.Allura.Func.SelectSurgicalView', 'SRS.Allura.Func.PatientOrientation', 'SRS.Allura.Func.DisplayOrientation', 'SRS.Allura.Func.RotateBeam', 'SRS.Allura.Func.BeamShapingControls', 'SRS.Allura.Func.SelectTableSide'] | [('text', '\n'), ('table', [['Legacy\n\n\n\n\n\n', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, ER>\nOperate User Interface (UI) control devices (e.g. joystick) intuitively, meaning that the movement direction of table/stand and shutters/wedges should be:\nAs much the same as the relative direction of the UI control device;\nAs expected by taking into account:\nThe location of the physician;\nThe type of procedure (e.g. ‘diagnostic view’ for regular procedures ‘surgical view’ for biopsy procedures);\nThe orientation of the displayed image.'], ['New', 'N.A.']]), ('text', '\n')]
UNS.SmartSuite.UserSpecificSettings | ['SRS.Allura.Func.Display/SetTimeAndDate', 'SRS.Allura.Func.ModifyPhysicianList', 'SRS.Allura.UI.Help.Customize', 'SRS.Allura.Func.SystemCustomization'] | [('text', ''), ('table', [['Legacy\n\n', 'The user wants to:\n<Clinical users/Physicists/ Service users, all clinical areas, CR/ER>:\nDefine system specific settings;\nTransfer specific settings from one system to another;\nDefine procedure specific settings for the hospital;\nCustomize hospital specific information e.g. physician list, ProcedureCards, dose in AK/DAP;\nImport/export ProcedureCards.\n'], ['New', 'N.A.']]), ('text', ''), ('text', ''), ('text', ''), ('text', ''), ('text', ''), ('text', ''), ('text', 'Management'), ('text', ' ')]
UNS.SmartSuite.VerticalArmConfiguration | ['SRS.Allura.Conf.Stand-RelatedItems'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0\n\xa0', 'The user wants to:\n<Clinical users/Physicists/ Service users, all clinical areas, ER>\nPlace the ceiling mounted system in high and low ceiling rooms with a resp. heights of 310 cm (FlexMove), 290 cm (Long L-arm) and 270 cm (Short L-arm).'], ['New', 'N.A.']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.VeteransAdministration | ['SRS.Allura.SLS.VA'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'The user wants to:\n<All users, all clinical areas, ER/CR>\nUse a system that complies with the mandatory requirements from Veterans Administration (VA)'], ['New', 'N.A.']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.VideoStreamIntegration | ['SRS.Allura.Intgr.VideoSlave'] | [('text', '\xa0'), ('table', [['Legacy', 'The user wants to:\n<All users, all clinical areas, CR/ER>\nStream video output signal to other rooms for purpose of i.e. communication, teaching and recording.'], ['New', 'N.A.']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.VideoSwitching | ['SRS.Allura.Intgr.VideoDisplayInterface', 'SRS.Allura.Intgr.SwitchableMonitors'] | [('text', ''), ('table', [['Legacy\n\n', 'The user needs to be able to:\n< Clinical users/Physicists/ Service users, all clinical areas , CR>\nEasily switch X-ray and external sources across the monitors in the Exam room (non-FlexVision).'], ['New', 'N.A.']]), ('text', ''), ('text', ' ')]
UNS.SmartSuite.ViewingNavigation | ['SRS.Allura.Func.MosaicOverview', 'SRS.Allura.Func.SelectViewingStudy', 'SRS.Allura.Func.DynamicViewing', 'SRS.Allura.Func.Cycle', 'SRS.Allura.Func.ShowStudyInfo', 'SRS.Allura.Qual.ClinicalResponseTimes.Review', 'SRS.Allura.Func.NavigateToImageData', 'SRS.Allura.Func.DisplayStudyList', 'SRS.Allura.Func.SwitchDisplayFocus'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, CR/ER>\nReview the acquired images of the current examination or images from earlier examinations (from same and different patients);\nRun/replay image by image;\nScan through available image data for overview and selection of relevant images;\nClearly see when the displayed patient is not the acquisition patient.'], ['New', 'N.A.']]), ('text', ''), ('text', '')]
UNS.SmartSuite.ViewingROI | ['SRS.Allura.Func.AngulateBeam', 'SRS.Allura.Func.RotateBeam'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'N.A.'], ['New', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, ER>\nView from various points the region of interest to ensure a correct 3D impression of the anatomy with optimal imaging conditions.']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.Workspot | ['SRS.Allura.UI.Workspot', 'SRS.Allura.Intgr.Interfaces.ControlFromAuxiliarySystems', 'SRS.Allura.Intgr.Interfaces.ControlAuxiliarySystems', 'SRS.Allura.Intgr.Interfaces.RealTimeImageLink', 'SRS.Allura.UI.TSM.Workspot'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, CR/ER>\nPerform all activities, corresponding to his/her role in the hospital, from a single control point at the location of his/her preference; Control auxiliary systems on those work spots.'], ['New', 'N.A.']]), ('text', '\xa0'), ('text', 'Imaging'), ('text', ' ')]
UNS.SmartSuite.X-rayBeamShaping | ['SRS.Allura.UI.TSM.Image', 'SRS.Allura.Func.Move/ResetWedges', 'SRS.Allura.Func.AutomaticWedgeFollowing', 'SRS.Allura.Func.SetElectronicShutters', 'SRS.Allura.Qual.ClinicalResponseTimes.Movement', 'SRS.Allura.Func.Move/ResetShutters', 'SRS.Allura.Func.BeamShapingControls', 'SRS.Allura.Func.ShutterAndWedgeIndicationWithoutRadiation'] | [('text', '\n'), ('table', [['Legacy\n\n', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, ER/CR>\nControl the radiated area;\nRestrict X-ray radiation to the clinical region of interest e.g.:\nManually positioning shutters;\nManually closing/opening shutters (in horizontal and vertical direction);\nAutomatically position wedges (depending on stand rotation and angulation angles) while moving the beam around the patient (should be configurable depending on the procedure);\nManually adjust wedges (also if ‘Automatic Wedge Follow’ is active);\nPrevent over-exposure at clinical region of interest (e.g. manually adjusting wedges that reduce X-ray intensity);\nBlank part of the displayed image by means of image cropping;\nSet shutters (positioning, opening/closing) and wedges (moving), with or without radiation;\nView actual and new positions on display (with and without radiation);\nSee the anticipated radiated area without radiation (e.g. displayed as overlay).\nAllow (graphical) control of shutters/wedges on the TSM;\nView actual and new positions on display and TSM Pro (with and without radiation) in relation to X-ray image.'], ['New', 'N.A.']]), ('text', '\n')]
UNS.SmartSuite.X-rayDetectorFieldSize | ['SRS.Allura.Conf.XperFDSeriesConfigurations', 'SRS.Allura.Conf.SupportedConfigurations', 'SRS.Allura.Func.SelectDetectorFieldsize'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'The user wants to:\n< Clinical users/Physicists/ Service users, all clinical areas, ER/CR> \nZoom in on displayed image;\nView the region of interest by select/adjust/switch Fields of View (FoV);\nIdentify the current level of magnification during FOV adjustment;\nHave\xa0the possibility to select a predefined Field Of View upon selection of a specific procedure.\nWith the FD12:\nSee the aortic arch and the valve in one view;\nSee the whole coronary tree in one view (cardiac myopathy);\nSee both the needle insertion point (subclavian or axillary vein) and the whole wire/lead towards the right atrium in one view.'], ['New', 'N.A.']]), ('text', '\xa0'), ('text', ' ')]
UNS.SmartSuite.processingfocus | ['SRS.Allura.Func.SelectProcessingFocus'] | [('text', '\xa0'), ('table', [['Legacy\n\xa0', 'The user wants to\n< Clinical users/Physicists/ Service users, all clinical areas, ER/CR>\nSelect if the processing applies to the images of both channels or to only one channel.'], ['New', 'N.A.']]), ('text', '\xa0'), ('text', ' ')]